## Original
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Anesthesia Department
Equipments
| Anesthesia Machine: Gas delivery: 3 gas version- Nitrous oxide, oxygen and air. Integrated safety functions: The Oxygen-Ratio Controller should guarantee a minimun oxygen in the fresh-gas flow. N20 cut off mechansim should ensure that no N2O is delivered in case of an oxygen failure. Audible oxygen supply failure alarm with oxygen flush. Pin index trolley: 2 yokes for oxygen and 2 yokes for N2O and 2 pressure gauges for oxygen and 2 for N2O: 2/3 vaporizer confirguation, Tool-free installation selectatic manifold interlock and isolates vaporizers. Number of position:2 or 3, Vaporizers: Halothane, Isoflurane or Sevoflurane. Ventilator: electronically controlled and driven/ pneumatically controlled and driven. Ventilator operating modes- volume control, pressure control with (SIMV/pressure support ,ASB) Tidal volume range: 20 to 1500 mL: One motion from mechanical to manual mode. Ventilator screen with parameter display with continuous monitoring of inspiratory oxygen, breathing frequency, airway pressure, minute volume and tidal volume. Gas supply: Pipeline connection and cylinder input. Power supply: 220 -240V, 50 HZ. Suitable UPS, Power cord with round 3 pin top. | Set | √ | √
| Ventlator: (ICU): Modes of ventilation: Assist/Control (A/C), SIMV, BIPAP, Spontaneous: Breath types: VCV in A/C or SIMV, PCV in A/C or SIMV, PSV SPONT or SIMV, Mandatory Settings: Patient range: adult and pediatric, Ventilation frequency: 2 to 80 bpm. Inspiration time: 0.2 - 10s, Tidal volume: 40 - 2000 ml, Inspiratory flow: 0 - 180 L/min, Inspiratory pressure: 5- 100 cmH2O, PEEP: 0 - 35 cmH2O, Plateau: 0 to 2 seconds, Flow pattern: Square or decelerating ramp, Oxygen concentration: 21% - 100%, Trigger sensitivity: 1 to 20 L/min, Operation-Set Alarm Limits: Heart rate, low insp. pressure, high pressure, low minute volume, high tidal volume, low tidal volume, airway pressure, apnoea alarm, Ventilator Status Indicator: Alarm, caution, normal, VENT INOP, safety valve open, On AC/battery charging, On internal battery, Manual INSP: delivers one mandatory breath, Operating data: Mainpower connection-100-240V, internal battery- approx 60 minutes, Standard accessories: collectors vials, ventilator breathing circuits with water traps:- adult and pediatric, Electronic humidifier with humidification chamber, standard cart, bacterial filters, test lung, and flex arm. Power supply: 220-240V. Suitable UPS, Power cord with round 3 pin top. should have one set of test lung. | Set | √ | √
| Portable ventilator: Mode; Assist/Control, SIMV, Spon. Breath types: VCV, PCV, PSV. Mandatory setting: Respiratory rate- 1 to 80 breath/min, Tidal volume- 50 to 2000 ml, Inspiratory time- 0.2 to 5 seconds, Spontaneous setting- Pressure support: 0 to 70 cmH2O, Sensitivity: 3 to 25 lpm, Common setting: PEEP/CPAP: 0 to 20 cmH2O, Trigger sensitivity: 3 to 25lpm, Oxygen %: 21% to 100%, Operator set alarm limits: Alarms classified as high and medium priority with audible and visual presentations, should have option to monitor respiratory mechanics (negative inspiratiry force, etc) Internal battery: operating time of more than 4 hours and rechargable, Power line 220 - 240 V, 50 Hz, Power cord with round 3 pin top. Accessories: Resuable adult and pediatric circuit and bacterial filters. | Set | √ | √
| Defibrillator machine: ECG Monitoring: 3-lead ECG cable, paddles and 10nos of pacing pads. Screen: High resolution display- heart rate, ECG, SPO2, NIBP, alarm on/off, AED functions and prompts, defibrillator test function, self test functions, pacer functions and delivered energy. Defibrillator: Biphasic waveform. Energy selection-upto 200 joules. Selected using control on sternum paddle or device front pannel. Quick charge time. Monitor display both selected and delivered energy. Charge controls on paddles and on device front pannel. Adult paddles slides off to expose pediatric paddles. AED function- auto-analyze and auto energy level selection, screen prompts and voice prompts. Recorder: Manual or automatic printing modes, on/off control on front pannel and on paddle. With the following modes: NIBP, EtCO2, IBP, Pacemaker, Temperature. Rechargeable battery. Power cord with three round pin plug top. | Set | √ | √ | √
| AED-Automatic External Defibrillator : Fully automatic, L+B11CD screen, Electrodes pads for adult and pediatric patients, Battery powered with alternate AC adapter, Compact unit. Voice prompt instruction- check airway, apply pad, etc. (should supply 10nos of pacing pads each with the AED) Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √ | √ | √ | √
| Intubating Fiberscope with light source (Adult): Working length of 55 cm to 65 cm flexible probe, 1.4 mm multipurpose channel, 90' field of vision, 2 - 50 mm focal range, 260' tip deflection and 25 mm angulation, and plus to minus diopter adjustment. Accessories: Cleaning brush, light source and adapters & Case. Power supply: 220-240V. Power cord with round 3 pin top. | Set | √
| Intubating Fiberscope with light source (Pediatric) : Flexible probe with multipurpose channel, 90' field of vision, 2 - 50 mm focal range, 260' tip deflection and 25 mm angulation, and plus to minus diopter adjustment. Accessories: cleaning brush, adapters, light source & case. Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √
| Intubating Fiberscope with light source (neonate) : Flexible probe with multipurpose channel, 90' field of vision, 2 - 50 mm focal range, 260' tip deflection and 25 mm angulation, and plus to minus diopter adjustment. Accessories: cleaning brush, adapters, light source & case (with AMBUSCOPE FLEXIBLE INTUBATING SCOPE MONITOR). Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √
| AMBUSCOPE FLEXIBLE INTUBATING SCOPE neonate | Piece | √
| AMBUSCOPE FLEXIBLE INTUBATING SCOPE adult | Piece | √
| AMBUSCOPE FLEXIBLE INTUBATING SCOPE pediatric | Piece | √
| Difficlut Airway Intubation Aid: aScope 3 system for difficult intubation: should have (1) high resolution reusable monitor, with inbuilt rechargeble battery. Power supply to power the set as well as recharge the battery. (2) 3 size ascope for Adult, Pediatric, and infant, all with connection for monitor, control lever for easy manupulation of the tip, suction port, working channel, etc. | Set | √
| Infusion Pump: Pump type: Peristaltic finger type, Control method: Drip rate control, Required infusion set: Standard PVC IV set with 3.4 - 4.5 mm OD tubing, Flow rate range: 1 - 450 ml/h, Purge rate: more than 450 ml/h, Alarm: Door open, Occlusion, Low battery, Internal malfunction, Air in line, Infusion complete, Empty container, Reminder, abnormal drips. Mounting: Pole or Bed frame. Power supply: 220 - 240V =/- 10% 50 Hz.power cord with round 3 pin top | Set | √ | √
| Syringe Pump: Flow rage (ml/h), 0.1 - 200 ml/h 10 ml syringe, 0.1 - 300 ml/h, 5 ml to 50 ml syringe, (0.1 ml/h increments), Rapid flow rate 200 - 900 ml/h depending on size of syringe, Accuracy: Within +/-1 to 2%, Alarm: occlusion, near empty, end, system malfunction, syringe disengage, reminder, low battery, Power source: 240 V, 50Hz and power cord with round 3 pin top. Built in battery, Mounting: Pole or bed frame. | Set | √ | √ | √ | √
| Patient warming system with drape warming blankets : Fast warm up, Temperature range Low 32', medium 38', and high 40' - 44'C, Multiple mounting options: cart, bedrail, IV pole or floor, Drape warming blankets: for upper and lower part of body (adult) & pediatric warming blanket ( set of 10 blankets each). AC power 220 - 240V, 50 Hz, Hose 5 ft. with nozzle clip. Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √ | √
| Blood/Fluid Warming Unit: The warming device should apply advanced heat technology to obtain the right fluid temperature.Should have a wide range of temperature for all routine clinical situations.The device should be able to warm blood/fluid at the rate of 125ml/minute to 300m/minute. Delivers normothermic fluids at flow rates from 15ml/min to 250ml/minute. Wide range of temperature-from 28'C to 42'C. Pole mounting and connects to standard electrical (220V - 240V), Easy maintenance and self calibrating, Accessories as required. Power supply: 220-240V. Power cord with round 3 pin top. | Set | √ | √ | √
| Patient monitor for OT: 15 -17 inch color TFT with maximum 8 waveforms. Standard configuration- NIBP, SPO2, 3/5 lead ECG, Pulse rate, Temperature, dual IBP, resp. & ETCO2 & sidestream, multi-gas/O2 measuring. Large font display mode.TOUCH SCREEN. Comprehensive arrhythmia review. Adjustable alarm sound. Battery and recorder. Power AC 220 - 240V. Patient range: adult, pediatric and neonate. Power cord with round 3 pin top. Accessories to be included: 10 numbers of CO2 sampling lines with water trap, 10 numbers of IBP transducers, Extra SPO2 probe & BP cuff for adult, child and neonate | Set | √ | √ | √
| Patient monitor for PACU & ICU: 10 - 12 inch color TFT display with 8 waveforms. Standard Configuration: NIBP, SPO2, 3/5 lead ECG, EtCO2, IBP, temprature and pulse rate, Arrhythmia review. Adjustable alarm sound. Patient range:adult, pediatric and neonate. TOUCH SCREEN. Power AC 220 - 240V. Battery and recorder. Power cord with round 3 pin top. Accessories to be included: 10 numbers of CO2 sampling lines with water trap,10 numbers of IBP transducers, Extra SPO2 probe & BP cuff for adult, child and neonate (with wall mount) | Set | √ | √
| Patient monitor for Wards : 8-10" color TFT display with 4 - 6 waveforms. Standard configuration; ECG, RESP. Dual-TEMP. NIBP, SPO2, Pulse rate. Adjustable alarm sound. Patient range; adult, pediatric and neonate. Power AC 220 - 240V. Battery and recorder. Accessories to be inlcuded :SPO2 probe & BP cuff for adult, child and neonate. | Set | √ | √ | √
| Vital Sign Monitor: Standard Configuration- NIBP and SPO2, adult and pediatric sensor probe. Power supply: AC 220-240V. Power cord with round 3 pin top | Set | √ | √ | √ | √
| Pulse Oximeter: SPO2 measurement range of 1% - 100%, Accuracy saturation (SpO2 =/- 1 SD) inbuilt rechargeable back up battery, High quality LCD graphical screen, Plethysmographic waveform or magnified vieww with pulse amplified indicator,visual indicators: Pulse search, audible alarm, silenced or off, artifact indicator, low battery and battery charging, sensor off, Alarm: audible and visual alarm for high and low saturation and pulse rate, audible and visual warning indicator for low battery, display saturation and pulse rate alarm limits. Standard accessories: Extension cable, sensor probe for adult, pediatric and neonate, power cord, operator manual.Variable pitch beep tone, Selection of adult and neonate alarm. Power supply: AC 220-240V, 50 Hz. Power cord with round 3 pin top | Set | √ | √ | √
| GlideScope-Video Laryngoscopes: Features anti-fogging mechanism, integrated high resolution camera and color monitor & video output/recording features, Video laryngoscope with baton for 1. adult, 2. pediatric & 3. neonate size. Power AC 220V- 240V with power power cord of round three pin top, with inbuilt battery backup | Set | √ | √
| Digital Ultrasonic Diagnostic Imaging System and for nerve block (Laptop Type): High resolution colour monitor 10 - 15 inch, High resolutions imaging from 2 to 18 MHz, upto 2-4 hours of plug free imaging, Transducer button operation: high speed linear array 6-18 MHz, curved array 2- 6 MHZ. Hockey stick transducer probe, frequency 5-10 MHz. Power supply: 220-240V. Power cord with round 3 pin top | Set | √
| Manual Jet Ventilator: with regulator and Gauge with jet ventilation catheter, 100 cm connecting tube with leur lock. | Set | √
| ICU Bed with mattress: Split safety sides with built-in patient and nurse control, Folding safety sides with control, Electrically operated back rest, knee-break and height adjustment, Electrically operated calf section adjustment, Electrically operated trendelenberg/reverse trendelenberg tilt, Auto-CPR button, Battery backup, Built-in bed extension with mattress retainer, Removal head and foot panels, Linked braking/stering system with four braking castors, Over all length and width 230 - 105 cm approximately, Power supply: 220-240V. Power cord with round 3 pin top | Piece | √ | √
| Arterial blood gas machine (ABG) : Advanced Analytics, Fast Stat Results, All Liquid Calibration Cartidges, Automates Quality Control, Automatic Quality Analysis, Automatic Data Storage and Reporting, Small Sample Volume, Easy, Safe Capillary or Syringe Use. Standard Measured Test, Calculated Results, Special Calculated Test, Measurement Range. including facility for calculation of anion gap and printer. Calibration: User-selectable, fully automatic two-point calibration, user-selectable single point calibration, manual calibration initiated at any time. Power: 220 - 240V, 50 Hz, Power cord of round 3 pin top. Data port. | Set | √ | √
| Patient Controlled Analgesia Machine: Programmable, work on a wide ranges of syringes/infusion pump type, comprehensive alarm packages, with various modes of infusion, bolus, continuous infusion, lock out, electronic security key, battery back facility, Power supply: 220-240V, 50 Hz, Power cord with round 3 pin top and suitable voltage stabilizer. | Piece | √
| Laryngoscope: Standard Macintosh blade- 4 blades Mac 1,2,3,4 with battery, handle type C. (stainless steel blades) with Fibreoptic light bulbs | Set | √ | √ | √ | √
| Laryngoscope :Standard Millers blade- MIL 1, 2, 3, 4 with battery, handle type C. (stainless steel blades) with Fibreoptic light bulbs | Set | √ | √ | √ | √
| Laryngoscope :Standard Millers blade- MIL 00, 0, 1 with battery, handle type AA. (stainless steel blades) with Fibreoptic light bulbs | Set | √ | √ | √ | √
| Macintosh FO blade : Flexible blade with integrated F.O light conductor, with bendable distal end, blade size adult,pediatric and neonate, stainless steel blade and battery handle type C. | Set | √
| Airway Management Stimulator :Head and Neck with Thyroid cartilage and bronchial tree fixed to base, exchangeable skin and thyroid cartilage, carina with ballons/bronchial tree | Set | √
| Manniquin: CPR Torso with light Indicator: Papable anatomical landmarks, Adult CPR training with optional child spring, Fully articulated head, neck and jaw, Palpable carotid pulse, Adult/child mode, Proper compression depth, proper ventilation volume, improper hand placement indicated by yellow, green and red light. Accessories; Mouth and nose pieces, child spring, disposable lower airway, light controller with batteries, disposable tracheal airway washable suit and carrying case.Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √
| CPR Torso Manniquin: Containing adult, child and neonate torso, mouth protection and carrying bag. | Set | √
| ECG Simulator with Arrhythmia | Piece | √
| Recovery Trolley: Light weight trolley, Swing type side rail protector, Provision for IV stand, Oxygen cylinder cage, Adjustable Back rest lifting 0 - 80 degrees, Adjustable height and facility to put in trendelenberg position. Frame made of stainless steel, Four 15 cm antistatic castor, lockable wheel, Foam mattress covered with synthetic leather. | Set | √ | √ | √
| Crash Trolley: Tubular framework, 6 removable bins, two lockable storage units with three drawers each, each cylinder carrier, stainless steel saline rod, mounted on 125 mm DB castor, two brake provision. | Set | √ | √ | √
| Peripheral Nerve Stimulator: Pulse amplitude adjustable from 2 to 100 mA. Modes: TWT, TET, TOF, DBS, PTC. Battery operated. | Set | √
| Cuff Pressure Gauge : 50 mm scale, release valve to allow sensitive adjustment of the desired cuff pressure, scale with cuff range from 20 to 30 cm H20. | Piece | √
| Patient transfer system: For patient transfer between trolley, OT table, ICU bed etc : Easy move or patslide, full body length, compatible with 200 kg patient | Piece | √ | √
| Wright's Respirometer: Tidal volume measurement with minute volume facility. Small inner dial records 0 - 100 liters, large outer dial measures 0 - 1 liter, Connector: inlet/outlet, 22 mm male BS/ISO taper. | Piece | √
| Sequential compression device (SCD) for prevention of DVT in ICU. Sleeves for the device: Knee length and thigh length ( Sizes: Small, Medium & Large) | Set | √
| Fluid warming cabinet for warming IV and irrigation fluids | Piece | √
| BiPAP machine: - portable- Autoset algorithym,- With humidifier, IPAP: 4-30 ppressure mmHg, EPAP: 4-30 preaaure mmHg,- With at least 3 modes: a)CPAP, b) Bipap, c) spontaneous mode | set | √
| Flexible bronchoscopy/fibreoptic bronchoscope: with xenon light source, forceps for bronchoscopic biopsy, channel for biopsy, and suction, biopsy channel diameter of 2.2mm, suction port, and cleaning brush, viewing angulation range of 130 /180o -video bronchoscope would be preferable | set | √
Consumables
| Oropharyngeal Airway : Size: 00 | Piece | √ | √ | √ | √
| Oropharyngeal Airway : Size: 0 | Piece | √ | √ | √ | √
| Oropharyngeal Airway : Size: 1 | Piece | √ | √ | √ | √
| Oropharyngeal Airway : Size: 2 | Piece | √ | √ | √ | √
| Oropharyngeal Airway : Size: 3 | Piece | √ | √ | √ | √
| Oropharyngeal Airway : Size: 4 | Piece | √ | √ | √ | √
| Sleeves for sequential compression device. Knee length | Piece | √ | √ | √ | √
| sleeves for sequential compression device. Thigh length | Piece | √ | √ | √ | √
Face mask
| Face mask transparent disposible : Size 00 (inflatable cushion) | Piece | √ | √ | √
| Face mask transparent disposible : Size 0 (Resuable inflatable cushion) | Piece | √ | √ | √
| Face mask transparent disposible : Size 1 (Resuable inflatable cushion) | Piece | √ | √ | √
| Face mask transparent disposible : Size 2 (Resuable inflatable cushion) | Piece | √ | √ | √
| Face mask transparent disposible : Size 3 (Resuable inflatable cushion) | Piece | √ | √ | √
| Face mask transparent disposible : Size 4 ( Resuable inflatable cushion) | Piece | √ | √ | √
| BiPAP masks: full face masks, soft cushion with proper seal, head strap, large and medium size compatible with BiPAP machine | Piece | √ | √ | √
| Capnography (etCO2 ) Sampling Line standard 200 tom400 cm line length, with leur lock male type at both end, single patient use, donot contain natural rubber latex, intended to be used in intubated, or face mask, adaptor dead space is <6.6cc | Piece | √ | √ | √
Endotracheal tube cuffed: Oral/Nasal:(O/N)
| Endotracheal tube : PVC, high volume,low pressure (Micro) cuff size I.D 2.5mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume,low pressure (Micro) cuff size I.D 3 mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume, low pressure (Micro) cuff size- 3.5mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume, low pressure (Micro) cuff size- 4 mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume, low pressure cuff size- 4.5mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume, low pressure cuff size- 5 mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume, low pressure cuff size- 5.5 mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume low pressure cuff - size I.D 6mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume low pressure cuff - size I.D 6.5mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume low pressure cuff - size I.D 7mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume low pressure cuff - size I.D 7.5mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume low pressure cuff - size I.D 8mm (O/N) | Piece | √ | √ | √
| Endotracheal tube :PVC, high volume low pressure cuff - size I.D 8.5mm (O/N) | Piece | √ | √ | √
Endotracheal tube non-cuffed: Oral/Nasal
| Endotracheal tube: PVC, non-cuffed-size I.D 2mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 2.5mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 3mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 3.5mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 4mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 4.5mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 5mm | Piece | √ | √ | √
| Endotracheal tube: PVC, non-cuffed-size I.D 5.5mm | Piece | √ | √ | √
RAE tube Cuffed PVC Oral
| RAE tube : PVC, micro cuffed oral- size ID 3mm | Piece | √ | √
| RAE tube : PVC, micro cuffed oral- size ID 3.5mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 4mm | Piece | √ | √
| RAE tube : PVC, cuffed oral - ID 4.5mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 5mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 5.5mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 6mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 6.5mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 7mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 7.5mm | Piece | √ | √
| RAE tube : PVC, cuffed oral- size ID 8mm | Piece | √ | √
RAE Tube Uncuffed Oral
| RAE tube : PVC, non-cuffed oral- size ID 2.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed oral- size ID 3mm | Piece | √ | √
| RAE tube : PVC, non-cuffed oral- size ID 3.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed oral- size ID 4mm | Piece | √ | √
| RAE tube : PVC, non-cuffed oral- size ID 4.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed oral- size ID 5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed oral- size ID 5.5mm | Piece | √ | √
RAE tube Cuffed PVC Nasal
| RAE tube : PVC, cuffed nasal- size ID 3.5mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 4mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 4.5mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 5mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 5.5mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 6mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 6.5mm | Piece | √ | √
| RAE tube : PVC, cuffed nasal- size ID 7mm | Piece | √ | √
RAE Tube PVC, Uncuffed Nasal
| RAE tube : PVC, non-cuffed nasal- size ID 2.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 3mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 3.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal-size ID 4mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 4.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 5.5mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 6mm | Piece | √ | √
| RAE tube : PVC, non-cuffed nasal- size ID 7mm | Piece | √ | √
Double Lumen Endobronchial Tube
| Double Lumen Endobronchial tube set : Left sided- size- 26 Fr | Piece | √
| Double Lumen Endobronchial tube set : Right sided- size 26 Fr | Piece | √
| Double Lumen Endobronchial tube set : Left sided- size 35 Fr | Piece | √
| Double Lumen Endobronchial tube set : Right sided- size 35 Fr | Piece | √
| Double Lumen Endobronchial tube set : Left sided size 37 Fr | Piece | √
| Double Lumen Endobronchial tube set : Right sidedsize 37 Fr | Piece | √
| Double Lumen Endobronchial tube set : Left sided - size 39 Fr | Piece | √
| Double Lumen Endobronchial tube set : Right sided- size 39 Fr | Piece | √
| Double Lumen Endobronchial tube set : Left sided - size 28 Fr | Piece | √
Flexo-Metallic Endotracheal Tube cuffed
| Flexo-Metalic Endotracheal Tube : Cuffed- size - I.D 5mm | Piece | √ | √
| Flexo-Metalic Endotracheal Tube : Cuffed- size - I.D 5.5mm | Piece | √ | √
| Flexo-Metalic Endotracheal Tube : Cuffed- size - I.D 6mm | Piece | √ | √
| Flexo-Metallic Endotracheal Tube: Cuffed- size -I.D 6.5 mm | Piece | √ | √
| Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 7mm | Piece | √ | √
| Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 7.5mm | Piece | √ | √
| Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 8mm | Piece | √ | √
| Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 8.5mm | Piece | √ | √
Microlaryngeal Tube
| Microlaryngeal tube : Size - I.D 5mm | Piece | √
| Microlaryngeal tube : Size - I.D 6mm | Piece | √
Laryngeal Mask Airway: (LMA)
| Laryngeal mask airway : Size 1, reusable, (curved/classic) with apparture bar | Piece | √ | √
| Laryngeal mask airway : Size 1.5, reusable (curved/classic) with apparture bar | Piece | √ | √
| Laryngeal mask airway : Size 2, reusable (curved/classic) with apparture bar | Piece | √ | √
| Laryngeal mask airway : Size 2.5, reusable (curved/classic) with apparture bar | Piece | √ | √
| Laryngeal mask airway : Size 3, reusable (curved/classic) with apparture bar | Piece | √ | √
| Laryngeal mask airway : Size 4, reusable (curved/classic) with apparture bar | Piece | √ | √
| Laryngeal mask airway : Size 5, reusable (curved/classic) with apparture bar | Piece | √ | √
Laryngeal Mask Airway, Flexible, Reinforced Tube
| Laryngeal mask airway : Flexible, reinforced tube, latex free, size 2 | Piece | √
| Laryngeal mask airway : Flexible, reinforced tube, latex free, size 2.5 | Piece | √
| Laryngeal mask airway : Flexible, reinforced tube, latex free, size 3 | Piece | √
| Laryngeal mask airway : Flexible, reinforced tube, latex free, size 4 | Piece | √
Laryngeal Mask Airway: Proseal reusable
| Laryngeal Mask Airway: Proseal reusable: Size 2 | Piece | √
| Laryngeal Mask Airway: Proseal reusable: Size 2.5 | Piece | √
| Laryngeal Mask Airway: Proseal reusable: Size 3 | Piece | √
| Laryngeal Mask Airway: Proseal reusable: Size 4 | Piece | √
| Laryngeal Mask Airway: Proseal reusable: Size 5 | Piece | √
| Laryngeal Mask Airway: Proseal reusable: Size 1.5 | Piece | √
Intubating LMA Airway
| Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 6, latex free. | Piece | √
| Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 6.5, latex free. | Piece | √
| Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 7, latex free. | Piece | √
| Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 7.5, latex free. | Piece | √
| Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 6, Latex free | Piece | √
| Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 6.5, Latex free | Piece | √
| Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 7, Latex free | Piece | √
| Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 7.5, Latex free | Piece | √
Laryngeal Tube
| Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone) - reusable size 2 | Piece | √ | √
| Laryngeal Tube with Drain Tube: Color code syringe and one biteblock (silicone)- reusable size 2.5 | Piece | √ | √
| Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone)- Reusable Size 3 | Piece | √ | √
| Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone)- Reusable Size 4 | Piece | √ | √
| Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone)- Reusable size 5 | Piece | √ | √
| Esophageal/Tracheal Double Lumen Airway (Combitube): Size: 37 ID | Piece | √ | √
Emergency Cricothyrotomy set
| Emergency Cricothyrotomy set : ID 2mm (child) | Piece | √ | √
| Emergency Cricothyrotomy set : ID 4mm (adult) | Piece | √ | √
| Percutaneous dilatational tracheostomy set (Adult) | Piece | √ | √
| Percutaneous dilatational tracheostomy set (pediatric) | Piece | √ | √
| Percutaneous dilatational tracheostomy set (Infant) | Piece | √ | √
Stylets
| Stylet : Size: OD Fr 6 (small) | Piece | √ | √ | √
| Stylet : Size: OD Fr 10 (medium) | Piece | √ | √ | √
| Stylet : Size: OD 14 (large) | Piece | √ | √ | √
| Gum elastic bougie (adult) | Piece | √ | √ | √
| Gum elastic bougie (child) | Piece | √ | √ | √
| Gum elastic bougie (neonate) | Piece | √ | √ | √
| Tube exchanger with markings and port for oxygenation adult | Piece | √ | √ | √
| Tube exchanger with markings and port for oxygenation pediatric | Piece | √ | √ | √
Resuscitation Bag
| Reusable Silicone Resuscitation Bag : Adult 2000 ML, with 40 cm H2O pressure relief, Air Cushion Face Mask # 4, Reservoir Bag 2600 ML with Valve, 2 m O2 tubing. | Set | √ | √ | √ | √
| Reusable Silicone Resuscitation Bag : Child 500 ML, with 40 cm H2O pressure relief, Air Cushion Face Mask # 2, Reservoir bag 1800 ML with Valve, 2m 02 Tubing | Set | √ | √ | √ | √
| Reusable Silicone Resuscitation Bag : Infant 250 ML with 40 cm H20 pressure relief, Mask round # 1, Reservoir Bag 600 ML with valve, 2 m 02 tubing | Set | √ | √ | √ | √
Oxygen Face Mask
| Oxygen face mask : With adjustable nose clip, elastic headstrap with reservoir bag with 2 meter tubing (Adult) | Set | √ | √ | √ | √
| Oxygen face mask : With adjustable nose clip, elastic headstrap with reservoir bag- pediatric, with 2 meter tubing (Pediatric) | Set | √ | √ | √ | √
| Oxygen nasal prong : With 2 meter tubing (Adult) | Set | √ | √ | √
| Oxygen nasal prong : With 2 meter tubing (Pediatric) | Set | √ | √ | √ | √
Neubilizer Mask Kit
| Adult nebulizer mask kit : Facemask with adjustable nose clip, head strap, nebulizer chamber with 2 meter tubing. | Set | √ | √ | √ | √
| Pediatric nebulizer mask kit : Face mask with adjustable nose clip, head strap, nebulizer chamber with 2 meter tubing. | Set | √ | √ | √ | √
| 22 mm adult breathing circuit : Corrugated, flexible and lightweight tubing, 150 cm, 22F-22F ventilator connections, parallel Y piece ports, 90'angled elbow with capno port, 22M/15F patient connections, straight connector 22M-22M/19F, sterile (PP/PE material), should have extension tube for reservior bag connection | Set | √ | √ | √
| 15 mm pediatric breathing circuit : 100 cm, 22F-22F ventilator connections, Y piece port, 90' angled elbow with capno port, 22M/15F patient connections, sterile. (PP/PE material), should have extension tube for reservior bag connection | Set | √ | √ | √
| Jackson REES breathing system : Corrugated, flexible, lightweight, fresh gas corrugated tube 180 cm, vertical gas inlet, 22M/15F angled patient connections with capno port, 15M ventilator connection, tube 50 cm with 22M connector, 1/2 lt open latex free bag, not sterile. | Set | √ | √ | √
| Mapleson D - Bain's circuit : 22 mm Adult PE Anesthesia circuit, 198 cm, 22M/15F patient connection, 90' angled elbow at patient side, vertical gas inlet, APL valve with 2 liter bag, tube length 200 - 250 cm. Fresh gas inlet at patient end. | Set | √ | √
| T-piece for ICU and recovery | Piece | √ | √
Breathing System Bacterial/Viral Filter for Ventilator
| Breathing System Filter for anesthesia circuit (Adult) : Filter material: Pleated hydrophobic membrane. Liquid borne bacterial/viral removal efficiency: 100%. Air borne bacterial/viral removal efficiency: >99.99%. Connections: Patient side 15 mm ID and 22 mm OD, breathing system side 15 mm OD 22 mm ID. Filter Volume and filter weight: Low filter volume and low weight. Construction: Transparent, non-conductive material. Port: CO2 monitoring port. (for OT only) | Piece | √ | √ | √
Breathing System Filter for anesthesia circuit (pediatric) : Filter material: Pleated hydrophobic membrane. Liquid borne bacterial/viral removal efficiency: 100%. Air borne bacterial/viral removal efficiency: >99.99%. Connections: Patient side 15 mm ID and 22 mm OD, breathing system side 15 mm OD 22 mm ID. Filter Volume and filter weight: Low filter volume and low weight. Construction: Transparent, non-conductive material. Port: CO2 monitoring port. (for OT only)
| HMEF Filter for icu - Adult : 100% bacterial Viral Efficiency, HME Membrane with CO2 port, 22M/15F ISO Connectors. (single use) ( for ICU only) | Piece | √ | √
| HMEF Filter for ICU - Pediatric : 100% bacterial Viral Efficiency, HME Membrane with CO2 port, 22M/15F ISO Connectors.(single use) ( for ICU only) | Piece | √ | √
| HMEF Filter - Neonate: 100% Bacterial Viral Efficiency, HME Membrane with CO2 port, 22M/15F connector.(single use) ( for ICU only) | Piece | √ | √
| 10 mm neonate breathing circuit: Disposable, lightweight,100cm, angled Y piece with Luer lock (capnoport).22mm connectors | Set | √ | √
| 22 mm adult breathing circuit :MRI COMPATIBLE Corrugated, flexible and lightweight tubing,300cm, 22F-22F ventilator connections, parallel Y piece ports, 90'angled elbow with capno port, 22M/15F patient connections, straight connector 22M-22M/19F, not sterile (PP/PE m | Set | √ | √
| 15 mm pediatric breathing circuit : MRI COMPATIBLE 300 cm, 22F-22F ventilator connections, Y piece port, 90' angled elbow with capno port, 22M/15F patient connections, not sterile. (PP/PE material) | Set | √ | √
| 10 mm neonate breathing circuit: MRI COMPATIBLE Disposable, lightweight,300cm, angled Y piece with Luer lock (capnoport) .22mm connectors | Set | √ | √
Smooth bore PVC Circuit (ICU Ventilator)
| Smoothbore PVC circuits (with accessories) : 22 mm adult PVC ICU circuit, standard Y-Piece with port, 22M/15F patient connections, inspiratory line 160 cm with adult water trap, expiratory line 160 cm with adult water trap, extra limb 100 cm, 22Flex-22Flex ventilator connections, sterile. | Set | √ | √
| Smoothbore PVC circuits (with accessories) : 22 mm Pediatric PVC ICU circuit, standard Y-Piece with port, 22M/15F patient connections, inspiratory line 160 cm with pediatric water trap, expiratory line 160 cm with pediatric water trap, extra limb 80 cm, 22Flex-22Flex ventilator connections, sterile. | Set | √ | √
Breathing Bag/Reservoir Bag
| Breathing Bags (Latex Free) : 0.5L breathing bag, latex free, 22F connector, not sterile | Piece | √ | √ | √
| Breathing Bags (Latex Free) : 1L breathing bag, latex free, 22F connector, not sterile | Piece | √ | √ | √
| Breathing Bags (Latex Free) : 2L breathing bag, latex free, 22F connector, not sterile | Piece | √ | √ | √
Pressure Infusion Cuff
| Pressure Infusion Cuff: With hand inflator and manometer fixed to the cuff: wrap around, reusable. 500 ml | Set | √ | √ | √
| Pressure Infusion Cuff: With hand inflator and manometer fixed to the cuff: wrap around, reusable. 1000ml | Set | √ | √
| CPAP Respiration Set : CPAP Mask, Harness, PEEP Valve adjustable from 5 - 20cmH2O, and adapter for 22mm tube connection | Set | √
Close Suction System, Replaceable Catheters, Endotracheal Length, MDI Port
| Size: 6Fr/Ch | Piece | √ | √
| Size: 8Fr/Ch | Piece | √ | √
| Size: 10Fr/Ch | Piece | √ | √
| Size: 12Fr/Ch | Piece | √ | √
| Size: 14Fr/Ch | Piece | √ | √
| Size: 16Fr/Ch | Piece | √
| Size: 18Fr/Ch | Piece | √
Biteblock
| Biteblock : Autoclavable- size 1 length 2.5 cm | Piece | √ | √
| Biteblock : Autoclavable- size 3 length 3 cm | Piece | √ | √
| Biteblock : Autoclavable- size 5 length 3.5 cm | Piece | √ | √
Central Venous Catheter
| CVP Line cava fix 45 cm | Set | √ | √
| CVP Line cava fix 70 cm | Set | √ | √
| CVP manometer : With connector (sterile) | Set | √ | √
| Triple lumen central venous cather sets : 7F, 18G x 150 - 200 mm catheter length. | Set | √ | √
| Triple lumen central venous cather sets : 5F, 20G x 100 - 150 mm catheter length. | Set | √ | √
| Double lumen central venous cather sets : 4F, 20G, 100- 150mm catheter length. | Set | √ | √
| Single lumen Central venous catheter : 3F, 22G, 100-150 mm catheter lendth | Set | √ | √
| Single lumen Central venous catheter : 2.5F, 22G, 60 - 100 mm catheter lendth | Set | √ | √
Intra-arterial catheter
| Arterial Catheter: Radial artery set 2.5F, lenght 60 mm | Set | √
| Arterial Catheter: Radial artery set 3F, lenght 60 mm | Set | √
| Closed loop arterial blood sampling disposable transducer set for the invasive Monitoring of Haemodynamics Pressure. | Set | √
Extension Tubes/Lines(sterile)
| Extension tube: With 3-way stop cock with leur lock: Length 50 cm | Piece | √ | √
| Extension tube: With 3-way stopcock with leur lock, length 100 cm (low volume) | Piece | √ | √
| Extension tube: With 3-way stopcock with leur lock, length 150 cm | Piece | √ | √
| Extension tube: Length 100-150mm Male and Female Leur Lock (Low volume for syringe pump) | Piece | √ | √
Epidural Needles and Catheters
| Epidural needles : 18 G with epidural needle 16 G, with accessories,(loss of air resistence syringe10ml, epidural catheter, sterile pack set with plastic tray | Set | √ | √
| Epidural needles: 18G, with Loss of resistence to air syringe 10ml , With plastic tray in sterile pack | Piece | √ | √
| Epidural needles: 20G, sterile pack. | Piece | √ | √
| Pediatric epidural set 18/20G, 50 mm with 5mm marking | Piece | √ | √
| Neonate/Infant epidural set 20/24G 35 or 50 mm with 5mm marking | Piece | √ | √
Spinal Needles
| Spinal needle (pencil point) : Disposable, size - 25 G-sterile ,with introducer needle | Piece | √ | √ | √
| Spinal needle (pencil point) : Disposable, size - 26 G- sterile ,with introducer needle | Piece | √ | √ | √
| Spinal needle (pencil point) : Disposable, size - 27 G- sterile with introducer needle | Piece | √ | √ | √
| Spinal needle (Quinckes): size 22, length 15 cm- sterile (for nerve block) | Piece | √
| Spinal needle (Quinckes) : Disposable, size 22 G- sterile | Piece | √ | √ | √
| Spinal needle (Quinckes) : Disposable, size 23 G- sterile | Piece | √ | √ | √
| Spinal needle (Quinckes) : Disposable, size 25 G- sterile | Piece | √ | √ | √
| Spinal needle (Quinckes) : Disposable, size 27G- sterile | Piece | √ | √ | √
| 90' Angled Connector : Dith 22M/15F patient connection, 15M ending with capno port with cap. | Piece | √ | √ | √
| Sodalime : Granule size 4 - 8 mesh or granules 1/4 - 1/8 inch in diameter, PH indicator turing from pink to white when exhausted. | Packet | √ | √ | √
Accessories for Medical Cylinders
| Regulator with Flow Meter & Humidifier Chamber for Oxygen cylinder : B/D types of Cylinders | Piece | √ | √ | √ | √
| Regulator for Oxygen cylinder : D types of Cylinders for Anesthesia | Piece | √ | √ | √
| Regulator for Nitrous oxide cylinder : D type | Piece | √ | √ | √
| Key and Wrench for Nitrous oxide cylinder: B & D-Type | Piece | √ | √ | √
| Key and Wrench for Oxygen cylinder : B & D-Type | Piece | √ | √ | √ | √
| Oxygen key: For A type of oxygen cylinders | Piece | √ | √ | √
| High pressure oxygen tubing : For extended connection from Oxygen cylinder to Anesthesia Machine. | Piece | √ | √ | √
| High pressure nitrous oxide tubing : For extended connection from Nitrous oxide cylinder to Anesthesia Machine. | Piece | √ | √ | √
Forceps
| Forceps MaGill - Adult | Piece | √ | √ | √
| Forceps MaGill - Pediatric | Piece | √ | √ | √
Consumables for EPOC - Hand held (Portable)ABG (Arterial Blood Gas) Analyzer
| Cartridge (Test Card) | 50 test/pkt | √
| Control Fluid (Ampoules) | 10pcs/pkt | √
| Epoc ABG report printing paper roll | Roll | √
Medical Gases (Gas Refilling)
| Hydrogen Refilling-D Type | Per Cylinder | √ | √ | √
| Carbon Dioxide Refilling- B Type | Per Cylinder | √ | √ | √
| Carbon Dioxide Refilling- D Type | Per Cylinder | √
| Oxygen Gas Refilling- AA Type | Per Cylinder | √ | √ | √
| Oxygen Gas Refilling- A Type | Per Cylinder | √ | √ | √
| Oxygen Gas Refilling- B Type | Per Cylinder | √ | √ | √ | √
| Oxygen Gas Refilling- C Type | Per Cylinder | √ | √ | √
| Oxygen Gas Refilling- D Type | Per Cylinder | √ | √ | √
| Nitrogen Gas Refilling- D Type | Per Cylinder | √ | √
| Nitrous Oxide Gas Refilling- D Type | Per Cylinder | √ | √ | √
| Nitrous Oxide Gas Refilling- A Type | Per Cylinder | √ | √ | √
| Nitrous Oxide Gas Refilling- B Type | Per Cylinder | √ | √ | √
| IV drip set compatible with infusion pump | Piece | √ | √
| BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 5cm (find out he length) | Piece | √ | √
| BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 8cm (find out he length) | Piece | √ | √
| BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 10cm | Piece | √ | √
| Hospital Emergency Crash cart: 1. ABS trolley body 2. Plastic-steel columns 3. Adjustable height IV pole attached to unit 4. Removable Dust basket 5. Utility container 6. File bag 7. Defibrillator shelf, power outlet & hooks 8. Sliding side shelf 9. Stainless steel guard rail 10.Concave top covered transparent soft plastic glass 11.Centralized lock 12.Two small & Three big drawer with adjustable partition which can be well organized inside the each drawer. 13.Drawers repositioned eaily 14.Luxurious noiseless castersinner with partitions can be well organized 15.ABS basket & drawer frame can be repositoned easily 16.Luxuriouos casters wheels which enables easy trolleying / movement, with two with brakes 17. CPR board 18. Oxygen tank holder 29. Needle disposal bottle | set | √
Surfactant connector
| Surfactant connector attaches to ET 3/3.5 with side port | Piece | √ | √
Mucosal atomizer device
| Laryngo-tracheal mucosal atomizer device without syringe | Piece | √ | √
Pressure monitoring lines
Non-expanding PMO line 200cm with luer lock
Connectors
Male to male luer lock connector
Male luer lock to female connector
Disposable straight connector 22F-22F
Disposable straight connector 22M-22F/15F
Disposable straight connector 22M/15F-22M/15F
| High pressure nitrous oxide tubing : For extended connection from Nitrous oxide cylinder to Anesthesia Machine. | Piece | √ | √ | √
Forceps
| Forceps MaGill - Adult | Piece | √ | √ | √
| Forceps MaGill - Pediatric | Piece | √ | √ | √
Consumables for EPOC - Hand held (Portable) ABG (Arterial Blood Gas) Analyzer
| Cartridge (Test Card) | 50 test/pkt | √
| Control Fluid (Ampoules) | 10pcs/pkt | √
Medical Gases (Gas Refilling)
| Hydrogen Refilling-D Type | Per Cylinder | √ | √ | √
| Carbon Dioxide Refilling- B Type | Per Cylinder | √ | √ | √
| Carbon Dioxide Refilling- D Type | Per Cylinder | √
| Oxygen Gas Refilling- AA Type | Per Cylinder | √ | √ | √
| Oxygen Gas Refilling- A Type | Per Cylinder | √ | √ | √
| Oxygen Gas Refilling- B Type | Per Cylinder | √ | √ | √ | √ | √
| Oxygen Gas Refilling- C Type | Per Cylinder | √ | √ | √
| Oxygen Gas Refilling- D Type | Per Cylinder | √ | √ | √
| Nitrogen Gas Refilling- D Type | Per Cylinder | √ | √
| Nitrous Oxide Gas Refilling- D Type | Per Cylinder | √ | √ | √
| Nitrous Oxide Gas Refilling- A Type | Per Cylinder | √ | √ | √
| Nitrous Oxide Gas Refilling- B Type | Per Cylinder | √ | √ | √
| IV drip set compatible with infusion pump | Piece | √ | √
| BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 5cm (find out he length) | Piece | √ | √
| BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 5cm (find out he length) | Piece | √ | √
| BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 8cm | Piece | √ | √
| Hospital Emergency Crash cart: 1. ABS trolley body 2. Plastic-steel columns 3. Adjustable height IV pole attached to unit 4. Removable Dust basket 5. Utility container 6. File bag 7. Defibrillator shelf, power outlet & hooks 8. Sliding side shelf 9. Stainless steel guard rail 10.Concave top covered transparent soft plastic glass 11.Centralized lock 12.Two small & Three big drawer with adjustable partition which can be well organized inside the each drawer. 13.Drawers repositioned eaily 14.Luxurious noiseless castersinner with partitions can be well organized 15.ABS basket & drawer frame can be repositoned easily 16.Luxuriouos casters wheels which enables easy trolleying / movement, with two with brakes 17. CPR board 18. Oxygen tank holder 29. Needle disposal bottle | set | √
Surfactant connector
| Surfactant connector attaches to ET 3/3.5 with side port | Piece | √ | √
Mucosal atomizer device
| Laryngo-tracheal mucosal atomizer device without syringe | Piece | √ | √
Pressure monitoring lines
Non-expanding PMO line 200cm with luer lock
Connectors
Male to male luer lock connector
Male luer lock to female connector
Disposable straight connector 22F-22F
Disposable straight connector 22M-22F/15F
Disposable straight connector 22M/15F-22M/15F
Chemotherapy
Equipments
| Negative pressure Isolator for chemotherapy drug mixing; Two glove ports, tw+A1o transfer chambers, exhaust spigot for connection to dedicated exhaust system (Ducted type). Should also include remote fans for the exhaust system. [OR Four glove ports, two transfer chambers, exhaust spigot for connection to dedicated exhaust system (Ducted type). Should also include remote fans for the exhaust system]. Additional features: Negative pressure for operation and safety. EU GMP Grade A unidirectional airflow work area for aseptic work and product protection. Transfer chambers with electromagnetic timed interlocks and high air change rates for total suppression of airborne contamination during material transfers in and out, and for rapid evaporation of disinfecting agents. Digital display for isolator functions and transfer chamber operation. Automated pressure decay leak test for indicating results and monitoring trends. PLC controls for easy set up, operation and maintenance. Good access to all potentially contaminated areas to facilitate cleaning. Safe to change main (primary) filter. 316L stainless steel polished internally and externally. 13 amp single splash proof socket. Service taps. Glove sleeve leak test disc with digital manometer. Remote fan for the exhaust system. Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √
Consumables
| Gown with cuffs and back closure; impermeable fabric | Set | √
| Gloves; for protection | Set | √
| Goggles | Pair | √
| Shoe cover; disposable | Pair | √
| Mask; disposable | Piece | √
| Absorbable towel | Piece | √
| Mask (N-95) | Piece | √
| Disposable sponges | Piece | √
| Container; punture and leak proof | Piece | √
| Container of desiccant powder or granules to absorb wet contents | Piece | √
| Disposable dust pan | Piece | √
| Plastic scraper | Piece | √
| Plastic bag (Ziplock ) | Piece | √
| Butterfly needle (scalp vein set), Size 23 | Piece | √
| Butterfly needle (scalp vein set); Size 24 | Piece | √
| Butterfly needle (scalp vein set); Size 25 | Piece | √
| Water for injection - plastic 50ml | Piece | √
| Water for injection-plastic 100ml | Piece
| Syringe with Luer lock cap 2ml (disposable) | Piece | √
| Syringe with Luer lock cap 5ml (disposable) | Piece
| Syringe with Luer lock cap 10ml (disposable) | Piece
| Syringe with Luer lock cap 20ml (disposable) | Piece
| Syringe with Luer lock cap 50ml (disposable) | Piece
CSSD
Equipments
| Autoclave horizontal (Large): Chamber size: 500mm×5000mm×950 (W×L×H). Electrical power: 18Kw, 3 phase, 440 V 50Hz AC supply. Working pressure: 1.2 to 2.2 kg/Sq. cm ( upto 21 psi), at 121 degree Celsius. Inner chamber: stainless steel ( thickness 5mm to 10 mm) with argon welding. | Piece | √ | √
| Autoclave vertical 400x600mm High grade stainless steel SS 304 or equivalent, triple walled, internal joint should be argon arc welded; Internal chamber Size (75 to 80 Ltrs): 400mm (diameter)X600mm (depth) ;supplied with two suitable stackable stainless steel drum for sterilization,waterlevel gauge ( Min and Max), temperature and pressure indicator, safety valve,air release valve, water inlet valve and drainage valve, sillicon gasket, automatic cutoff protection for low water, stainless steel drum stand, . Fully automatic programmable controller with digital display: working pressure 15 to 20 psi , working temperature 115 deg to 132 deg. Power supply: 220-240 AC, 50Hz. Power cord of round 3 pin top. Operating enviromental condition: Humidity > 80%, temperature -10 deg to + 40 deg . Should conform to IEC 60601/ ISI or higher certification (CE, FDA,ISO ). | Set | √ | √ | √ | √
| Autoclave double door: 304 / 316 / 316L Stainless Steel autoclave horizontal (Large) Double door,Size 1100 X 950 X 950 or 1300x1000x900mm (working pressure: 5psi-20psi; working temperature: 115 deg to 132 deg ; shape:square; working pressure :2.5kgf/cm sq.g(35psi) standard controls through PLC (Programmable Logic Controller): multioperating valve, quick vaccum drying apparatus,accidental vaccum Breaker, safety valve, air release valve, water inlet and drainage valve, pressure display, chamber condensate line, pressure & compound gauge). Accessories: low water protection device, Automatic pressure controsl switch (Pressure control switch, Automatic water cut-off device, digital temperature indicator, timer with alarm system, vacuum pump Power supply : 3 phase AC (440V), 50Hz, Power cord of round 3 pin top. Operational enviromental operating condition: Humidity > 80%, temperature -10 deg to + 40 deg . Should conform to IEC 60601/ ISI or higher certification (CE, FDA,ISO ). | piece | √
| Programmable sterilizer with high vacuum process including low temperature cycle for heat sensitive goods, leak test, Bowie/Dick tape test, environmental operating condition:-10±40° C, RH>80%,altitude>2000m,ISO/TUV/CE certified, Specify loading trays, Volume>1200lts,Chamber depth>1500mm.3 pin round type plug top | Piece | √ | √ |  |  | CSSD
| Automatic Instrument Washing Machine: (10 DIN): Power supply: 220V-230V, Power cord of round 3 pin round type plug top. | Piece | √ | √ |  |  | Where there is functional OT .
| Vaccum, Ultrasonic and plasma cleaning and drying equipment: Power supply: 220V-230V, Power cord of round 3 pin round type plug top. | Piece | √ | √ |  |  | CSSD
| Automatic instrument washer, disinfector and dryer (15DIN trays): Camber volume 500litres ,Basket volume 350litres . Reduction of instrument cleaning cycle up to 30 minutes | Piece | √ | √ |  |  | CSSD
| Perforated stainless steel instrument tray with lid: size 23" x 15" x 5.5" | Piece | √ | √ |  |  | CSSD
| Instrument tray: size 16.5" x 8.5" x 2" | Piece | √ |  |  |  | CSSD
Reagents for Plasma sterrad 100s
| Sterrad 100S sterilliant cassete: code 10113 | 5pcs/ pkt | √ | √
| Chemical Indiactor strip: Code 14100 | 1000pcs/pk | √ | √
| Chemical Indiactor Tape: Code 14202 | 6pcs/roll/pkt | √ | √
| Cycle Sure Biological Indiactor: Code14324 | 4pcs/roll/box | √ | √
| Tyvek Roll 4" x 70m: Code 12410 | 6pcs/roll/pkt | √ | √
| Tyvek Roll 6' x 70m: Code 12415 | 4pcs/roll/pkt | √ | √
| Tyvek Roll 10'x 70m: Code 12425 | 4pcs/roll/pkt | √ | √
| Tyvek Roll 14' x 70m: Code 12435 | 2pcs/roll/pkt | √ | √
| Ribbon catridge red/blk | piece | √ | √
| Casette collection box | piece | √ | √
| Vapourizer plate | piece | √ | √
| Printing paper | roll | √ | √
Consumables
| Autoclave indicating tape: 20mmx50meter | roll | √ | √ | √ | √
| Checmical indicators strip for stream sterlization | strips per pkt | √ | √ | √
| total length 8 inch brush length 2.52 ich - Med 100.46 | piece | √ | √ | √
| total length 7.2 inch Nylon - Med 100.43 | piece | √ | √ | √
| laparoscopic brush total length 7 inch Med 100.33 | piece | √
| Flexible bone reamer cleaning brush total length 30 inch brush 3mm - Med 100.57 | piece | √
| Suction pipe cleaner, total length 12 inch brush 1/8 inch Med 100.34 | piece | √
| Suction pipe cleaner, total length 12 inch brush 1/4 inch Med 100.35 | piece | √
| Utility elbow length gloves size 6.5 | piece | √
| Utility elbow length gloves size 7 | piece | √
| Enzymatic Detergent | litre | √
| Hinge Free Instrunment lubicant spray | 250 ml | √
| sterlization Perforated instrument tray 12/10 inch | piece | √
| sterlization Perforated instrument tray 15/17 inch | piece | √
| 5 shelf shelving/wire mesh with wheels | piece | √
Dental Department
General
Equipments
| Bone Rongeurs Luer 14.5cm | Piece | √
| Browne lingual flap retractor | Piece | √
| Chin Retractor Obwegeser 16.5cm | Piece | √
| Condyle retractor - thin - AK's new | Piece | √
| Cleft lip surgery - set | Set | √
| Cleft plate surgery set | Set | √
| Crown wire twisting forceps: 16cms/6 1/4 inch | Piece | √
| Dingmans mouth gag | Piece | √
| Dissecting Forceps, Adson-brown 12cm | Piece | √
| Dissecting Forceps, Gillies 15.5cm | Piece | √
| Dissecting Forceps, Potts-Smith 16cm | Piece | √
| Dissecting Forceps, Potts-Smith 18cm | Piece | √
| Dissecting Forceps, Potts-Smith 23cm | Piece | √
| Finger protector: langenbeck | Piece | √
| Ferguson jaw opener | Piece | √
| Hammers length 230mm, weight 350g. | Piece | √
| Heister jaw opener - SS | Piece | √
| Jaw Splint - Small | Piece | √
| Jaw Splint - Large | Piece | √
| Lip & Check Retractor - Cawood - Minnesota | Piece | √
| Lip & Check Retractor - Oringer | Piece | √
| Mandibular Modular Fixation System with following contents: |  | √
| Mandible Trauma Module Set (2.0mm) consisting of: | Set | √
| Mandible Trauma Module Set (2.4mm) consisting of: | Set | √
| Maxillofacial Modular Fixation System with following contents: |  | √
| Module 1.5mm | Piece | √
| Module 2.0mm | Piece | √
| Maxillary awl Obwegeser 15cm | Piece | √
| Maxillary awl Obwegeser 17.5cm | Piece | √
| Mouth gag Mckesson - adult | Piece | √
| Mouth gag Mckesson - Child | Piece | √
| Needle holder 16cm | Piece | √
| Needle holder ,Mathieu,170mm | Piece | √
| Needle holder 23cm | Piece | √
| Obwegeser Retractor 24cm | Piece | √
| Osteotomy Chisel 10mm | Piece | √
| Osteotomy Chisel 15mm | Piece | √
| Periosteal Elevator Joseph | Piece | √
| Ramus retractor | Piece | √
| Retractor Adson | Piece | √
| Rowe's Mandible bone holding forceps | Piece | √
| Rowe's Maxilla disimpaction clamp - left | Piece | √
| Rowe's Maxilla disimpaction clamp - right | Piece | √
| Rowe's zygomatic elevator | Piece | √
| Scissor Iris 11.5cm/ 4.5 inch | Piece | √
| Scissor Metzenbaum - fino 18cm/ 7 inch | Piece | √
| Scissor Tenotomy 115mm | Piece | √
| Sinus Lift curettes 15cm | Piece | √
| Sinus Lift curettes 18.5cm | Piece | √
| Surgical Skin hooks | Piece | √
| Third Molar Retractor Obwegeser 24cm | Piece | √
| Tissue retractor Austin | Piece | √
| Tongue depressor(wider) | Piece | √
| Wire cutter,Straight | Piece | √
| Wire ligature pliers for circum mandibuler | Piece | √
| Wire ligature pliers for circum zygomatic | Piece | √
| Wire pusher Rowe | Piece | √
| Zygoma elevator Gillies 20cm | Piece | √
| Zygomatic bone awl curved/ zygomatic arch awl | Piece | √
| Zygomatic bone hook with T handle | Piece | √
| Maxillofacial Modular Fixation System with following contents: a) Obwegeser Retractor 24cm; b)Third Molar Retractor Obwegeser 24cm | Set | √
| Elcomed, contra angle, straight HP, motor cable and motor holder, surgical hand piece (W&H) | set
| Tungsten carbide burs for surgical- Round (2.35mmX30-40mm) | piece
| Tungsten carbide burs for surgical-Taper criss-cut (2.35mmX30-40mm) | piece
| Tungsten carbide burs for surgical-Cylindrical (2.35mmx30-40mm) | piece
| Drill Bits with stops, for Stryker coupling, length- 44.5mm, diameter-1.55mm | piece
| Drilling depth 6mm | piece
| Drilling depth 8mm | piece
| Drilling depth 12mm | piece
| Bur block stainless steel (sterilizable) for surgical | Piece
| Surgical drill-Stryker COMMAND2 50K DRILL. | piece
| Cord for Surgical Drill Stryker command2 –FISHER 104 | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 6 mm for stryker coupling j-latch type | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 8 mm for stryker coupling j-latch type | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 10 mm for stryker coupling j-latch type | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 12 mm for stryker coupling j-latch type | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 6mm,pure titanium , | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 8mm,pure titanium , | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 10mm,pure titanium , | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 12mm,pure titanium , | piece
| 2.4 mm Emergency Screw length 6mm,pure titanium | piece
| 2.4 mm Emergency Screw length 8mm,pure titanium | piece
| 2.4 mm Emergency Screw length 10mm,pure titanium | piece
| 2.4 mm Emergency Screw length 12mm,pure titanium | piece
| Mandible Mini Plates 2.0,centre space, 4holes,length 28mm,pure titanium | piece
| Mandible Mini Plates 2.0,centre space, 6holes,length 40mm,pure titanium | piece
| Mandible Mini Plates 2.0,centre space, 20holes,length 119mm,pure titanium | piece
| Chin plates 2.0,double bend,5 holes,pure titanium , | piece
| Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece
| Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece
| 1.0mm MF Cortex Screw,Self tapping, Length 4mm,pure titanium, | piece
| 1.0mm MF Cortex Screw,Self tapping, Length 5mm,pure titanium, | piece
| 1.0mm MF Cortex Screw,Self tapping, Length 6mm,pure titanium, | piece
| Double Drill Sleeve 2.4/1.8 | piece
| Depth Gauge for 1.5 to 2.0 screws | piece
| IMF Screw Set with stainless steel implants with lid, 1. Wire twister | piece
IMF Screw Set with stainless steel implants with lid, 1. Wire Cutter
| Dental Amalgamator for silver alloy powder & mercury,Power 220-240V, power cord with round 3 pin top. | no | √ | √ | √
| Acrylizer bath electrical with temperature regulator. Temperature range 0- 100°C, Voltage 220-240,power cord with round 3 pin top. | No | √ | √ | √
| Dental Air Compressor, oil free, 1-4 hp, pressure upto 8 bar, noise free, capicity 25 - 30L, Voltage 220-240V, power cord with round 3 pin top. | No | √ | √ | √
| Dental Chair, High end complete set( with low and high speed Hand pieces, Scaler, Air Compressor unit, suction machine, dentsit and assisat chairs stools, three way syringe, light cure unit, x ray viewer,oral camera with screen) Voltage 220-240v, 50Hz. Suitable Voltage stabilizer. All power cords of round 3 pin plug tops | Set | √
| Dental Chair Standard, Complete set (with low and high speed hand pieces, air compressor unit, x ray viewer,two stools, three way syringe, light cure unit), Suitable Voltage stabilizer having high/low cut off with delay timer, Power cord with round 3 pin tops. | Set | √ | √ | √
| Water distilllation flask, Capacity 5 Litres, Voltage supply 220-240v, 50Hz. Power cord with round 3 pin top | No | √ | √ | √
| Dental iontophoresis device, direct current/pulse,0.4mV current,12V, Input volatge supply 220-240, three pin round tops, | No | √ | √
| Light curing unit (cordless) with light guide, Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
| Lathe polishing machine with builtin burchuck & eye shield double arm. Power supply 220-240v, 50Hz. Power cord with round 3 pin top,Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. | No | √ | √ | √
| Lathe machine Trimmer , nonstop 220v builtin burchuck, inbuilt light-system, suction chamber & eye shield. Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
| Micromotar, rpm 1000-50,000, Electrical, 220V-240V, table top, with CHP,50Hz & foot control, with oil,Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
| Micromotor , rpm 1000-50,000, with Cable for SHP attachment.Power Supply 220-230v, 50Hz, Power cord with round 3 pin top. |  | √ | √ | √
| OPG macine with Cephalometric CDR program,Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
| Vacuum alginate/plaster/stone mixer, Table top with inbuilt Vacuum pump/gauze digital display along mixing 170 & 560 cc Bowls/covers/blades. Power supply 220-240v, 50Hz. Power cord with round 3 pin top. |  | √ | √ | √
| Dental Vibrator, Power supply 220-240v, 50Hz. Power cord with round 3 pin top. | No | √ | √ | √
| Absorbable gelatin haemostatic sponge | 10pcs/ pkt | √ | √ | √
| Dispensing gun, periodontal dressing material. | Set | √ | √ | √
| Disposable suction tube, standard zsize,dental | 100pcs/ pkt | √ | √
| Dycal cement, Calcium Hydroxide Composition, (Base 13 gm, catalyst 11 gm with mixing pad) | Packet/set of Base 13 gm and catalyst 11 gm | √ | √
| Cocoa-butter tube for glass ionomer | Set | √ | √ | √
| Glass Ionomer, Light cured material set | Set | √ | √
| Miracle Mix, restorative | Set | √ | √ | √
| Mortar and pestle, Glass/porcelain | Set | √ | √ | √
| Silver alloy powder | 30gm | √ | √ | √
| Iodoform paste, Paedo root canal sealant | Set | √ | √ | √
| X-ray lead apron with thyroid collar- flexible |  | √ | √ | √
| ZnOE Impression (base-catalyst)paste tube | Set | √ | √ | √
| Forceps Cow-Horn, Lower | Piece | √ | √ | √
| Forceps Cow-Horn, Upper | Piece | √ | √ | √
| A.R.T. (autraumatic filling) Hand instrument set of 8 with head lamp. | Set | √ | √ | √
| CHP Hand piece for Micromotor, compatible with exisiting machine | Piece | √ | √ | √
| Spot welder (Dental use) | Piece | √ | √ | √
| Whitening Accelerator (UV light) for bleaching | Piece | √ | √
| Scissor Crown curved, | Piece | √ | √ | √
| Universal Wire Cutter Bandage Scissor, Size 120 mm, SS with golden handle. | piece | √ | √ | √
| Osada Model XL-S30 (Oral surgery),Oral surgery and implant, Handpiece system,( All electric High torque/low speed), Micromotor handpiece and built-in Peristaltic irrigation system. Micromotor with AMC: LVSA,Hand piece:Speed range 1600-60,000 rpm. SH-31 SP Double speed Twist Chuck Handpiece,LVSA micromotor/MCC 8-8 ft coiled cord | Piece | √
| Portable All-in-One PEDO-30W,Power Console ( Floor Unit ), Variable Speed ( 800-30,000 rpm), Standard Podiatric Handpiece: SH-29 S/ LVS /MC S8 on HHR Rubber,Output Speed Range 550-11,000 rpm | Set | √
Consumables
| Absorbable gelatin haemostatic sponge | 10pcs/ pkt | √ | √ | √
| Acrylic powder, veined, heat cure with liquid Cold mold seal, separating sheet | Set | √ | √ | √
| Acrylic, powder, clear, heat cure with liquid, cold mold seal, separating sheet | Set | √ | √ | √
| Acrylic, powder, pink, heat cure with liquid, cold mold seal, separating sheet | Set | √ | √ | √
| Acrylic, powder, pink, cold cure with liquid | Set | √ | √ | √
| Acrylic, powder, clear, cold cure with liquid | Set | √ | √ | √
| Acrylic trimmer, SHP,steel, round – small, std size | Piece | √ | √ | √
| Acrylic trimmer, steel, SHP round – medium, std size | Piece | √ | √ | √
| Acrylic trimmer, steel, SHP,round – large, std size | Piece | √ | √ | √
| Acrylic trimmer, steel, SHPflame shape – small, std size | Piece | √ | √ | √
| Acrylic trimmer, steel, SHP flame shape – medium, | Piece | √ | √ | √
| Acrylic trimmer, steel, SHP flame shape – large, std size | Piece | √ | √ | √
| Articulating paper, standard-40/60µm | Packet of 12 books of 12 sheets | √ | √ | √
| Band for matrix retainer, Ivory No.1 | 10pcs/ pkt | √ | √ | √
| Band for matrix retainer, Tofflemire contra-angle, junior | 10pcs/ pkt | √ | √ | √
| Band for matrix retainer, Tofflemire contra-angle, adult | 10pcs/ pkt | √ | √ | √
| Bristle brush, mounted, CHP | Piece | √ | √ | √
| Bur holder,revolving 7” diameter 4” high(400x343) | Piece | √ | √ | √
| Bur cleaning brush | Piece | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 006 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 008 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 010 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 012 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 014 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 016 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 018 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 006 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 008 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 010 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 012 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 014 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Tapered fissure-long Shank, size ISO 016 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Tapered fissure-Long Shank, size ISO 018 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 006 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 008 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 010 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 012 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 014 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 016 | 5pcs/pkt | √ | √ | √
| Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 018 | 5pcs/pkt | √ | √ | √
| Carborundum discs | 100disc/ pkt | √ | √ | √
| Cellophane/ Mylar strips | 100dics/ pkt | √ | √ | √
| Chamois leather for amalgam | Piece | √ | √ | √
| Clove oil | 500 ml | √ | √ | √
| Dispensing gun, periodontal dressing material. | Set | √ | √ | √
| Cold mold seal of Sodium alginate solution (450 ml) | Bottle | √ | √ | √
| Composite light cure material, complete set | Packet/ set | √ | √ | √
| Composite finishing kit | Packet/ kit | √ | √ | √
| Cotton buff wheel, Large for Lathe,4½” | Piece | √ | √ | √
| Cotton buff wheel, Small for SHP, mounted,3” | Piece | √ | √ | √
| Crown preparation set |  | √ | √ | √
| Dappen glass, different colours | Piece | √ | √ | √
| Dental wedges wooden, readymade, assorted of 4 different colours. | 400pcs pkt | √ | √ | √
| Dental floss | Piece | √ | √ | √
| Dental Plaque disclosing Solution/ gel | 45ml | √ | √ | √
| Denture adhesive powder | 15gm | √ | √ | √
| Denture cleaning powder | 15gm | √ | √ | √
| Disposable suction tube, standard zsize,dental | 100pcs/ pkt | √ | √ | √
| Dry socket paste | set | √ | √ | √
| Dycal cement, Calcium Hydroxide Composition, (Base 13 gm, catalyst 11 gm with mixing pad) | Packet/set | √ | √ | √
| Endo: Broaches, barbed, 15-40, 21 mm length | Pkt | √ | √ | √
| Endo: Broaches, barbed, 15-40, 25mm length | Pkt | √ | √ | √
| Endo: EDTA gel for canal irrigation, 5gm | Piece | √ | √ | √
| Endo: Finger spreader assorted, 15-40, | Pkt | √ | √ | √
| Endo: Forma cresol solution, | 10 ml | √ | √ | √
| Endo: GP points, assorted pkt, size 15-40 | Pkt | √ | √ | √
| Endo: GP points, assorted pkt, size 45-80 | Pkt | √ | √ | √
| Endo: Hedstroem file NiTi, 25mm, assorted pkt, size 45 - 80 | Pkt | √ | √ | √
| Endo: Hedstroem file NiTi, 25mm, assorted pkt, size 15 - 40 | Pkt | √ | √ | √
| Endo: Hedstroem file NiTi, 21mm, assorted pkt, size 15 - 40 | Pkt | √ | √ | √
| Endo: K- file, NiTi, 25mm, assorted pkt, size 15- 40 | Pkt | √ | √ | √
| Endo: K- file, NiTi, 21mm, assorted pkt, size 15- 40 | Pkt | √ | √ | √
| Endo: K- file, NiTi, 21mm, assorted pkt, size 15- 40 | Pkt | √ | √ | √
| Endo: K - file, NiTi, 25mm, assorted pkt, size 45 - 80 | Pkt | √ | √ | √
| Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 15 - 40 | Pkt | √ | √ | √
| Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 45 -80 | Pkt | √ | √ | √
| Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 15 - 40, Lock grip assorted | Pkt | √ | √ | √
| Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 45 -80, lock grip assorted | Pkt | √ | √ | √
| Endo: Lentulo - Spiral, NiTi, 25mm, Lock grip, assorted, | Pkt | √ | √ | √
| Endo: Lentulo-spiral, assorted,19mm lock grip | Pkt | √ | √
| Endo: Orifice opener | Pkt | √ | √ | √
| Endo: Paper point, assorted packet, size 15-40 | Pkt | √ | √ | √
| Endo: Paper point, assorted packet, size 45 - 80 | Pkt | √ | √ | √
| Endo: Probe, periodontal (graded) | Pkt | √ | √ | √
| Endo: RCT obturation paste, Endomethasone N, Powder 14 gm and liquid 10 ml | Set | √ | √ | √
| Endo: RCT plugger | Piece | √ | √ | √
| Endo: Root canal filling solvent | 50ml | √ | √ | √
| Endo: Sodium hypochlorite 3% | 450ml | √ | √ | √
| Endo: pulp devitilizer(Paraformaldehyde, Phenol, Lidocaine Base) | Ml | √ | √ | √
| Endo RCT sealer Endomethasone powder and liquid | Set | √ | √ | √
| Endobox, aluminuim, with stand, small (104x44x40mm) | Piece | √ | √ | √
| Endobox, aluminuim, with stand, large (204x105x54mm) | Piece | √ | √ | √
| Endo: RCT measurement kit | Piece | √ | √ | √
| Erich ARCH bar, 12 inch | Piece | √ | √
| Eye-shield Plastic transparent, spectacle type | Piece | √ | √ | √
| Felt cone, large (1”x4”) | Piece | √ | √ | √
| Felt cone, medium (1”x3”) | Piece | √ | √ | √
| Glass Ionomer, Fuji VII,VII, IX | Set | √ | √ | √
| Cocoa-butter tube for glass ionomer | Set | √ | √ | √
| Glass Ionomer, Light cured material set | Set | √ | √
| Glass Ionomer, Pit & Fissure Sealant material | Set | √ | √ | √
| Miracle Mix, restorative | Set | √ | √ | √
| Glass mixing slab, 6"x 3" x 1/2" | Piece | √ | √ | √
| Gum paint, Gumex | 60ml | √ | √ | √
| Handpiece lubricating oil spray | Cane | √ | √ | √
| Impression Compound | 5pcs/pkt | √ | √ | √
| Impression material Sodium-Alginate, normal setting | 450gm | √ | √ | √
| Lathe stone wheel – small (2”x¼”) | Piece | √ | √ | √
| Lathe stone wheel - medium(3”x¼”) | Piece | √ | √ | √
| Lathe stone wheel - large(4”x¼”) | Piece | √ | √ | √
| Mandrel for Sand paper SHP |  | √ | √ | √
| Mandril for disc, CHP, L.G. | Piece | √ | √ | √
| Mandril for disc, SHP | Piece | √ | √ | √
| Mercury liquid, 30gm (100%) | Piece | √ | √ | √
| Mortar and pestle, Glass/porcelain | Set | √ | √ | √
| Mould guide, rubber, lower | Piece | √ | √ | √
| Mould guide, rubber, upper | Piece | √ | √ | √
| Mouth gag , rubber adult | Piece | √ | √ | √
| Plaster of Paris | 50kg/ drum | √ | √ | √
| Plastic disc, soflex polishing | 100dics/ pkt | √ | √ | √
| Plastic face shield | 10pcs/ pkt | √ | √ | √
| Polishing rubber cups, CHP, L.G. | Piece | √ | √ | √
| Polishing brush for lathe | Piece | √ | √ | √
| Polysulfide impression Paste, Base/Catalyst set/kit, Regular: Kerr | Set | √ | √ | √
| Polysulfide rubber base tray Adhesive: Kerr | Set | √ | √
| Prophylaxis paste, Dose Assorted | Piece | √ | √ | √
| Provi-paste filling material | 20gm | √ | √ | √
| Pumice powder | 5kg/pkt | √ | √ | √
| Rubber bowl – large (12.5cm) | Piece | √ | √ | √
| Rubber bowl – small (9.5cm) | Piece | √ | √ | √
| Sand paper disc, assorted | 100dics/ pkt | √ | √ | √
| Sharpening stone, Arkansas, double side | Piece | √ | √ | √
| Shellac Base Plate - Lower | 50pcs/ pkt | √ | √ | √
| Shellac Base Plate - Upper | 50pcs/ pkt | √ | √ | √
| Silver alloy powder | 30gm | √ | √ | √
| Stone mounted, SHP, cylindrical – Small,std size | Piece | √ | √ | √
| Stone mounted, SHP, cylindrical – Medium std size, | Piece | √ | √ | √
| Stone mounted, SHP, cylindrical – Large, std size | Piece | √ | √ | √
| Stone mounted, SHP, flame shaped – Small, std size | Piece | √ | √ | √
| Stone mounted, SHP, flame shaped – Medium, std size | Piece | √ | √ | √
| Stone mounted, SHP, flame shaped – Large, std size | Piece | √ | √ | √
| Stone mounted, SHP, round – Small, std size | Piece | √ | √ | √
| Stone mounted, SHP, round – Medium, std size | Piece | √ | √ | √
| Stone mounted, SHP, round – Large, std size | Piece | √ | √ | √
| Stone mounted, SHP, tapered – Small, std size | Piece | √ | √ | √
| Stone mounted, SHP, tapered – Medium, std size | Piece | √ | √ | √
| Stone mounted, SHP, tapered – Large, std size | Piece | √ | √ | √
| Stone plaster | 1kg/pkt | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, lower anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, lower anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, lower anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, upper & lower posterior small set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, upper & lower posterior medium set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, upper & lower posterior large set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, set of 28 upper & lower full set small | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, set of 28 upper & lower full set medium | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, set of 28 upper & lower full set large | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, upper anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, upper anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 24, upper anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, lower anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, lower anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, lower anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, upper & lower posterior small set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, upper & lower posterior medium set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, upper & lower posterior large set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, set of 28 upper & lower full set small | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, set of 28 upper & lower full set medium | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, set of 28 upper & lower full set large | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, upper anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, upper anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 25, upper anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, lower anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, lower anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, lower anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, upper & lower posterior small set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, upper & lower posterior medium set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, upper & lower posterior large set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, set of 28 upper & lower full set small | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, set of 28 upper & lower full set medium | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, set of 28 upper & lower full set large | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, upper anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, upper anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 26, upper anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, lower anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, lower anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, lower anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, upper & lower posterior small set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, upper & lowerposterior medium set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, upper & lower posterior large set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, set of 28 upper & lower full set small | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, set of 28 upper & lower full set medium | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, set of 28 upper & lower full set large | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, upper anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, upper anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 27, upper anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, lower anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, lower anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, lower anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, upper & lower posterior small set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, upper & lower posterior medium set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, upper & lower posterior large set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, upper anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, upper anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 28, upper anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, lower anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, lower anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, lower anterior large, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper & lower posterior small set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper & lower posterior medium set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper & lower posterior large set of 16 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper & lower posterior large set of 17 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper anterior small, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper anterior medium, set of 6 | Set | √ | √ | √
| Teeth, acrylic cross linked, shade - 29, upper anterior large, set of 6 | Set | √ | √ | √
| Iodoform paste, Paedo root canal sealant | Set | √ | √ | √
| Temporary SS crown standard set | Set | √ | √ | √
| Topical fluoride gel (1 ltr Jar/Bottle) | 1ltr/bottle | √ | √ | √
| Tracing green stick | 10pcs/pkt | √ | √ | √
| Wax modeling | 10pcs/pkt | √ | √ | √
| Wax sticky | 10pcs/pkt | √ | √ | √
| Whitening powder | 1kg/pkt | √ | √ | √
| Wire spool, SS, G -19 | Piece | √ | √ | √
| Wire spool, SS, G -20 | Piece | √ | √ | √
| Wire spool, SS, G -21 | Piece | √ | √ | √
| Wire spool, SS, G -22 | Piece | √ | √ | √
| Wire spool, SS, G -23 | Piece | √ | √ | √
| Wire spool, SS, G -27 | Piece | √ | √ | √
| Wire spool, SS, G -28 | Piece | √ | √ | √
| X- ray Developer (powder) | 1kg/pkt | √ | √ | √
| X- ray Fixer (Powder) | 1kg/pkt | √ | √ | √
| X- ray film, Occlusal, 57x76mm (2 ¼” x 3”) No.4,E-speed, minimum 2yrs of expiry date | 50pcs/ pkt | √ | √ | √
| X- ray film, IOPA, No. 0,E-speed, minimum 2yrs of expiry date | 50pcs/ pkt | √ | √ | √
| X-ray film, IOPA, No. 2, E-speed, minimum 2yrs of expiry date | 50pcs/ pkt | √ | √ | √
| X-ray film Bitewing holder | 50pcs/ pkt | √ | √ | √
| X-ray film lateral ceph, Size - 8x10 inch' | 50pcs/ pkt | √ | √ | √
| X-ray lead apron with thyroid collar- flexible |  | √ | √ | √
| X-ray OPG film, 6x12" | 50pcs/ pkt | √ | √ | √
| Zinc Phosphate cement , podwder-liquid set | Set | √ | √ | √
| ZnOE cement, powder & liquid set | Set | √ | √ | √
| ZnOE Impression (base-catalyst)paste tube | Set | √ | √ | √
| Needdle,1”,27G (0.40x21 mm) | 100pcs/box | √ | √ | √
| Cerclage Wire, 0.6mm dia, without eye, length 175mm | Piece | √
| Surgical Cutting Burs, SS, HM 141 204 040 | Piece | √
| Surgical Cutting Burs, SS, 165 205 023 | Piece | √
| Tungsten carbide SS, HM 141 205 040 | Piece | √
| Tungsten carbide SS, HM 141 205 050 | Piece | √
| Tungsten carbide, HM 33T 104 016 | Piece | √
| Tungsten carbide, HM 151 316 016 | Piece | √
| Tungsten carbide, HM 152 315 014 | Piece | √
| Tungsten carbide, HM 162 104 016 | Piece | √
Instrunment
| Amalgam burnisher, double ended,(1mm,1.1mm) small | Piece | √ | √ | √
| Amalgam burnisher, double ended, medium(1.8mm, 3mm) | Piece | √ | √ | √
| Amalgam burnisher, double ended(1.6mm,2.4mm) | Piece | √ | √ | √
| Amalgam carver, diamond shape | Piece | √ | √ | √
| Amalgam carver, double ended, ½ Hollenbeck | Piece | √ | √ | √
| Amalgam condenser, serrated double ended (1.2mm,1.5mm) small | Piece | √ | √ | √
| Amalgam carrier, plastic, conventional, straight | Piece | √ | √ | √
| Amalgam carrier, plastic, curved 45˚ | Piece | √ | √ | √
| Bone Chisel, osteotome 4" pointed curved, triangular | Piece | √ | √ | √
| Bone chisel, osteotome 4" semi circular, curved | Piece | √ | √ | √
| Bone files, Small, 6 inch, double ended | Piece | √ | √ | √
| Bone nibbler dental, Small, 6 | Piece | √ | √ | √
| Bone Rongeur | Piece | √ | √ | √
| Cement plastic filling, medium, double ended | Piece | √ | √ | √
| Cement plastic filling, small, double ended | Piece | √ | √ | √
| Cement plugger condenser, medium, double ended | Piece | √ | √ | √
| Cement plugger condenser, small, double ended | Piece | √ | √ | √
| Cheek retractor blade | Piece | √ | √ | √
| Coupland, Apexo elevator, set of 2 | Piece | √ | √ | √
| Dental Cartridge Syringe 1.8ml | Piece | √ | √ | √
| Dental flask, brass | Piece | √ | √ | √
| Elevator, Bristows | Piece | √ | √ | √
| Elevator, cryer's - left | Piece | √ | √ | √
| Elevator, cryer's - right | Piece | √ | √ | √
| Elevator, Periosteal, SS - Small | Piece | √ | √ | √
| Elevator, Periosteal, SS - large | Piece | √ | √ | √
| Elevator, Rowes | Piece | √ | √ | √
| Elevators, Coupland, straight - Small | Piece | √ | √ | √
| Elevators, Coupland, straight - Medium | Piece | √ | √ | √
| Elevators, Coupland, straight - Large | Piece | √ | √ | √
| Elevators, warwick jams, set of 3 | Set | √ | √ | √
| Endo-spoon excavator, double ended - small | Piece | √ | √ | √
| Endo-Spoon excavator, double ended - large | Piece | √ | √ | √
| File, volcanite round | Piece | √ | √
| File, volcanite, 1/2 round | Piece | √ | √ | √
| Forcep - bone cutting – lower | Piece | √ | √ | √
| Forcep - bone cutting – upper | Piece | √ | √ | √
| Forceps - dissecting, tooth, 4 inch, Adson | Piece | √ | √ | √
| Forceps - plain serrated tissue, 4 inch | Piece | √ | √ | √
| Forceps - toothed tissue 4 inch | Piece | √ | √ | √
| Forceps Cow-Horn, Lower | Piece | √ | √ | √
| Forceps Cow-Horn, Upper | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Lower Anterior, Klein pattern | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Lower Molar, Klein pattern | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Lower Root, Klein pattern | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Upper Anterior, Klein pattern | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Upper Molar, Klein pattern | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Upper Root, Klein pattern | Piece | √ | √ | √
| Forceps, Extraction, Deciduous Teeth: Upper Universal, Klein pattern | Piece | √ | √ | √
| Forceps-extraction, permanent teeth Lower Anterior | Piece | √ | √ | √
| Forceps-extraction, permanent teeth lower molar | Piece | √ | √ | √
| Forceps-extraction, permanent teeth lower premolar | Piece | √ | √ | √
| Forceps-extraction, permanent teeth lower root | Piece | √ | √ | √
| Forceps-extraction, permanent teeth Lower Third Molar | Piece | √ | √ | √
| Forceps-extraction, permanent teeth Upper Anterior | Piece | √ | √ | √
| Forceps-extraction, permanent teeth upper molar, left | Piece | √ | √ | √
| Forceps-extraction, permanent teeth upper molar, right | Piece | √ | √ | √
| Forceps-extraction, permanent teeth Upper premolar | Piece | √ | √ | √
| Forceps-extraction, permanent teeth Upper root | Piece | √ | √ | √
| Forceps-extraction, permanent teeth Upper Third Molar | Piece | √ | √ | √
| A.R.T. (autraumatic filling) Hand instrument set of 8 with head lamp. | Set | √ | √ | √
| Heavy wire cutter | Piece | √ | √ | √
| Lecrons carver | Piece | √ | √ | √
| Mallet surgical, for nylon headed | Piece | √ | √ | √
| Mallet prosthetic SS | Piece | √ | √ | √
| Mirror, mouth Top, sterilizable, no.5, adult | Piece | √ | √ | √
| Mirror, handle sterilizable, no. 5 | Piece | √ | √ | √
| Plaster cutter plier type | Piece | √ | √ | √
| Plaster knife | Piece | √ | √ | √
| Plaster spatula | Piece | √ | √ | √
| Pliers Young s universal | Piece | √ | √ | √
| Pliers, Adams medium, | Piece | √ | √ | √
| Pliers, Clasp bending | Piece | √ | √ | √
| Pliers, loop forming | Piece | √ | √ | √
| Probe, double ended | Piece | √ | √ | √
| RCT Plugger | Piece | √ | √ | √
| Scaler, cingulum | Piece | √ | √ | √
| Scaler, double ended inter-dental | Piece | √ | √ | √
| Scaler, Jaquette | Piece | √ | √ | √
| Scaler, sickle - Small | Piece | √ | √ | √
| Scaler, sickle - Large | Piece | √ | √ | √
| Scaler, universal | Piece | √ | √ | √
| Crown Bandage Scissor, Saw edge, Curved, size 110 m, SS with golden handle. | Piece | √ | √ | √
| Scissors Surgical straight, sharp pointed 5 " | Piece | √ | √ | √
| Scissors Surgical straight, sharp pointed 6 " | Piece | √ | √ | √
| Scissors Surgical curved, sharp pointed 5 " | Piece | √ | √ | √
| Scissors Surgical curved, sharp pointed 6 " | Piece | √ | √ | √
| Spatula, cement, SS | Piece | √ | √ | √
| Spirit lamp, steel/brass, | 250 ml | √ | √ | √
| Sub - gingival Scaler complete set | Piece | √ | √ | √
| Tweezer dental, fine tip | Piece | √ | √ | √
| Wax knife double ended | Piece | √ | √ | √
| Wax spatula double ended | Piece | √ | √ | √
| Articulator Hanau H2 with 6 pairs of mounting rings | Piece | √ | √ | √
| Articulator, free plane, mean Value | Piece | √ | √ | √
| Articulator, hinge | Piece | √ | √ | √
| Chip blower with nozzle | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, lower, size – 0 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, lower, size – 1 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, lower, size – 2 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, lower, size – 3 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, lower, size – 4 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, upper, size – 0 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, upper, size – 1 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, upper, size – 2 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, upper, size – 3 | Piece | √ | √ | √
| Impression tray, Edentulous, Non preforated, upper, size – 4 | Piece | √ | √ | √
| Impression tray Dentulous, perforated lower size – 0 | Piece | √ | √ | √
| Impression tray Dentulous, perforated lower size – 1 | Piece | √ | √ | √
| Impression tray Dentulous, perforated lower size – 2 | Piece | √ | √ | √
| Impression tray Dentulous, perforated lower size – 3 | Piece | √ | √ | √
| Impression tray Dentulous, perforated lower size – 4 | Piece | √ | √ | √
| Impression tray Dentulous, perforated, upper size – 0 | Piece | √ | √ | √
| Impression tray Dentulous, perforated, upper size – 1 | Piece | √ | √ | √
| Impression tray Dentulous, perforated, upper size – 2 | Piece | √ | √ | √
| Impression tray Dentulous, perforated, upper size – 3 | Piece | √ | √ | √
| Impression tray Dentulous, perforated, upper size – 4 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, lower size – 1 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, lower size – 2 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, lower size – 3 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, lower size – 4 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, lower size – 5 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, upper size – 1 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, upper size – 2 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, upper size – 3 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, upper size – 4 | Piece | √ | √ | √
| Impression tray Dentulous, non-perforated, rimlock, upper size – 5 | Piece | √ | √ | √
| Endodontic HP apex locator/triauto | Piece | √ | √ | √
| Flask clamp | Piece | √ | √ | √
| Flask, table clamp | Piece | √ | √ | √
| Gingivectomy set, complete | Set | √ | √ | √
| Gum Scissor Iris straight 14cm | Piece | √ | √ | √
| Hot plate(metal) | Piece | √ | √ | √
| Matrix retainer ivory no. 1 | Piece | √ | √ | √
| Matrix retainer toflemire no. 8 | Piece | √ | √ | √
| Spot welder (Dental use) | Piece | √ | √ | √
| Whitening Accelerator (UV light) for bleaching | Piece | √ | √
| Bending Pliers with beak, for plates up to 2.7 | Piece | √
| Bending-Cutting Pliers 1.0 to 2.0, right | Piece | √
| Bone Curettes Hemingway 18.5cm | Piece | √
| Bone Curettes Volkmann 170mm | Piece | √
| Bone cutting forcep 18cm | Piece | √
| Bone file 18cm | Piece | √
| Bone Holding forcep 13cm | Piece | √
| Bone Holding forcep 14cm | Piece | √
| Bone Holding forceps 18.5cm | Piece | √
| Bone Holding forceps, self-centering, length 191mm, softlock. | Piece | √
| Bone Reduction forcep with points, narrow 161mm softlock | Piece | √
| Bone Reduction forcep with points, wide 127mm softlock | Piece | √
| Bone Rongeur 12.5cm | Piece | √
| Bone Rongeur 16.5cm | Piece | √
| Bone Rongeurs Luer 14.5cm | Piece | √
| Browne lingual flap retractor | Piece | √
| Chin Retractor Obwegeser 16.5cm | Piece | √
| Condyle retractor - thin - AK's new | Piece | √
| Cleft lip surgery - set | Set | √
| Cleft plate surgery set | Set | √
| Crown wire twisting forceps: 16cms/6 1/4 inch | Piece | √
| Dingmans mouth gag | Piece | √
| Dissecting Forceps, Adson-brown 12cm | Piece | √
| Dissecting Forceps, Gillies 15.5cm | Piece | √
| Dissecting Forceps, Potts-Smith 16cm | Piece | √
| Dissecting Forceps, Potts-Smith 18cm | Piece | √
| Dissecting Forceps, Potts-Smith 23cm | Piece | √
| Finger protector: langenbeck | Piece | √
| Ferguson jaw opener | Piece | √
| Hammers length 230mm, weight 350g. | Piece | √
| Heister jaw opener - SS | Piece | √
| Jaw Splint - Small | Piece | √
| Jaw Splint - Large | Piece | √
| Lip & Check Retractor - Cawood - Minnesota | Piece | √
| Lip & Check Retractor - Oringer | Piece | √
| Mandibular Modular Fixation System with following contents: |  | √
| Mandible Trauma Module Set (2.0mm) consisting of: | Set | √
| Mandible Trauma Module Set (2.4mm) consisting of: | Set | √
| Maxillofacial Modular Fixation System with following contents: |  | √
| Module 1.5mm | Piece | √
| Module 2.0mm | Piece | √
| Maxillary awl Obwegeser 15cm | Piece | √
| Maxillary awl Obwegeser 17.5cm | Piece | √
| Mouth gag Mckesson - adult | Piece | √
| Mouth gag Mckesson - Child | Piece | √
| Needle holder 16cm | Piece | √
| Needle holder ,Mathieu,170mm | Piece | √
| Needle holder 23cm | Piece | √
| Obwegeser Retractor 24cm | Piece | √
| Osteotomy Chisel 10mm | Piece | √
| Osteotomy Chisel 15mm | Piece | √
| Periosteal Elevator Joseph | Piece | √
| Ramus retractor | Piece | √
| Retractor Adson | Piece | √
| Rowe's Mandible bone holding forceps | Piece | √
| Rowe's Maxilla disimpaction clamp - left | Piece | √
| Rowe's Maxilla disimpaction clamp - right | Piece | √
| Rowe's zygomatic elevator | Piece | √
| Scissor Iris 11.5cm/ 4.5 inch | Piece | √
| Scissor Metzenbaum - fino 18cm/ 7 inch | Piece | √
| Scissor Tenotomy 115mm | Piece | √
| Sinus Lift curettes 15cm | Piece | √
| Sinus Lift curettes 18.5cm | Piece | √
| Surgical Skin hooks | Piece | √
| Third Molar Retractor Obwegeser 24cm | Piece | √
| Tissue retractor Austin | Piece | √
| Tongue depressor(wider) | Piece | √
| Wire cutter,Straight | Piece | √
| Wire ligature pliers for circum mandibuler | Piece | √
| Wire ligature pliers for circum zygomatic | Piece | √
| Wire pusher Rowe | Piece | √
| Zygoma elevator Gillies 20cm | Piece | √
| Zygomatic bone awl curved/ zygomatic arch awl | Piece | √
| Zygomatic bone hook with T handle | Piece | √
| Maxillofacial Modular Fixation System with following contents: a) Obwegeser Retractor 24cm; b)Third Molar Retractor Obwegeser 24cm | Set | √
| Elcomed, contra angle, straight HP, motor cable and motor holder, surgical hand piece (W&H) | set | √
| Tungsten carbide burs for surgical- Round (2.35mmX30-40mm) | piece | √
| Tungsten carbide burs for surgical-Taper criss-cut (2.35mmX30-40mm) | piece | √
| Tungsten carbide burs for surgical-Cylindrical (2.35mmx30-40mm) | piece | √
| Drill Bits with stops, for Stryker coupling, length- 44.5mm, diameter-1.55mm | piece | √
| Drilling depth 6mm | piece | √
| Drilling depth 8mm | piece | √
| Drilling depth 12mm | piece | √
| Bur block stainless steel (sterilizable) for surgical | Piece | √
| Surgical drill-Stryker COMMAND2 50K DRILL. | piece
| Cord for Surgical Drill Stryker command2 –FISHER 104 | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 6 mm for stryker coupling j-latch type | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 8 mm for stryker coupling j-latch type | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 10 mm for stryker coupling j-latch type | piece
| Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 12 mm for stryker coupling j-latch type | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 6mm,pure titanium , | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 8mm,pure titanium , | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 10mm,pure titanium , | piece
| 2 mm Mandible Cortex Screw , Self- tapping,L 12mm,pure titanium , | piece
| 2.4 mm Emergency Screw length 6mm,pure titanium | piece
| 2.4 mm Emergency Screw length 8mm,pure titanium | piece
| 2.4 mm Emergency Screw length 10mm,pure titanium | piece
| 2.4 mm Emergency Screw length 12mm,pure titanium | piece
| Mandible Mini Plates 2.0,centre space, 4holes,length 28mm,pure titanium | piece
| Mandible Mini Plates 2.0,centre space, 6holes,length 40mm,pure titanium | piece
| Mandible Mini Plates 2.0,centre space, 20holes,length 119mm,pure titanium | piece
| Chin plates 2.0,double bend,5 holes,pure titanium , | piece
| Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece
| Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece
| 1.0mm MF Cortex Screw,Self tapping, Length 4mm,pure titanium, | piece
| 1.0mm MF Cortex Screw,Self tapping, Length 5mm,pure titanium, | piece
| 1.0mm MF Cortex Screw,Self tapping, Length 6mm,pure titanium, | piece
| Double Drill Sleeve 2.4/1.8 | piece
| Depth Gauge for 1.5 to 2.0 screws | piece
| IMF Screw Set with stainless steel implants with lid, 1. Wire twister | piece
IMF Screw Set with stainless steel implants with lid, 1. Wire Cutter
Accessories
| Ultra sonic scaler insert tips with keys, interdental for sirona scaler - HP, compatible with the (Specify Name; Brand & Model) Ultra sonic scaler. | Piece | √ | √ | √
| Ultra sonic scaler inserts tips with keys, subgingival for sirona scaler, HP, compatible with the (Specify Name; Brand & Model) Ultra sonic scaler. | Piece | √ | √ | √
| Turbine hand piece high speed, press button type, 4 holes, Standard | Piece | √ | √ | √
| Turbine Hand piece high speed, fibre optic, press button, standard head, 4 holes, T1/T2, compatible to the existing Dental chair |  | √ | √ | √
| Turbine hand piece high speed, fibre optic, press button, small head, 4 holes, T1/T2, compatible to the existing Dental chair |  | √ | √ | √
| Fibre optic quick coupling attachment/agent with bulb for hand piece,T1/T2 CHP compatible to the existing Dental Chair |  | √ | √ | √
| Fibre Optic high speed, CHP, Press button, standard head with quick coupling, compatible to the existing Dental chair |  | √ | √ | √
| Exposure / firing switch cat 5154307 with cable for wall mount KODEX, IOPA, X-Ray Machine model no. KODAK 2100, 230 V-240V, 50/60 Hz, 60 Kv, 7 AMP (compatible to the existing X ray machine) |  | √ | √ | √
| CHP Hand piece for Micromotor, compatible with exisiting machine | Piece | √ | √ | √
| SHP Hand piece for Micromotor, W&H | Piece | √ | √ | √
| Composite polishing paste | 2 Syringes/ Pkt | √
| Mineral trioxide aggregate (MTA) powder and liquid | Set | √
| Tooth forceps comella pattern (to extract wisdom tooth roots) | piece | √
| Universal Wire Cutter Bandage Scissor, Size 120 mm, SS with golden handle. | piece | √ | √ | √
Oral Surgery and Implants
Equipments
| Double Speed Contra Handpiece | Piece | √ | √
| Double Speed Twist Chuck Handpiece/ LVSA Micromotor/MCC8, 8 - Ft.: | Set | √ | √
| Low Speed High Torque Electric Handpiece Systems for Oral Surgery and Implants. | Piece | √ | √
Consumables
| Cerclage Wire, 0.6mm dia, without eye, length 175mm | Piece | √ | √
| Surgical Cutting Burs, SS, HM 141 204 040 | Piece | √ | √
| Surgical Cutting Burs, SS, 165 205 023 | Piece | √ | √
| Tungsten carbide SS, HM 141 205 040 | Piece | √ | √
| Tungsten carbide SS, HM 141 205 050 | Piece | √ | √
| Tungsten carbide, HM 33T 104 016 | Piece | √ | √
| Tungsten carbide, HM 151 316 016 | Piece | √ | √
| Tungsten carbide, HM 152 315 014 | Piece | √ | √
| Tungsten carbide, HM 162 104 016 | Piece | √ | √
Instrunment
| Bending Pliers with beak, for plates up to 2.7 | Piece | √ | √
| Bending-Cutting Pliers 1.0 to 2.0, right | Piece | √ | √
| Bone Curettes Hemingway 18.5cm | Piece | √ | √
| Bone Curettes Volkmann 170mm | Piece | √ | √
| Bone cutting forcep 18cm | Piece | √ | √
| Bone file 18cm | Piece | √ | √
| Bone Holding forcep 13cm | Piece | √ | √
| Bone Holding forcep 14cm | Piece | √ | √
| Bone Holding forceps 18.5cm | Piece | √ | √
| Bone Holding forceps, self-centering, length 191mm, softlock. | Piece | √ | √
| Bone Reduction forcep with points, narrow 161mm softlock | Piece | √ | √
| Bone Reduction forcep with points, wide 127mm softlock | Piece | √ | √
| Bone Rongeur 12.5cm | Piece | √ | √
| Bone Rongeur 16.5cm | Piece | √ | √
| Bone Rongeurs Luer 14.5cm | Piece | √ | √
| Browne lingual flap retractor | Piece | √ | √
| Chin Retractor Obwegeser 16.5cm | Piece | √ | √
| Condyle retractor - thin - AK's new | Piece | √ | √
| Cleft lip surgery - set | Set | √ | √
| Cleft plate surgery set | Set | √ | √
| Crown wire twisting forceps: 16cms/6 1/4 inch | Piece | √ | √
| Dingmans mouth gag | Piece | √ | √
| Dissecting Forceps, Adson-brown 12cm | Piece | √ | √
| Dissecting Forceps, Gillies 15.5cm | Piece | √ | √
| Dissecting Forceps, Potts-Smith 16cm | Piece | √ | √
| Dissecting Forceps, Potts-Smith 18cm | Piece | √ | √
| Dissecting Forceps, Potts-Smith 23cm | Piece | √ | √
| Finger protector: langenbeck | Piece | √ | √
| Ferguson jaw opener | Piece | √ | √
| Hammers length 230mm, weight 350g. | Piece | √ | √
| Heister jaw opener - SS | Piece | √ | √
| Jaw Splint - Small | Piece | √ | √
| Jaw Splint - Large | Piece | √ | √
| Lip & Check Retractor - Cawood - Minnesota | Piece | √ | √
| Lip & Check Retractor - Oringer | Piece | √ | √
Mandibular Modular Fixation System with following contents:
| Mandibular Modular Fixation System with following contents: | Set | √ | √
| Mandible Trauma Module Set (2.4mm) consisting of: | Set | √ | √
Maxillofacial Modular Fixation System with following contents:
| Module 1.5mm | Piece | √ | √
| Module 2.0mm | Piece | √ | √
| Maxillary awl Obwegeser 15cm | Piece | √ | √
| Maxillary awl Obwegeser 17.5cm | Piece | √ | √
| Mouth gag Mckesson - adult | Piece | √ | √
| Mouth gag Mckesson - Child | Piece | √ | √
| Needle holder 16cm | Piece | √ | √
| Needle holder ,Mathieu,170mm | Piece | √ | √
| Needle holder 23cm | Piece | √ | √
| Obwegeser Retractor 24cm | Piece | √ | √
| Osteotomy Chisel 10mm | Piece | √ | √
| Osteotomy Chisel 15mm | Piece | √ | √
| Periosteal Elevator Joseph | Piece | √ | √
| Ramus retractor | Piece | √ | √
| Retractor Adson | Piece | √ | √
| Rowe's Mandible bone holding forceps | Piece | √ | √
| Rowe's Maxilla disimpaction clamp - left | Piece | √ | √
| Rowe's Maxilla disimpaction clamp - right | Piece | √ | √
| Rowe's zygomatic elevator | Piece | √ | √
| Scissor Iris 11.5cm/ 4.5 inch | Piece | √ | √
| Scissor Metzenbaum - fino 18cm/ 7 inch | Piece | √ | √
| Scissor Tenotomy 115mm | Piece | √ | √
| Sinus Lift curettes 15cm | Piece | √ | √
| Sinus Lift curettes 18.5cm | Piece | √ | √
| Surgical Skin hooks | Piece | √ | √
| Third Molar Retractor Obwegeser 24cm | Piece | √ | √
| Tissue retractor Austin | Piece | √ | √
| Tongue depressor(wider) | Piece | √ | √
| Wire cutter,Straight | Piece | √ | √
| Wire ligature pliers for circum mandibuler | Piece | √ | √
| Wire ligature pliers for circum zygomatic | Piece | √ | √
| Wire pusher Rowe | Piece | √ | √
| Zygoma elevator Gillies 20cm | Piece | √ | √
| Zygomatic bone awl curved/ zygomatic arch awl | Piece | √ | √
| Zygomatic bone hook with T handle | Piece | √ | √
| Maxillofacial Modular Fixation System with following contents: a) Obwegeser Retractor 24cm; b)Third Molar Retractor Obwegeser 24cm | Set | √ | √
| Elcomed, contra angle, straight HP, motor cable and motor holder, surgical hand piece (W&H) | set | √ | √
| Tungsten carbide burs for surgical- Round (2.35mmX30-40mm) | piece | √ | √
| Tungsten carbide burs for surgical-Taper criss-cut (2.35mmX30-40mm) | piece | √ | √
| Tungsten carbide burs for surgical-Cylindrical (2.35mmx30-40mm) | piece | √ | √
| Drill Bits with stops, for Stryker coupling, length- 44.5mm, diameter-1.55mm | piece | √ | √
| Drilling depth 6mm | piece | √ | √
| Drilling depth 8mm | piece | √ | √
| Drilling depth 12mm | piece | √ | √
| Bur block stainless steel (sterilizable) for surgical | Piece | √ | √
Orthodontics
Consumables
| Orthocryl Powder Ploymer, Green | 1 kg/ pkt | √
| Orthocryl liquid Monomer / Green | 50ml | √
| Orthocryl Powder Ploymer, Neon | 1kg/pkt | √
| Orthocryl liquid Monomer / Neon | 50ml | √
| Orthocryl Powder Ploymer, Purple | 1kg/pkt | √
| Orthocryl liquid Monomer / Purple | 50ml | √
| Orthocryl Powder Ploymer , Yellow | 1kg/pkt | √
| Orthocryl liquid Monomer/ Yellow | 50ml | √
| Low fusing silver soldering flux | spool | √
| Anti-flux | 50g | √
| Self-Igniting Butane Micro-Torch (refillable): gas torch that creates a precise flame emitting from a nozzle head, allowing direct heat application to an object. | pcs | √
| Pressurized refill butane Gas for Self-Igniting Butane Micro-Torch | 250ml/can | √
| Mouth gag rubber, pediatric | Piece | √ | √
| Orthocryl Powder Ploymer 1 kg/ pkt, Green | Packet | √ | √
| Orthocryl liquid Monomer / Green | 50ml | √ | √
| Orthocryl Powder Ploymer, Neon | 1kg/pkt | √ | √
| Orthocryl liquid Monomer / Neon | 50ml | √ | √
| Orthocryl Powder Ploymer, Purple | 1kg/pkt | √ | √
| Orthocryl liquid Monomer / Purple | 50ml | √ | √
| Orthocryl Powder Ploymer , Yellow | 1kg/pkt | √ | √
| Petralit cement, (Silico-phosphate cement), powder 60 gm & liquid 22 ml | Set | √ | √ | √
Equipments
| Dual wheel model trimmer,ower supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top. | Piece | √
| Essix Vacuum Thermoforming Machine, Power supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top. Appropriate voltage stabilizer should be provided with the unit. | Piece | √
| Sterilization sleave sealer machine, Power supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top. | Piece | √
| 022 Damon brackets with cuspid hook | Set | √
| 3-dimensional screw – (Bertoni) | Piece | √
| APC- adhesive precoated brackets MBT 022 slot | Set | √
| Arch hooks Rectangular | Packet | √
| Ball end clasps 028, 032 | Packet | √
| Bicuspid bands kit, upper right & left, with 022 bracket prewelded | Packet | √
| Bicuspid bands kit, lower right & left, with 022 bracket prewelded | Packet | √
| Bite turbos | Packet | √
| Bondable buttons | Packet | √
| Bondable lingual retainer 3 x 3 | Packet | √
| Bondable lingual retainer 4 x 4 | Packet | √
| Bracket storage box (for individual brackets) | Box | √
| Brass wire | Spool | √
| Clarity Ceramic bracket 5 - 5 (022 slot) | Set | √
| Copper NiTi 017x0.025 upper | Packet | √
| Copper NiTi 017x0.025 lower | Packet | √
| Copper NiTi 0.014 inch upper | Packet | √
| Copper NiTi 0.014 inch lower | Packet | √
| Copper NiTi 0.016 inch upper | Packet | √
| Copper NiTi 0.016 inch lower | Packet | √
| Copper NiTi 0.018 inch upper | Packet | √
| Copper NiTi 0.018 inch lower | Packet | √
| Crimpable hook 019 x 025 | Packet | √
| Cuspid bands kit, upper,left/right 022 bracket prewelded, Kit/s | Pack | √
| Cuspid bands kit, lower,left/right,022 bracket prewelded, Kit/s | Pack | √
| E-chain elastics closed, black | Spool | √
| E-chain elastics closed, blue | Spool | √
| E-chain elastics closed, Clear | Spool | √
| E-chain elastics closed, Green | Spool | √
| E-chain elastics closed, orange | Spool | √
| E-chain elastics closed, pink | Spool | √
| E-chain elastics closed, purple | Spool | √
| E-chain elastics closed, red | Spool | √
| E-chain elastics continuous, Clear | Spool | √
| E-chain elastics continuous, Gray | Spool | √
| E-chain elastics open, Clear | Spool | √
| E-chain elastics open, Gray | Spool | √
| Elastic separators | Stick | √
| Elastic thread 50ft length | Spool | √
| Etchant gel (9.6% hydrofluoric acid) 4 syringes of 1.2ml and 8 pre-bent applicator tips. | Piece (Syringe of 1.2 ml) | √
| Etchant gel (38% phosphoric acid) with 50 pre –bent applicator tips, Thixotropic | Piece (Syringe of 25 ml) | √
| Expansion screw safety keys | Packet | √
| Extraoral elastics 1/2" 8 OZ | Packet | √
| Extraoral elastics 1/2" 14 Oz | Packet | √
| Extraoral elastics ¼” 8 Oz | Packet | √
| Extraoral elastics ¼” 14 Oz | Packet | √
| Extraoral elastics 3/16” 8 Oz | Packet | √
| Extraoral elastics 3/16” 14 Oz | Packet | √
| Extraoral elastics 3/8” 8 Oz | Packet | √
| Extraoral elastics 3/8” 14 Oz | Packet | √
| Extraoral elastics 5/16” 8 Oz | Packet | √
| Extraoral elastics 5/16” 14 Oz | Packet | √
| Fan-type expansion screw | Piece | √
| Flexible Double Ended plastic self-retaining Cheek Retractor, Paedo Small | Piece | √
| Forsus Kit (3M) | Pkt | √
| Gemini Roth 022 brackets with 4 tubes (UT/LD) | Set | √
| Glass Ionomer Band Cement | Pkt | √
| Headgear kit | Pkt | √
| Hyrax expansion screw, 11mm | Piece | √
| Hyrax expansion screw, 13mm | Piece | √
| Hyrax expansion screw, 8mm | Piece | √
| Hyrax expansion screw, 9mm | Piece | √
| Intraoral elastics ¼" 3.5 Oz | Pkt | √
| Intraoral elastics ¼” 6 Oz | Pkt | √
| Intraoral elastics 1/8" 2 Oz; | Pkt | √
| Intraoral elastics 1/8” 3.5 Oz | Pkt | √
| Intraoral elastics 3/16” 3.5 Oz | Pkt | √
| Intraoral elastics 3/8” 2 Oz | Pkt | √
| Intraoral elastics 3/8” 3.5 Oz | Pkt | √
| Intraoral elastics 5/16” 2 Oz | Pkt | √
| Intraoral elastics 5/16” 3.5 Oz | Pkt | √
| Intraoral elastics 5/16” 4.5 Oz | Pkt | √
| Intraoral stops 2 mm length | Pkt | √
| Intraoral stops 1mm length | Pkt | √
| Lingual cleats | Pkt | √
| Lip bumper with loops | Pkt | √
| Lower expander screw | Piece | √
| Mandibular bow screw | Piece | √
| Micro sectional screw | Piece | √
| Mini Twin Vs 022 brackets | Set | √
| NiTi 0.014 lower | Pkt | √
| NiTi 0.014 upper | Pkt | √
| NiTi 0.016 lower | Pkt | √
| NiTi 0.016 upper | Pkt | √
| NiTi 0.017x0.025 lower | Pkt | √
| NiTi 0.017x0.025 upper | Pkt | √
| NiTi 0.018 lower | Pkt | √
| NiTi 0.018 upper | Pkt | √
| NiTi 0.019 x 0.025 lower | Packet | √
| NiTi 0.019 x 0.025 upper | Packet | √
| NiTi closed coil long | Spool | √
| NiTi closed coil medium | Spool | √
| NiTi closed coil small | Spool | √
| NiTi Closed coil spring- heavy force 200g | Packet | √
| NiTi Closed coil spring- light force 100g | Packet | √
| NiTi Closed coil spring- medium force 150g | Packet | √
| NiTi distalizing (stop wound coil) spring | Packet | √
| NiTi open coil spring - heavy | Packet | √
| NiTi open coil spring - light | Packet | √
| NiTi open coil spring - medium | Packet | √
| NOLA dry field system-cheek retractor, Tongue Guards,"Y" Connector, Suction tubes and adapter | Set | √
| O-ring modules Black | Pkt | √
| O-ring modules Blue | Pkt | √
| O-ring modules Clear | Pkt | √
| O-ring modules Green | Pkt | √
| O-ring modules Light Blue | Pkt | √
| O-ring modules Maroon | Pkt | √
| O-ring modules Neon Green | Pkt | √
| O-ring modules Orange | Pkt | √
| O-ring modules Pink | Pkt | √
| O-ring modules Purple | Pkt | √
| O-ring modules Red | Pkt | √
| O-ring modules Silver | Pkt | √
| O-ring modules Teal | Pkt | √
| O-ring modules Tooth color | Pkt | √
| O-ring modules Violet | Pkt | √
| O-ring modules Yellow | Pkt | √
| Patient relief wax | Pkt | √
| Preformed Kobayashi ligature ties | Packet | √
| Preformed Left lower first molar band with cleats kit, lower double tube | Set | √
| Preformed Left upper first molar band with cleats kit, triple tube | Set | √
| Preformed ligature long ties | Packet | √
| Preformed Right lower first molar band with cleats kit,lower double tube | Set | √
| Preformed Right upper first molar band with cleats kit, triple tube | Set | √
| Preformed second molar band kits lower left, single tube | Set | √
| Preformed second molar band kits lower right, single tube | Set | √
| Preformed second molar band kits upper left, single tube | Set | √
| Preformed second molar band kits upper right, single tube | Set | √
| Preformed short ligature ties | Packet | √
| Rapid molar distalizer | Packet | √
| Retraction screw - medium | Piece | √
| Retraction screw - Mini | Piece | √
| Rotation wedges | Pkt | √
| Sectional screw - Medium | Piece | √
| Sectional screw - Mini | Piece | √
| Stainless steel wire 0.014 lower | Packet | √
| Stainless steel wire 0.014 upper | Packet | √
| Stainless steel wire 0.016 lower | Packet | √
| Stainless steel wire 0.016 upper | Packet | √
| Stainless steel wire 0.016 x 0.022 lower | Packet | √
| Stainless steel wire 0.016 x 0.022 upper | Packet | √
| Stainless steel wire 0.017 x 0.022 lower | Packet | √
| Stainless steel wire 0.017 x 0.022 upper | Packet | √
| Stainless steel wire 0.018 lower | Pkt | √
| Stainless steel wire 0.018 upper | Pkt | √
| Stainless steel wire 0.018x0.025 lower | Pkt | √
| Stainless steel wire 0.018x0.025 upper | Pkt | √
| Stainless steel wire 0.019x0.025 lower | Pkt | √
| Stainless steel wire 0.019x0.025 upper | Pkt | √
| Stainless steel wire 0.021x0.025 lower | Packet | √
| Stainless steel wire 0.021x0.025 upper | Packet | √
| Stop/spacer 2mm - 022 | Pkt | √
| Stop/spacer 2mm - 022 x 028. | Packet | √
| Stop/spacer 2mm - 030 | Packet | √
| Stripping Strips, Diamond studded | Packet | √
| Thermoforming materials - Essix A+ 1.5mm | Pkt | √
| Thermoforming materials - Essix A+ 1mm | Pkt | √
| Thermoforming materials - Essix Embrace 1mm | Pkt | √
| Thermoforming materials - Essix Embrace 2mm | Pkt | √
| Thermoforming materials - Essix Night Guard | Pkt | √
| Thermosealer - 10mm, Role | Roll | √
| Thermosealer - 5mm, Role | Roll | √
| Thermosealer - 7.5mm, Role | Roll | √
| Tissue guard (clear), 0.027 inch | Spool | √
| TMA preformed pendulum springs – left and right | Packet | √
| TMA Reverse curve with T-loops lower 0.017x0.025 (broad) | Packet | √
| TMA Reverse curve with T-loops lower 0.016x0.022 (broad) | Packet | √
| TMA Reverse curve with T-loops upper 0.017x0.025 (broad) | Packet | √
| TMA Reverse curve with T-loops upper 0.016x0.022 (broad) | Packet | √
| TMA Upper 0.017x0.025 | Pkt | √
| TMA Lower 0.017x0.025 | Pkt | √
| TMA upper 0.016x0.022 | Packet | √
| TMA lower 0.016x0.022 | Packet | √
| TMA upper large 0.018, U 016 x 022 | Packet | √
| TMA lower large 0.018, L 016 x 022 | Packet | √
| Tooth stripper files | Packet | √
| Traction screw | Piece | √
| Transbond adhesive kit plus primer + Etchant, self cure | Set | √
| Transbond kit (light cure), plus primer and etchant | Piece | √
| Trapezoidal Expansion screw | Piece | √
| Universal expansion screw - Maxi | Piece | √
| Universal expansion screw - medium | Piece | √
| Universal expansion screw - mini | Piece | √
| Weldable buttons | Packet | √
| Wilcock SS round wire, 012 Premium plus | Spool | √
| Wilcock SS round wire, 014 Premium plus | Spool | √
| Wilcock SS round wire, 016 Premium plus | Spool | √
| Wilcock SS round wire, 018 Premium plus | Spool | √
| Wilcock SS round wire, 020 Premium plus | Spool | √
| Orthocryl Powder Ploymer, Green | 1 kg/ pkt | √
| Orthocryl liquid Monomer / Green | 50ml | √
| Orthocryl Powder Ploymer, Neon | 1kg/pkt | √
| Orthocryl liquid Monomer / Neon | 50ml | √
| Orthocryl Powder Ploymer, Purple | 1kg/pkt | √
| Orthocryl liquid Monomer / Purple | 50ml | √
| Orthocryl Powder Ploymer , Yellow | 1kg/pkt | √
| Orthocryl liquid Monomer/ Yellow | 50ml | √
| Low fusing silver soldering flux | spool | √
| Anti-flux | 50g | √
| Self-Igniting Butane Micro-Torch (refillable): gas torch that creates a precise flame emitting from a nozzle head, allowing direct heat application to an object. | pcs | √
| Adhesive removing pliers with diamond inset | Piece | √
| Bird beak pliers with cutter | Piece | √
| R.E.D. Distraction Hallo | Set | √
| Arch hook crimping pliers | Piece | √
| Arch wire turret | Piece | √
| Band contouring pliers (Johnson) | Piece | √
| Band cutting scissors-curved | Piece | √
| Band slitting pliers | Piece | √
| Bite plane thermoplier | Piece | √
| Bite stick (band adapter) autoclavable | Piece | √
| Bonding Organization tray | Piece | √
| Bracket holder with alignment tip | Piece | √
| Cinch back pliers | Piece | √
| Convertible cap removal tool kit with replacement tip | Piece | √
| Direct bracket remover pliers | Piece | √
| Distal end cutter safety hold | Piece | √
| Distal twister | Piece | √
| Double ended band seater | Piece | √
| Elastic hook-forming thermoplier | Piece | √
| Essix IPR Deluxe kit Raintree USA | Piece | √
| Heavy wire bending pliers (Headgear pliers set) | Piece | √
| How plier | Piece | √
| Jarabak plier | Piece | √
| Ligature cutter small | Piece | √
| Ligature cutter standard | Piece | √
| Ligature tucker | Piece | √
| Mallet handle band pusher (Mershon) | Piece | √
| Mandibular tooth moving thermoplier | Piece | √
| Maxillary tooth moving thermoplier | Piece | √
| Mesial-distal thermoplier | Piece | √
| Micro ramp thermoplier | Piece | √
| Mini Mathieu needle holder | Piece | √
| Nance Tie Back Plier | Piece | √
| Optical pliers | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed lower L1 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed lower L2 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed lower L3 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed lower L4 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed Upper U1 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed Upper U2 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed Upper U3 | Piece | √
| Orthodontic Impression tray, Rim Locked, Paed Upper U4 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult lower L1 | Piece | √
| Orthodontic Impression trays perforated, rim locked,adult lower L2 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult lower L3 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult lower L4 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult, upper U1 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult, upper U2 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult, upper U3 | Piece | √
| Orthodontic Impression trays perforated, rim locked adult, upper U4 | Piece | √
| Photo Cheek retractors, large, set of 3 | Piece | √
| Photo Cheek retractors, Small, set of 3 | Piece | √
| Posterior band removing pliers | Piece | √
| Rectangular arch forming pliers (Tweed) | Piece | √
| Slot forming thermoplier | Piece | √
| Stackable Band Organizer | Piece | √
| Surgical hook crimping pliers | Piece | √
| Torquing pliers | Piece | √
| Tweed loop forming pliers | Piece | √
| Undercut Enhancing thermoplier | Piece | √
| Utility arch/loop forming pliers (Rickettts) | Piece | √
| V-bending pliers | Piece | √
| Weingart utility pliers | Piece | √
| Wire contouring pliers | Piece | √
Paedodontic
Consumables
| Acidulated Phosphate Fluoride(APF) 1.23% gel(12,300 ppmF) strawberry flavor 17.5 oz | Bottle of 17.5 Oz | √ | √
| Sodium Fluoride rinse 0.2% (920 ppmF) | 1Ltr/ bottle | √ | √
| Sodium Fluoride rinse 0.05%(230 ppmF) | 250ml/ bottle | √ | √
| Fluoride Tablets 0.25mg | 1000 Tabs/ Pkt | √ | √
| Fluoride Tablets 0.50mg | 1000 Tabs/ Pkt | √ | √
| Sodium Fluoride Varnish 5% ( 2.26% F)10 ml tube with applicator brush | 50tubes/ pkt | √ | √
| Precontoured Primary Molar Bands | Set | √ | √
| Precontoured Permanent First Molar bands | Set | √ | √
| Stainless steel crowns for primary dentitions(Pre-contoured) | Set | √ | √
| Crown scissors, curved serrated blade 4" | Piece | √ | √
| Crown Scissors,straight,110 mm | Piece | √ | √
| Transparent Pediatric Strip Crown Forms | Set | √ | √
| Molt Mouth Gag, external with silicon tip, child/infant, 11cm | Piece | √ | √
| Molt Mouth Gag, external with silicon tip, medium,14 cm | Piece | √ | √
| Polycarbonite lip cheek retractor with wings 0118WC (child) | Box of 3 pieces | √ | √
| Polycarbonite lip cheek retractor with wings 01189WM(Medium) | Box of 3 pieces | √ | √
| Papoose board padded, small with chest strap and abdominal flaps (30” x 7”) | Piece | √ | √
| Papoose board padded, regular with chest strap and abdominal flaps (36” x 11”) | Piece | √ | √
| Dual Arch Fluoride plastic tray, Pedo(aqua) | Set | √ | √
| Dual Arch Fluoride plastic tray, Large (blue) |  | √ | √
| Dual Arch Fluoride plastic tray, medium (White) | Set | √ | √
| Dual Arch Fluoride plastic tray, small (yellow) | Set | √ | √
| DSP IOPA X-ray film holder DR540501(Snap-A-Ray) | 1pcs/ pkt | √ | √
| X-ray apron with collar, pediatric 0.3mm (21" x 22") | Piece | √ | √
| Rubber dam 5”x5”, thin green, Mint (Box) | 52pcs/ box | √ | √
| Rubber dam clamp, SS White,No.206 | Set | √ | √
| Rubber dam clamp, SS White,No.207 | Set | √ | √
| Rubber dam clamp, SS White No. 208 | Set | √ | √
| Rubber dam clamp No.12 | DOPC | √ | √
| Rubber dam clamp No. 13 | Set | √ | √
| Rubber dam clamp, Ivory,No.14 | Set | √ | √
| Rubber Dam Clamp No.0 | Set | √ | √
| Rubber Dam Clamp No.209 | Set | √ | √
| Rubber dam clamp, Ivory, No.8-A | Set | √ | √
| Rubber dam clamp, Ivory, No. 14-A | Set | √ | √
| Rubber dam punch, Ainsworth 160mm | Piece | √ | √
| Rubber Dam Clamp Forceps, Ivory 170mm | Piece | √ | √
| Rubber dam frame(Young’s) 5”x5” | Piece | √ | √
| Abell Contouring pliers N0.112 | Piece | √ | √
| Reynolds contouring pliers 125mm/5” | Piece | √ | √
| Gordon Contouring Pliers 5¼ inch. | Piece | √ | √
| Johnson contouring pliers 135mm/5¼ “. | Piece | √ | √
| Contouring pliers ultra thin No.007-118 | Piece | √ | √
| Crimping Pliers,No.417 | Piece | √ | √
| How plier,straight,135mm | Piece | √ | √
| Transparent Polyester Matrix, 0.05 mm/0.02” thick x 59 ft | Roll | √ | √
| soft stainless matrix, 0.025 mm/0.001” thick x20 foot | Roll | √ | √
| Ferric sulphate 15.5% dropper bottle of 59ml(2oz) | 59ml/ bottle | √ | √
| Reinforced ZOE cement (IRM) Powder 40g,15ml liquid | Set | √ | √
| Needle, 1" 30G (0.40x21mm) short 25 mm | 100pcs/ pkt | √ | √
| Benzocaine topical anesthetic gel 20% 1 oz/29.6ml jar ,raspberry flavour | 25pcs/ pkt | √ | √
Prostodontic
Consumables
| Temporary ST. steel 3M , readymade crowns, standard set/pkt, 3M/Bosworth | Set | √ | √ | √
| Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-597-00 | Piece | √ | √ | √
| Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-598-00 | Piece | √ | √ | √
| Cloth polishing buff SHP mounted,Denta, 181-100-01 | Piece | √ | √ | √
| Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-582-00 | Piece | √ | √ | √
| Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-583-00 | Piece | √ | √ | √
| Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-584-00 | Piece | √ | √ | √
| Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-583-00 | Piece | √ | √ | √
| Optra-Dam clamp-free rubber dam. Blue, Box of 5o, Ivoclar | 50sheets/ Pkt | √ | √ | √
| Optra-Gate medium, Box of 5o, Ivoclar | 50/kt | √ | √ | √
| Bur Silicone acrylic small, flame shaped. Shofu- HP/0423-6/Pcs, Grit-Course-Gray | Packet/set | √ | √ | √
| Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0427- Grit-Medium Brown | 6pcs/ pkt | √ | √ | √
| Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0428- Grit-Fine Light-Gray | 6pcs/ pkt | √ | √ | √
| Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0426-6/Pcs, Grit-Course-Gray | Packet/set | √ | √ | √
| Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0427-6/Pcs, Grit-Medium Brown | Packet/set | √ | √ | √
| Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0428-6/Pcs, Grit-Fine Light-Gray | Packet/set | √ | √ | √
| Silicone polishing Cone, Green medium grit, Small | Packet/set | √ | √
| Silicone polishing Cone, Green, fine grit, Small | Packet/set | √ | √
Instrunment
| Hanau Alcohol hand Torch plastic for prostho-lab waxing | Piece | √ | √ | √
| Hoopler duplicator, upper/lower set,for CD | Set | √ | √ | √
| Willis gauze for CD for OVD measurement | Piece | √ | √ | √
| Silicone-Rubber Model Base former U/L set (Large) | Set | √
| Silicone-Rubber Model Base former U/L set (Medium) | set | √
| Silicone-Rubber Model Base former U/L set (Small) | set | √
List for Porcelain-Ceramic Work (consumables)
| Porcelain Vita dentin B1 | 50gm/ pkt | √
| Porcelain Vita dentin B2 | 50gm/ pkt | √
| Porcelain Vita dentin B3 | 50gm/ pkt | √
| Porcelain Vita dentin B 3.5 | 50gm/ pkt | √
| Porcelain Vita dentin B 4 | 50gm/ pkt | √
| Porcelain Vita dentin C1 | 50gm/ pkt | √
| Porcelain Vita dentin C2 | 50gm/ pkt | √
| Porcelain Vita dentin C3 | 50gm/ pkt | √
| Porcelain Vita dentin C3.5 | 50gm/ pkt | √
| Porcelain Vita dentin C4 | 50gm/ pkt | √
| Porcelain Vita Opaque A1 | 50gm/ pkt | √
| Porcelain Vita Opaque A2 | 50gm/ pkt | √
| Porcelain Vita Opaque A3 | 50gm/ pkt | √
| Porcelain Vita Opaque A3.5 | 50gm/ pkt | √
| Porcelain Vita Opaque A4 | 50gm/ pkt | √
| Porcelain Vita Opaque B1 | 50gm/ pkt | √
| Porcelain Vita Opaque B2 | 50gm/ pkt | √
| Porcelain Vita Opaque B3 | 50gm/ pkt | √
| Porcelain Vita Opaque B3.5 | 50gm/ pkt | √
| Porcelain Vita Opaque B4 | 50gm/ pkt | √
| Porcelain Vita Opaque C1 | 50gm/ pkt | √
| Porcelain Vita Opaque C2 | 50gm/ pkt | √
| Porcelain Vita Opaque C3 | 50gm/ pkt | √
| Porcelain Vita Opaque C3.5 | 50gm/ pkt | √
| Porcelain Vita Opaque C4 | 50gm/ pkt | √
| Porcelain Vita Opaque dentine | 50gm/ pkt | √
| Porcelain Vita Wash Opaque | 50gm/ pkt | √
| Gutta –cut, rechargeable unit; 220-230 voltage, complete sets | set | √
| Ever-stick fabre, GC/Crown &Bridge,1.5mm/pkt | pkt | √
| Ever-stick fabre, GC/Perio, 1.5mm/pkt | pkt | √
| Ever-stick fabre, GC/post & core systems, 1.5mm/pkt | pkt | √
| Compomer TM Brilliant Veneer systems complete sets, Wheledent, pkts | pkt | √
| Crown cutter carbide burs, 8G, Striaght fissure, Wheledent | pcs | √
| Magic foam cord- Gingival retraction complete sets with Dispenser and syringe /pkts, | 10/ pkts | √
Dermatology Department
Equipments
| Carbon dioxide (CO2) Laser and accessories | Set | √
| Fibro Optic (spotlight) for enhancing Leisons in genitals. | Set | √
| RF-300 (Radio Frequency) and accessories | Set | √
| Ultra Violet Ray-A (UVA)-Chamber and accessories | Set | √
| Fluorescent Pulsed Light (FPL-Plasmalite) and accessories | Set | √
| Dermatoscope- Dermlite DL3 or Welsh Allyn | Set | √
| Patch test kit - International standard series, with finn chambers | Set | √
Consumables
| Biopsy Punch (which is sterile, disposable skin biopsy punches ranging from 5mm to 8mm) | Pkt | √
| Biopsy punch - disposable, size 2mm | Pkt
| Biopsy punch - disposable, size 3mm | Pkt
| Biopsy punch - disposable,size 4mm | Pkt
| TCA (Trichlor Acitic Acid) with applicator | Jar | √
Instrunment
| Instrument automatic washer/disinfectors | Set | √
| Ultrasonic cleaning basin, integrated in a work bench | Piece | √
| Ultrasonic cleaner with lifting system for trays, with lid | Set | √
| Ultravoilet Woods lamp set for | Set | √
| Cryo surgery instruments with accessories and liquid nitrogen (very small liquid nitrogen barrel) Power supply 220-240v, 50Hz. Suitable stabilizer. Power cord with round 3 pin top | Set | √
| Skin hook | Piece | √
| Comedo extracter | Piece | √
| Skin curette | Piece | √
| Iris scissor | Piece | √
| Gradle scissor | Piece | √
| Baby Metzenbaum undermining scissor | Piece | √
| UV diagnostic Wood's lamp | set | √
Emergency Department
Equipments and consumables added with the OT and aneathesia section
ENT Department
Equipments
| Xenon Cold Light Fountain Compact unit with built in 300 Watt XENON lamp with manually variable light intensity from 0 - 100 %, Provides a light intensity which meets all requirements even for use with fiberscope or narrow-lumen rigid endoscopes, Color Temperature of 6000 Kelvin for daylight illumination, Highest Color Rendering Index (CRI) for naturally color recognition, Maximum light intensity immediately after power up, Fan for silent air-ventilation Novel type of light source adapter for improved heat dissipation at the proximal ends of fiber optic light cables.Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √ | √
| Head light With ball joint and vertical height adjustment, front section movable in all directions, fiber optic light cable included, with fully adjustable comfortable head band. Power cord with round 3 pin top | Set | √ | √
| Otoscope With rubber bulb and spare light bulb with inbuilt rechargeable battery | Set | √ | √ | √ | √
| Rhino-Laryngo-Fibre scope Instrument channel 1.2 mm direction of view 0 degree, angle of view 110 degree, working length 34 cm, outer diameter 3.5 mm.Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √
| Diagnostic Audiometer, Available Signals: Stimulus: Pure tone, warble tone, 2 external inputs and USB audimetry. MIC input for live speech audiometry.Masking: NBN, WN, SN SIGNALS SPECIFICATION Attenuator step: 1 and 5 dB, Presentation: Continuous, Pulsed (0.5, 1 and 2 Hz) Warble: 5 Hz sin wave modulating signal AVAILABLE OUTPUTS AND TRANSDUCERS AC: Headphones (TDH-39 / DD45) HF headphones (HDA-300) Insert earphones (ER-3A / ER-5A) BC: B-71 bone vibrator Free field Insert masking earphone: IME-100 AVAILABLE TESTS • Pure Tone audiometry (HL and UCL)• Auto threshold (modified Hughsone-Westlake) • HF audiometry: from 8 to 20 kHz (optional) • Speech audiometry (SRT, WRS, MCL, and UCL) • Speech audiometry (SRT, WRS, MCL, and UCL) • SISI test • 2 independent channels Master Hearing Aid • QuickSIN® test (optional) • Visual Reinforcement Audiometry (VRA) and Conditioned Play Audiometry (CPA) test (optional) with video reinforces INCLUDED PARTS • TDH-39 supra-aural headphones • B71 bone vibrator • Patient response switch • Monitor headset with boom microphone • Clip-on microphone for patient-to-operator communication • Medical grade power supply • USB connection cable • User manual • Inventis Software Suite CD • Noah Maestro module • Daisy software and Daisy Maestro module • Four magnetic covers OPTIONAL ACCESSORIES (with order code) | Set | √ | √
| OPTIONAL ACCESSORIES • ER-3A insert earphones • ER-5A insert earphones • IME-100 insert masking earphone • DD45 complete headset • Amplivox Audiocups noise excluding enclosures for TDH- 39 and DD45 headphones • Desktop microphone for live speech tests • Cable set for sound booths with jack sockets • Cable set for sound booths with RCA sockets • One active speaker FBT J 5A (230V model) • One active speaker FBT J 5A (120V model) |  | √
| Play audiometer | Set | √
| OTOREAD SCREENER TEOAE Technical Specifications • Measurement type: OtoAcoustic Emissions • Frequency range DPOAE: 1.5, 2, 2.5, 3, 3.5, 4, 5, 6, 8, 10, 12 kHz • Frequency range TEOAE: 0.7, 1, 1.4, 1.5, 2, 2.5, 2.8, 3.5, 4 kHz• Stimulus intensity: 40 to 70 dB SPL (DPOAE). 83 dB SPL (TEOAE) • Maximum output (Protection): 90 dB (This level is well within OSHA permissible limits of 90 dBA for 8 hours) • Microphone system noise:-20 dB SPL @ 2 KHz(1 Hz bandwidth)-13 dB SPL @ 1 KHz (1 Hz bandwidth) • Battery life: Approximately 300 tests • Display: LCD-display 4 line x 10 character • Weight: 300 g/ 10.6 oz. including batteries • Internal software languages: English, German, French, Spanish, Russian • Screener 6 bands (1 fixed protocol) 1.5-4 kHz Standard accessories • Handheld OtoRead™ unit including probe cord and cradle • Printer including power supply and power cable • Printer Cable • Carrying Bag • Probe cord for extension (100 cm/39 inches) • 2 Thermal rolls paper • Box of 146 eartips (10 size; 3, 4, 6, 7, 8, 9, 10, 11, 12, 13 mm) • 4 probe tips • 4 AA/UW3/R6 alkaline batteries | Set | √ | √ | √
| Tympanometer (Ipsilateral & contralateral) | Set | √ | √
| Brainstem evoked response audiometer(ABR), complete set with software | Set | √
| portable Audiometer: ●Audiometric Headset ● Bone Conductor ● Patient Response Button●External Power Supply ●Audiogram Charts ●Operation Manual CD ●Multilingual CE Instructions for Use ● Amplivox Audiocups, noise reducing headset ●The lightweight carrying case will hold the as well as a noise reducing headset and audiogram charts etc. ●Pure, Warble and Pulsed tones ●Automatic Threshold test ● SISI and Bekesy tests ● Stenger and ABLB tests ●Tone in Noise test (Langenbeck) ● 125Hz - 8000Hz ……….(125-250hz to 8000hz) ●-10dB to 120dBHL output ● Air and Bone Conduction ● WN and NB masking ●5dB and 1dB attenuators ● Insert Phones (optional) | Set | √ | √
| Hearing Aid analyzer/programmer specifications: ●Frequency Range: 125 to 12500 Hz ● Stimuli: Pink, whilte, speech noise, ISTS, tone sweep ●Stimulus Level: 50-90 dB SPL, 1 or 5 dB steps ● Frequencies per Octave (sweep): 24, 12, 6, 3 ●Frequency Resolution: 1/24, 1/12, 1/6, 1/3 ● Dynamic Range:> 100 dB ● Coupler Mic Intensity:20-150 dB ●Reference Mic Intensity:20-100 dB ●Power:15V / 2 Amp DC ●Size: (W x D x H) 13.7 x 13.7 x 5.5 in. (348 x 348 x 140 mm.) ●Weight:13 lbs. (5.9 kg.) PC SYSTEM REQUIREMENTS: ●Pentium-class or equivalent PC with 1.6 GHz or faster processor ●2 Gb RAM (more is recommended) ●CD-ROM reader ●Screen resolution of 1024x768 or higher ●Available USB port ●150 Mb of available hard disk space ●NOAH 3.7 or later ●Microsoft Windows XP (SP3), Windows 7 or 8 | Set | √
| HOPKINS Straight Forward Telescope: 0 degree diameter 2.9 mm, length 6 cm | Piece | √
| HOPKINS Straight Forward Telescope: 0 degree diameter 4 mm, length 18 cm | Piece | √ | √
| HOPKINS Straight Forward Telescope: 30 degree diameter 2.9 mm, length 6 cm | Piece | √
| HOPKINS Straight Forward Telescope: 30 degree diameter 4 mm length 18 cm | Piece | √ | √
| HOPKINS Straight Forward Telescope: 70 degree diameter 4 mm length 18 cm | Piece | √ | √
| Dr. Speech (software version 4) : With all standard accessories | Set | √
| Naso- meter II - Model 6400 | Set | √
| Triveni Speech Trainer: TST- 100 | Set | √
| Auditory Brainstems Response-complete set | Piece | √
| Telecam SL II Camera Control unit: Color systems PAL/NTSC, with integrated digital Image Processing Module, Power Supply 100-240 VAC 50/60 Hz. Consisting of : 20213020 Telecam SL II Camera Control Unit 400A Mains cord 2x20221070 Connecting cable, for controlling peripheral devices, Lengh,180cm, 536MK BNC/BNC Video Cable, Length 180cm, 20040082, DV Connecting Cable, 6pin to 4 pin length 450cm.Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √ | √
| Unidrive multifunction unit for ENT surgery | Piece | √ | √
| Cordless rechargeable LED head light with charger | Set | √ | √
| Mastoid drill micromotor - It should be compatible with Unidrive multifunction unit for ENT surgery | Set | √ | √
| Mastoid drill hand piece - It should be compatible with Unidrive multifunction unit for ENT surgery | Set | √ | √
| Shaver system for paranasal sinuses micro-shaver handpiece, shaver blade straight, shaver blade curved: (It should be compatible with Unidrive multifunction unit for ENT surgery) Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √ | √
| Binocular surgical loupe, magnification of 3.5 X, HRP optics, working distance of 420 mm, with extended field of view , adjustable optics right and left to select exact inter-pupillary distance, with spectacle frame and LED lighting plus other accessories. | set | √
| Mobile suction unit with wide mouthed 2 X 2 Ltr jars,mtr. capacity of - 700 mm Hg +/- 10 regulable , flutter free vaccum control knob, tubing of 8mm ID X 2 | set | √ | √
| Foot suction: with capacity of -600mmHg t 10 at 200m1 Per Stroke, Heavy Piston Pump, Non - Collapsible Suction Tubing, 1500mml capacity jar wilh over flow system & ABS Lid | set | √ | √
| Operating ENT Microscope | set | √ | √
Consumables
| SHEEHY, Ventilation tube, Teflon, sterile, ID 1.25 mm. | Piece | √
| Absorbable Gelatin Sponge | Piece | √
| Portex cuffed tracheostomy tubes, Size: 8.0mm | Piece | √
| Portex cuffed tracheostomy tubes, Size: 7.5mm | Piece | √
| Portex cuffed tracheostomy tubes: Size: 7.0mm | Piece | √
| Portex cuffed tracheostomy tubes: Size: 6.0mm | Piece | √
| Portex cuffed tracheostomy tubes: Size: 4.0mm | Piece | √
| Hearing aids: Behind the ear (Power- mild) | Piece | √
| Hearing aids, Behind the ear (Power - moderate | Piece | √
| Hearing aids, Behind the ear (Power - high) | Piece | √
| Hearing aids, Body worn (Power - mild) | Piece | √
| Hearing aids, Body worn (Power -moderate) | Piece | √
| Hearing aids- Body worn (Power - high) | Piece | √
| Portex cuffled tracheostomy tubes 5mm OD | Piece | √
| Portex cuffed tracheostomy tubes,3.0 mm OD | Piece | √
| Portex tracheostomy tubes. 2.5 mm OD | Piece | √
| Shiley Tracheostomy Tube Cuffless with Disposable Inner Cannula Fenestrated, size 4, ID 5mm, length 62mm | Piece | √
| Shiley Tracheostomy Tube Cuffless with Disposable Inner Cannula Fenestrated, size 4, ID 5mm, length 62mm | Piece | √
| Shiley™ Tracheostomy Tube Cuffless with Disposable Inner Cannula Fenestrated, size 8, ID 7.6mm, length 79mm | Piece | √
| Shiley™ Tracheostomy Tube XLT Cuffless Distal with Disposable Inner Cannula, ID 5mm, Total length 90mm, Distal extension | Piece | √
| Shiley™ Tracheostomy Tube XLT Cuffless Distal with Disposable Inner Cannula, ID 7mm, Total length 100mm, Distal extention | Piece | √
| Shiley™ Tracheostomy Tube XLT Cuffless Distal with Disposable Inner Cannula, ID 7mm, Total length 100mm, Proximal extention | Piece | √
Instrunment
| Adaptor for suction cannula: With cut off hole, LUER cone length 10cm | Piece | √
| Barany Noise Apparatus, with ear olive, SS | Piece | √
| Beales elevator : Curved, width 1.5 mm, length 16 cm, SS | Piece | √
| Belluci scissors blade: Length 8 mm, working length 8 cm, SS | Piece | √
| Belluci scissors blade: Length 8 mm, working length 8 cm, curved to left, SS | Piece | √
| adaptor for suction cannula: Length 8 mm, working length 8 cm, curved to right, SS | Piece | √
| Billeau ear loop: Length 16 cm - size 1, SS | Piece | √ | √ | √
| Billeau ear loop: Length 16 cm - size 2, SS | Piece | √ | √ | √
| Ear Hook: 90 degree, blunt, SS | Piece | √ | √ | √
| Fisch articulated retroauricular retractor, 3x3 prongs, sharp, length 17cm, SS | Piece | √
| FISCH Ear Forcep, Delicate, pointed, SS | Piece | √
| FISCH Scissor, Delicate, curved, sharp/blunt, length 10.5 cm, SS | Piece | √
| Hartmann ear dressing forcep, SS | Piece | √ | √ | √
| Hartmann ear forcep: Very fine, serrated, working length 10 cm, SS | Piece | √ | √ | √
| HARTMANN Ear Speculum, Size O.D 3 mm, SS | Piece | √ | √ | √ | √
| HARTMANN Ear Speculum, Size O.D 4 mm, SS | Piece | √ | √ | √ | √
| HARTMANN Ear Speculum, Size O.D 5 mm, SS | Piece | √ | √ | √ | √
| HARTMANN Ear Speculum, Size O.D 6 mm, SS | Piece | √ | √ | √ | √
| HARTMANN Ear Speculum, Size O.D 7.5 mm, SS | Piece | √ | √ | √ | √
| HOLMGREEN self retaining endaural speculum, Outer diameter 5 mm , SS | Piece | √
| House Dieter malleus nipper, Upbiting, working length 8 cm, SS | Piece | √
| Instrument rack, 22cm x 12.5 cm, with holder for 5 micro instruments and 5 drills, SS | Piece | √
| Jobson-Horne probe, With serrated ring and end, length 18 cm, SS | Piece | √ | √ | √ | √
| LEMPERT Elevator, Length 19 cm, 3 mm wide, SS | Piece | √
| PLESTER Ear retractor 2x2 prongs, SS | Piece | √
| PLESTER Elevator, Length 18 cm, 8 mm wide, SS | Piece | √
| PLESTER Knife, Round, vertical, length 16 cm, standard size 3.5x2.5mm, SS | Piece | √
| PLESTER retractor with biprong blade on left and solid blade on right, length 11cm, SS | Piece | √
| PLESTER retractor with biprong blade on right and solid blade on left, length 11cm, SS | Piece | √
| PLESTER sickle knife, Double edge, length 16cm, standard, medium curved, SS | Piece | √
| PLESTER Sickle knife, Double edge, length 16 cm, standard, slightly curved, SS | Piece | √
| PLESTERS sickle knife, Standard, medium curved, length 16 cm, SS | Piece | √
| POLITZER Ear Knife, Working length 16 cm, curved, SS | Piece | √
| Seeker with ball end, Angled 45 degree, ball end diameter 1mm | Piece | √
| SHEA Vein Press, Length 13 cm | Piece | √
| SPRATT Mastoid Curette, Length 17 - size 1, SS | Piece | √
| SPRATT Mastoid Curette, Length 17 - size 2, SS | Piece | √
| SPRATT Mastoid Curette, Length 17 - size 4, SS | Piece | √
| Suction cannula: Angular LUER lock outer diameter 0.5 mm, working length 6 cm | Piece | √
| Suction cannula: Angular LUER lock outer diameter 0.7 mm, working length 6 cm | Piece | √
| Suction cannula: Angular LUER lock outer diameter 1.0 mm, working length 6 cm | Piece | √
| Suction tube: Conical, LUER, outer diameter 3.5 mm, working length 13 cm | Piece | √
| Suction tube: Conical, LUER, outer diameter 4.5 mm, working length 13 cm | Piece | √
| Suction tube: Cylindrical LUER, outer diameter 0.5 mm, working length 9 cm | Piece | √
| Suction tube: Cylindrical LUER, outer diameter 0.7 mm, working length 9 cm | Piece | √
| Suction tube: Cylindrical LUER, outer diameter 1.0mm, working length 9 cm | Piece | √
| TRAUTMANN Ear Knife, Working length 19 cm, Delicate Lance, SS | Piece | √
| Tuning Fork C-512, SS | Piece | √ | √ | √
| WAGENER Ear Hook: 90 degree, blunt, SS | Piece | √ | √ | √
| WULLSTEIN ear forcep: Very fine, oval cupped jaws, working length 12.5 cm, diameter 1mm, SS | Piece | √
| WULLSTEIN ear forcep, Very fine, oval cupped jaws, working length 8 cm, diameter 0.9 mm, curved to left, SS | Piece | √
| WULLSTEIN ear forceps, Very fine, oval cupped jaws, working length 8 cm, diameter 0.9 mm, curved to right, SS | Piece | √
| WULLSTEIN ear forceps, Very fine, oval cupped jaws, working length 8 cm, diameter 0.9 mm, curved upwards, SS | Piece | √
| WULLSTEIN Needle, Slight curve, length 16.5 cm, SS | Piece | √
| WULLSTEIN Retractor, 3x3 prongs, length 13 cm, SS | Piece | √
| WULLSTEIN Scissor, Curved, sharp/blunt, length 14 cm, SS | Piece | √
| Anterior Pillar Retractor and tonsil dissector with tooth | Piece | √
| Bronchoscope Tube Universal, Length 43cm - size 6.5 | Piece | √
| Bronchoscope Tube Universal, Length 43cm - size 7.5 | Piece | √
| Bronchoscope Tube Universal, Length 43cm - size 8.5 | Piece | √
| Cotton Carrier, Length 35 cm | Piece | √
| Cotton Carrier , Length 55 cm | Piece | √
| Cotton Carrier , with spring handle length 35 cm | Piece | √
| Cotton Carrier, with spring handle length 55 cm | Piece | √
| DAVIS - MEYER Mouth Gag, SS | Piece | √
| Draffin Suspension bipod for use with tongue blade and mouth gags | Piece | √
| Esophagoscope Alligator, Grasping, double action jaws: Diam 2.5 mm working length 55 cm, SS, for use with oesophagoscope | Piece | √
| Esophagoscope circular biopsy : Diam 2.5 mm working length 55 cm, SS, for use with oesophagoscope | Piece | √
| Esophagoscope denture cutting forceps: Diameter 2.5 mm working length 55 cm, for use with oesophagoscope SS | Piece | √
| Esophagoscope peanut grasping forceps: Diam 2.5 mm working length 55 cm, for use with oesophagoscope | Piece | √
| Esophagoscope Tube, Length 20 cm - size 3 | Piece | √
| Esophagoscope Tube, Length 30 cm - size 4 | Piece | √
| Esophagoscope Tube, Length 30 cm - size 6 | Piece | √
| Esophagoscope Universal biopsy grasping forcep: Diam 2.0 mm Working length 55 cm, for use with oesophagoscope | Piece | √
| Esophagoscope Universal biopsy grasping forcep: Diam 2.0 mm Working length 35 cm, for use with laryngoscope | Piece | √
| Eve's tonsillar snare, Length 27 cm, SS | Piece | √
| Forceps circular cup, biopsy, working length 35 cm, for use with laryngoscope | Piece | √
| Forceps Alligator Grasping double action jaws, Diam 1.5 mm, working length35 cm, SS for use with laryngoscope | Piece | √
| Forceps Alligator Grasping double action jaws, Diam 2.0 mm, working length 55 cm, SS, for use with laryngoscope | Piece | √
| Forceps Peanut Grasping double action jaws, Diam 2.0 mm working length 35 cm, for use with laryngoscope | Piece | √
| Fuller's Trachea Tube with fenestration, Sterling silver - size 28 | Piece | √
| Fuller's Trachea Tube with fenestration, Sterling silver - size 30 | Piece | √
| Fuller's Trachea Tube with fenestration, Sterling silver - size 36 | Piece | √
| Fuller's Trachea Tube with fenestration, Sterling silver - size 32 | Piece | √
| Fuller's Trachea Tube with fenestration, Sterling silver - size 34 | Piece | √
| Halle Birkett Tonsillar artery forceps, curved, 20 cm, SS | Piece | √
| Insulated cannula for suction/coagulation, with ergonomic handle, insulated and fitting for unipolar coagulation, working length 23 cm, diameter 3.5 cm | Piece | √
| KLEINSASSER Cutting Forcep, curved left with 2 mm cupped jaws, working length 23 cm, SS | Piece | √
| KLEINSASSER Cutting Forcep, curved right with 2 mm cupped jaws, working length 23 cm, SS | Piece | √
| KLEINSASSER Cutting Forcep, straight, with 2 mm cupped jaws, working length 23 cm, SS | Piece | √
| Kleinsasser Laryngeal knife, oval straight, SS | Piece | √
| Kleinsasser laryngeal sickle knife, curved pointed, SS | Piece | √
| KLEINSASSER Operating Laryngoscope, ength 13 cm, child size, small | Piece | √
| KLEINSASSER Operating Laryngoscope, length 15 cm, child size | Piece | √
| KLEINSASSER Operating Laryngoscope, length 18 cm, adult size, medium | Piece | √
| KLEINSASSER Scissor, curved left working length 23 cm, SS | Piece | √
| KLEINSASSER Scissor, curved right working length 23 cm, SS | Piece | √
| KLEINSASSER Scissor , straight working length 23 cm, SS | Piece | √
| KLEINSASSER Scissor, Angled 45 degree, Working length 23 cm, SS | Piece | √
| Laryngeal Mirror with handle, size: 5 | Piece | √ | √ | √
| Laryngeal Mirror with handle, size: 6 | Piece | √ | √ | √
| Laryngeal Mirror with handle, size: 7 | Piece | √ | √ | √
| Laryngeal Mirror with handle, size: 8 | Piece | √ | √ | √
| Laryngeal Mirror with handle, size: 4 | Piece | √ | √ | √
| Laryngoforce II scissor, curved right, cutting blade 10 mm, working length 23 cm, SS | Piece | √
| Laryngoforce II scissor, left, cutting blade 10 mm, working length 23 cm, SS | Piece | √
| Laryngoforce II scissor, straight, cutting blade 10 mm, working length 23 cm, SS | Piece | √
| Laryngoscope holder and chest support | Piece | √
| Metzenbaum scissor, curved, length 18 cm, SS | Piece | √
| Negus Adenoid Curette, curved, length 21cm, with cage - size 2, SS | Piece | √
| Negus Adenoid Curette, curved, length 21cm, with cage - size 3, SS | Piece | √
| Negus Adenoid Curette, curved, length 21cm, with cage - size 4, SS | Piece | √
| Negus forcep, large curved, 20 cm, SS | Piece | √
| Negus ligature carrier | Piece | √
| Rigid suction tube with rubber tip, straight, working length 35 cm | Piece | √
| Rigid Suction Tubes with rubber tip, straight working length 55 cm | Piece | √
| ROBERTS - JESBERG Hypopharyngoscope, length 20 cm | Piece | √
| ROBERTS - JESBERG Oval Operating Esophagoscope, length 30 cm | Piece | √
| ROBERTS - JESBERG Oval Operating Esophagoscope, length 50 cm | Piece | √
| RUSSELL - DAVIS Tongue Blade with central groove, size: 1 | Piece | √
| RUSSELL - DAVIS Tongue Blade with central groove, size: 2, SS | Piece | √
| RUSSELL - DAVIS Tongue Blade with central groove, size: 3, SS | Piece | √
| RUSSELL - DAVIS Tongue Blade with central groove, size: 4, SS | Piece | √
| Tele-Laryngo-Pharyngoscope with HOPKINS, Lateral telescope 90 degree | Piece | √
| Tracheal dilator, two bladed, length 15 cm, SS | Piece | √
| Universal Biopsy and Grasping forcep, working length 35 cm, diameter 2 mm, for use with laryngoscope | Piece | √
| Universal Biopsy and Grasping forcep, working length 55 cm, diameter 2.5 mm, for use with oesophagoscope | Piece | √
| YANKAUER Suction Tube, length 29 cm | Piece | √
| COTTLE Metal Mallet, SS | Piece | √
| COTTLE scissor, 10.5 cm, curved, SS | Piece | √
| CRAIG Nasal Bone Forcep, straight, SS | Piece | √
| Frazier suction tube, with cut off hole and stylet, angled, working length 10 cm, outer diameter 2 mm | Piece | √
| Frazier suction tube, with cut off hole and stylet, angled, working length 10 cm, outer diameter 3 mm | Piece | √
| Freer Elevator sharp and blunt, length 20 cm, SS | Piece | √
| Freer's suction elevator, with stylet, length 19 cm | Piece | √
| Heyman nasal scissor, medium, standard model, length 9.5 cm, SS | Piece | √
| Joseph Elevator, slightly curved, sharp, length 17.5 cm, SS | Piece | √
| Joseph nasal saw, bayonet type, left, SS | Piece | √
| Joseph nasal saw, bayonet type, right, SS | Piece | √
| Joseph retractor, length 15.5 cm, SS | Piece | √
| JOSEPH scissor, 14 cm, curved, SS | Piece | √
| Kerrison ronguer, size: 2, SS | Piece | √
| Killian- Claus septal chisel, V - shaped cutting edge, length 16.5 cm, SS | Piece | √
| Killian- osteotome | Piece | √
| KILLIAN STRUYCKEN Nasal Speculum with set screw, length 15 cm, blade length - 65 mm, SS | Piece | √
| KILLIAN STRUYCKEN Nasal Speculum with set screw, length 15 cm, blade length 40 mm, SS | Piece | √
| KILLIAN STRUYCKEN Nasal Speculum with set screw, length 15 cm, blade length 90mm, SS | Piece | √
| Thudicum Nasal Speculum, SS | Piece | √ | √ | √
| Bipolar Coagulation Forceps, with suction canal, nasal, blunt, insulated, bayonet shaped, length 16 cm, SS | Piece | √
| Blakesley nasal forceps, 45degrees upturned, through cutting, width 3.5 mm, length 11 cm, SS | Piece | √
| Blakesley Rhinoforce II Ethmoid forceps, straight, size 1, length 17 cm, SS | Piece | √
| Blakesley Wilde nasal forcep, 90 degrees upturned, width 3.5 mm, length 11cm, SS | Piece | √
| Blakesley Wilde nasal forcep, straight, width 2.5 mm, length 11 cm, SS | Piece | √
| Blakesley Wilde Rhinoforce II, Ethmoid forceps 45 degree upwards curved, length 17 cm, SS | Piece | √
| Blakesley Wildenasal forceps, straight, width 4.8 mm, length 11 cm, SS | Piece | √
| Citelli antrum punch through cuttung, size: 1.5 mm, working length 17 cm, SS | Piece | √
| Eicken Antrum cannula, Luer lock, long curved, working length 11 cm, outer diameter 3 mm, SS | Piece | √
| Grunwald-Henke Nasal Cutting forceps, length 11 cm , SS | Piece | √
| Heuwieser antrum grasping forcep, fixed jaw curved 90 degrees, curved downwards, movable jaw backward opening 1200, jaw curved downwards, length 10 cm, SS | Piece | √
| Lichtwitz antrum trochar needle Luer lock, outer diameter 1.5 mm, working length 10 cm | Piece | √
| Nasal Sickle knife, pointed, length 19.5cm | Piece | √
| Probe, double ended, maxillary sinus ostium seeker, ball shaped, length 19cm, SS | Piece | √
| Rhinoforce Nasal scissors, curved to left, length 13 cm, SS | Piece | √
| Rhinoforce Nasal scissors, curved to right, length 13 cm, SS | Piece | √
| Rhinoforce Nasal scissors, straight, length 13 cm, SS | Piece | √
| Stammberger Antrum punch, backward cutting, sheath 360 degree rotatable with fixing screw, working length 10 cm with cleaning adaptor, SS | Piece | √
| Takahashi Nasal Forceps, straight, length 11cms, SS | Piece | √
| Trocar and cannula for Sinoscopy, O.D 3.3 mm, length of the cannula 7.5 cm, oblique beak, SS | Piece | √
| Trocar and cannula for Sinoscopy, O.D 4 mm, length of the cannula 7.5 cm, oblique beak, SS | Piece | √
| Trocar and cannula for Sinoscopy, O.D 5 mm, length of the cannula 8.5 cm, fenestrated beak, SS | Piece | √
| Bronchoscope tube universal, 4.5 mm | Piece | √
| Lack's tongue depressor | Piece | √ | √ | √ | √ | √
| Tilley's nasal dressing forceps | Piece | √ | √ | √ | √
| Noise generator device (tinnitus masker) | Piece | √
| Vancouver Optical Forceps for pediatric bronchoscopy, diameter 2.9mm, length 36cm, with spring action handle for controlled removal of peanuts and soft foreign bodies | Piece | √
| Bipolar Coagulation Forcep, 16 cm, pointed, SS | Piece | √ | √
| Stammberger Telescope handle, flat, standard model, length 11 cm for use with HOPKINS telescope 4 mm, length 18 cm | Piece | √
| Tumarkin metal aural speculum slotted, set of 8, Nos 0 to 7, black finish. | Piece | √
| Halle Birkett Tonsillar artery forceps, straight, 20 cm, SS | Piece | √
| Dennis Brown tonsil holding forceps | Piece | √
| Tonsillar scissors, curved | Piece | √
| Lacrimal punctum dilator | Piece | √
| Tilley Lichtwitz trocar and cannula for antral wash out | Piece | √
Accessories
| Fiber Optic light cable, size 3.5cm, | Piece | √
| Fiber Optic light cable, size 3.5cm, Length 230 cm | Piece | √
| Diamond straight shaft burrs with smooth shaft, length 7 cm, size 006 - 070, set of 15, SS | Piece | √
| Straight shaft burrs, Tungsten Carbide, length 7 cm, sizes 006 - 070, set of 14, SS | Piece | √
| Adaptors from Bronchoscope to any respiration equipment | Piece | √
| BENJAMIN -HAVAS Clip for proximal illumination | Piece | √
| Flexible Brush for Cytology Handle, Length 35 cm, SS | Piece | √
| Kleinsasser handle for use with above two instruments, SS | Piece | √
| Prismatic Light Deflector for use with laryngoscope | Piece | √
| Specimen collector for bronchoscopy | Piece | √
| Protecting Tube for HOPKINS telescopes | Piece | √
| Relevant cables and tubes, For Unidrive multifunction unit | Set | √
| Higginson syringe for antral wash out | Pcs | √
| Bulbs for operating microscope, Compatible to existing operating microscop | Piece | √
| Bulbs for Otoscope, Compatible to existing otoscope | Piece | √ | √ | √ | √
| Bulbs for Cold light source, Compatible to existing | Piece | √
| Polysomnography. Complete set. Sleep study |  | √
Hartmann ear forcep Alligator type (Crocodile forcep), Very fine, serrated, working length 12 cm, SS
Hartmann ear forcep Alligator type (Crocodile forcep), Very fine, serrated, working length 10 cm, SS
| Audiometer accessories head phone, TDH 39 |  | √
| Audiometer accessories BC vibrator, B71 |  | √
| Sterilizer dental Autoclave, Table top, Power supply 220-240v, 50Hz. Operating environmental condition: Temperature -10 to 40 ºC, Humidity > 80% , Suitable Voltage stabilizer having high/low cut off with delay timer, Power cord with round 3 pin top. |  | √ | √ | √
| Back biter for FESS surgey |  | √ | √
| Crocodile forcep 7cm long |  | √ | √ | √ | √
| Wax hook |  | √ | √ | √ | √
| 90 degree hook |  | √ | √ | √ | √
| Tillis forceps |  | √ | √ | √ | √
Gyne & Obst Department
Equipments
| Hysteroscope: Telescope 0 & 30 degree, autoclaveable, Telescope sheaths, Operating accessories (stainless steel); (Biopsy, forceps, scissors, graspers, coagulation electrode, colen's knief, roller ball - 6 Piece, Endovision display system - colour monitor (14"). Medical grade monitor with three chip camera, Xenon light source with six spare bulbs, Hysteromat: automatic infusion, suction device for infusing fluids in uterus from 0.1 to 3.0 litre/mintue. air bubble detection facility. Containers for collecting outflowing fluids, Recording facility for operative procedure. Servo voltage stabilizer should be supplied. Power Supply: AC 220-230V . Power cord with round 3 pin top. | Set | √ | √
| Bed for Obstetric and Labour : with Epoxy coated. I.V Rod (72x28x32"), with Raxine mattress, Adjustable height by hydraulic operated foot pump, kick bucket with ss bowl, leggings 2 nos, handrest 2 nos and with railing with adjustable head end and detachable lower part. | Set |  |  | √ | √
| Portable video colposcope: Class II, Camera for medical use. Pixel : 4 Million Pixel. CCD : 1/1.8" High Dnsity CDD. Image size : 2272 x 1704, 1024 x 68. Lens : 4 Times Zoom. Focus Distance : F=7.85~32.0mm (35mm Format calculation 38~155. Iris : F5.1 - F10.3. Lens composition : 8 Clusters 10 Pieces: 1. Digital Zoom: Maximum 4 Times. Autofocus : TTL AF. Shot Distance : 30 Cm~oo. Memory Tool : 16M Compact Flash Card. Sidelight Form : TTL. Shutter speed : 1 - 1/2300 Second. Interface : USB. Requester Program 1 copy. Power supply: AC 220V- 230V, Power cord with round 3 pin top and suitable voltage stabilizer. To be supplied with LCD monitor. | Set | √ | √ | √ |  | To be supplied in district hospitals where there are Gynecologists.
| CTG (cardiotocograph) machine twin monitor: with facility for fetal heart rate, uterine contraction recording in single and twin pregnancies. FHR to be displayed from 0-240 beats/minute, uterine contraction to be displayed from 0 - 100 mmHg '- Facility for maternal BP, ECG and Pulse oxymetry (optional) '- Facility for fetal pulse oximetry (optional)'- Ultrasonic transducers (four) with Piece machine '- Tocotransducers (two) with Piece machine '- Remote event marker '- Facility for alarm in case of very low and high fetal heart rate '- Facility for hoOK up to central nursing station (optional)'- Facility for hoOK up to central comptuer facility to store data on hard disk '- Thermal paper to last one year supply. CTG table.Power supply: AC 220V- 230V, to be supplied with two cardio probes. Power cord with round 3 pin top and suitable voltage stabilizer ( delete) | Set | √ | √ | √ |  | supplied to where there are gynecologist
| Doppler fetal heart monitor: with numerical display-2.5MHz, speaker 0.6W output, 50-280bpm, AC230V, Single phase, with rechargeable battery, DC9.6V-12V | Set | √ | √ | √ | √
| Laparoscopic set : with all accessories for laparoscopic surgery including family planning | Set | √ | √
| LEEP Machine: RSF - 4000K (220 - 240 volts). Mode - Cut, Cut & amp; Coag, Coag, Fulguration, Bipolar. Output voltage - Variable from 500 to 600 +_ 20%. Output power - 100 W. Output Frequency - Monopolar mode > 1 MHz. Bipolar mode > 1MHz. Mode & amp; output Waveform. Cut (Fully Filtered). Cut & amp; Coag (Fully Rectified). Coag (Partially Rectified). Fulguration (Spark - Gap). Blend. Class I type BF. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √ |  | Supplied to where there are gynecologist at DH
| Cryotherapy set with different sizes probes connection pipe to carbon dioxide cylinder with adaptor with accessories (cryoprobes for ectocervix) | Set | √ | √ | √ |  | Supplied to where there are gynecologist at DH
| Ultrasound Unit: Perinatal Medicine colour Doppler, 3.5 MHz probe 5-10 MHz probe, Transvaginal probe, Printer, Cart for the Unit, CINE Loop Facility, Clinically analysis package, Vascular package, OB/GYN package, General imaging package, In built 3D package (part of the OB/GYN package), DIACOM comparable for networking communication, 3D & Panoramic image display technology, Independent monitor tilt and swived adjustment. TRANSDUCER: 2 to 5 MHz broadbank curved array transducer for abdomen, OB.GYN application. 3 to 12 MHz broadband Linear transducer for vascular & Small parts application. 4 to 8 MHz broadband curved endovaginal application. Colour printer for inbuilt image management system. CD ROM recorder, USB, DVD. DIACOM media with MOD. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS. | Set | √
| Vaccum extractor for instrunmental delivery. with siliconized cups (40 mm, 50 mm, 60 mm) with electric motar and twin bottle. Complete standard set. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √
| Manual vacuum aspiration (MVA) kit, Disposable: Vaccum aspirator syringe of 50 ml volume with adaptors and suction cannulas of different sizes (2mm to 12mm size), Made of plastic | Set | √ | √ | √ |  | supplied to where there are trained doctors in BHU I level
| Delivery bed: Electrohydraulic operation. Lithotomy position with all standard accessories such a stirrup, clamp rods etc. Detacheable leg section, leg section should be split in two parts. Arm supports - one pair. Separate head section. Lateral tilt 15 - 30 degrees. Trendlenberg 30 degres. antitrendlenberg 20 degree - 30 degree. Protected against blood, liquor or waste sepage into machine. Hand clutches - one pair, to hold while bearing down during childbirth, Remote control, Detachable fluid collector ( according to bed size) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √
| Ultrasound machine (Portable) with only abdominal Probe : (Small ultrasound machine with gray scale monitor with abdominal probe only) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ |  |  | Only for impatient department
| Bed for Gyne Examination and procedures with leg stirrups for lithotomy position and motorised table with the capacity to hold upto 250kg of weight | Set | √ | √ | √ |  | supplied to where there are gynecologist at DH level
| Breast pumps (Electric) with variable pressure regulating device, Siliastic cups, on/off switch, reusable containers, silent fan and battery back-up.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Piece | √ | √ | √
| Baby cot cradle with mattress | Set | √ | √ | √ |  | Only where there are cesarean facility
Consumables
| CTG paper compatible with Philips Machine. | Piece | √ | √ | √ |  | where there is CTG machine and Gynae
| CTG paper compatible with Mastro Mediline Machine. | Piece | √ | √ | √ |  | where there is CTG machine and Gynae
| CTG paper compatible with CTG Huntleigh Machine. | Piece | √ | √ | √ |  | where there is CTG machine and Gynae
| Glossy printing paper compatible with ultrasound camera | Piece | √ | √
| Mattress for Obstetric & labour bed: compatible with the Obstetric & labour bed (72x28x32"). Removable 2 parts mattress with perineal cut out seat section. Polyurethane foam with reduction cover. (Measurement: Upper part with 136 cm in length, 87 cm in breadth and 18 cm thick. Lower Part with 87 cm length, 60 cm breadth and 18 cm thick.) | Piece | √ | √ | √ | √
| Mattress for Motorized Birthing Bed compatible with Janak Healthcare : Model No. Arrive 1512, Size : 2155mm L x 1010 mm W x 515mm to 880mm H. comprises of 2 pieces. (i) Body Section with concave in middle (ii) Foot Section with convex in the middle | Piece | √ | √
| Spatula for Pap smear: Sterile, Made of plastic, Shaped like allys bury spatula | Pkt of 100 | √ | √ | √ | √ | discuss with DoMSHI and RH for local production
| Cytobrush for Pap smear : | Piece | √ | √ | √ | √
| Silicone Ring pessaries of different sizes from 5 cm to 10 cm. For use in Uterine prolapse | Set | √ | √ | √ |  | where there is Gynae (No. of pieces in the set to be defined)
| Umblical cord clamp: ETO sterile pyrogen free security lock to prevent accident opening made from non-toxic medical grade polymer with solid grip for easy cord excision | Piece | √ | √ | √ | √
| Endometrial suction device/curette (Pipette): plastic and disposable | Pkt of 25 | √ | √
| Intrautrine Insemination (IUI) Cannulas / Catheter (For use in infertility treatment, to deposit enriched sperms into uterine cavity), length-23cm, external diameter 2.6mm, internal diameter 2.2mm, internal cannula diameter 1.1mm, catheter volume 0.35ml (cannula should be flexible) | Piece | √ | √
| Siliconized vacuum cup: Set of 40mm, 50mm, 60mm with tubings | Set | √ | √ | √
| Laminaria Tent (Seaweed, Hygroscopic material for Osmotic dialation of cervics) | Piece | √
| Lugol's Iodine (Potassium Iodide Solution for Colposcopy & LEEP) | Jar | √ | √ | √ |  | Where there are gynecologist
| Fischie's Clip for use in laproscopic tubal ligation | Packet | √ | √
Instrunments
| Amniotome (hooked), ss | Piece | √ | √ | √
| Cervical biopsy forceps (minibite), ss | Piece | √ | √ | √ |  | where there is colposcopy
| Cervical punch biopsy forceps, ss | Piece | √ | √ | √
| Dilator, Hegars, set of 8 ss | Set | √ | √ | √
| Doyen's myoma screw, ss | Piece | √ | √ | √ |  | where there is OT
| Doyen's retractors, size: Medium, ss | Piece | √ | √ | √ |  | where there is OT
| Endocervical biopsy curette,ss | Piece | √ | √ | √
| Forceps, Green-Armitage, straight,ss | Piece | √ | √ | √ |  | where there is OT
| Forceps, obstetric, Wrigley's,ss | Pair | √ | √ | √ |  | where there is Gynecologist.
| Forceps, ovum, size: Large , ss | Piece | √ | √ | √
| Forceps, ovum , size : Small, ss | Piece | √ | √ | √
| Forceps, tenaculum - single toothed,,ss | Piece | √ | √ | √ | √
| Forceps, vulsellum - multiple toothed ,ss | Piece | √ | √ | √
| Leech Wilkinson's cannula for HSG, size: Large , ss | Piece | √ | √ |  |  | where there is Radiologist
| Leech Wilkinson's cannula for HSG , size: Medium, ss | Piece | √ | √ |  |  | where there is Radiologist
| Leech Wilkinson's cannula for HSG , size: Small, ss | Piece | √ | √ |  |  | where there is Radiologist
| Scissor, episiotomy, Blunt tip, size: 6", ss | Piece | √ | √ | √ | √
| Speculum, cusco's, size: Large ss | Piece | √ | √ | √ | √
| Speculum, cusco's, size: Medium, ss | Piece | √ | √ | √ | √
| Speculum, cusco's, size: Small, ss | Piece | √ | √ | √ | √
| Speculum, sim's, size: Large, ss | Piece | √ | √ | √ | √
| Speculum, sim's, size : Medium, ss | Piece | √ | √ | √ | √
| Speculum, sim's, size: Small, ss | Piece | √ | √ | √ | √
| Uterine curette double ended, sharp/blunt with 14.5mm head size - Large, ss | Piece | √ | √ | √
| Uterine curette double ended, sharp/blunt with 11.5mm head size - Medium, ss | Piece | √ | √ | √
| Uterine curette double ended, sharp/blunt with 6.8mm head size - Small , ss | Piece | √ | √ | √
| Uterine sound , ss | Piece | √ | √ | √ | √
| Wound exploring probes, ss | Piece | √ | √ | √ |  | where there is OT
| Heaney hysterectomy forceps, ss | Piece | √ | √ | √ |  | where there is OT
| Kogan's endocervical speculum, ss | Piece | √ | √ | √ |  | where there is Colposcopy
| Richardson's Retractor, ss | Piece | √ | √
| Wartheim's clamps, size: Large , ss | Piece | √ | √
| Wartheim's clamps, size: Medium, ss | Piece | √ | √
| Scissor, Cord cutting, size: 6 inch, ss | Piece | √ | √ | √ | √
| CuT ( Intra Utrine Device: IUD) removable, hook | Piece | √ | √ | √ |  | where there are gynecologists
| Intra Utrine Device : (Copper "T"380) | Piece | √ | √ | √ | √
| Auvard's speculum for vaginal hysterectomy, ss | Piece | √ | √
| Shirodkar's uterine clamp, ss | Piece | √ | √
Leep Accessories
| LOOP Electrode for LEEP: size: W1 (15 mm), W2 (20 mm), size W3 (25 mm), Endocervical electrodes F2-8mm & Ball electrode 4 inch long W18 - 5mm | Set | √ | √ | √
| LEEP Patient cable (Hand held piece) | Piece | √ | √ | √
| LEEP Cautery (Patient) plate for Machine | Piece | √ | √ | √
| LEEP Smoke evacuation vaginal speculum : Large 835-5355 stamless | Piece | √ | √ | √
| LEEP Smoke evacuation vaginal speculum : standard 835-5355 stamless | Piece | √ | √ | √
| Ferric Subsulfate solution for hemostasis (Monsel), 500ml | Jar | √ | √ | √ |  | Required where there are LEEP machine
Accessories of Laparoscope
| Uterine Mannipulator | Piece | √ | √
Accessories of Hysteroscope
Ob/Gyn Operating Hysteroscopic parts ( Richard Wolf)
| A: 5.5mm Optiflow Hysteroscope sheath with 5 Fr working channel includes: 1 : Optiflow sheath-5.5mm 219mm 5 Fr. channel, Part # 8974.3511 2: Operative outer sheath-5mm, 205mm WL, Part # 8974.152 3: Operative inner sheath-205mm, WL, 7 Fr channel, Part # 8974.251 4: Operative sheath, 4mm, 245mm WL, 3 Fr channel, Part # 8974.121 5: Hysteroscope-30 ° 2.7mm, Part # 8974.402 | Set | √ | √
| B : Hand Equipments (Richard Wolf) |  | √ | √ | √ | √
| Biopsy Forcep: 3 Fr.410mm WL, Part # 829.131 | Piece | √ | √
| Micro Scissor: 5Fr.355mm WL, Part # 8642.681 | Piece | √ | √
| Grasping Forcep: 5 Fr.355mm WL, Part # 828.131 | Piece | √ | √
| Punch Biopsy Forcep : 5 Fr.340mm WL , Part # 8642.631 | Piece | √ | √
| C: Bipolar Electrode: (Richard Wolf) 1: BipoTrode electrode 5Fr. 36mm WL, Part # 823.005 2: ERBE Bipolar cable, 3M, Part # 8108.135 3: ConMed, Valley lab Bipolar cable, 3M 4: ERBE VIO 300D Bipolar generator | Set | √ | √
| D: Fluid Manager System: (Richard Wolf) -2222.223-complete system, includes Hystero pump II (2222.011), Fluid monitoring module (2223.011) - Inflow Tubing, disposable, sterile boc of 10x2 (4170.223) | Set | √ | √
| E: Princess 21Fr. Bipolar resectoscope: ( Richard Wolf) - 8653.018-19Fr. Continous flow inner sheath - 8653.021-21 Fr. Continuous flow outer sheath - 8653.18- Obturator | Set | √ | √
Accessories
| Probe for "Huntleigh "CTG machine: Model No. BD400XS-2, Serial No. 714X0201563-08 | Piece | √ | √ | √ |  | where there are CTG Machine
| Probe for "Mastro Mediline "CTG machine: Model No. NEWCGT12, Serial No. 1201006 | Piece | √ | √ | √ |  | where there are CTG Machine
| Transvaginal Probes compatible with Philips | Piece | √
| Transabominal Probes compatible with Philips | Piece | √
| Camera for Ultrasound Machine | Piece | √
| Camera for Colposcopy Machine compatible with Pro-Mis | Piece | √ | √
| Delivery Light : Ceiling mounted, Dome shadowless- Two light source | Piece | √ | √ |  |  | for laparoscopic tubal ligation
| Toco probe for CTG machine (Hiuntleigh) | Piece | √ | √ | √ |  | where there is CTG
| TVS probe with groove for biopsy needle compatable with existing machine. | piece | √ |  |  |  | for TVS guided pelvic biopsy
Internal Medicine Department (Medical Department)
Cardiology
| Echo cardiography (Echo Lab) machine; with Adult and Paediatric Probes, TEE Probe with printer (Black and white). Power supply 220-240v, 50Hz. Suitable UPS. Power cord with round 3 pin top | Set | √ | √
| Tread Mill Test (TMT) machine; Power supply 220-240v, 50Hz. Suitable Stabilizer, Power cord with round 3 pin top | Set | √
Consumables
| Pigtail catheter, Size: 4F | Piece | √
| Pigtail catheter, Size: 5F | Piece | √
| Pigtail catheter, Size: 6F | Piece | √
| Printing paper for Echo machine; Black and white | Roll | √ | √
| Short guide wire, J-tipped | Piece | √
| TMT printing paper | Roll | √
| Vascular Access sheath, with dilator, size: 4F | Piece | √
| Vascular Access sheath, with dilator, size: 5F | Piece | √
| Vascular Access sheath, with dilator, size: 6F | Piece | √
| Pericardiocentesis kit | Kit | √
Operation Thearter
Equipments
| Instrument Trolley: stainless steel (length -75 cm, breadth - 55 cm. Wheel size: 3 inch and railings on three sides) | Piece | √ | √ |  |  | where there is OT
| Mayo Table: Stainless steel structure, height adjustable from 85 cm to 120 cm, manually operated, stainless steel top tray detachable, 650x530x20mm. | Piece | √ | √ |  |  | where there is OT
| Mayo table (stain less steel): Dimension of the tray: 73-75 cm X47-50 cm X 6 cm Adjustable Height: 75-140 cm Weight: 10-15 Kgs Wheel: total three castors, two of them should have brake. | Piece | √ | √ |  |  | where there is ortho surgeon
| Operation table Radiolucent table top, power adjustable height, back, trendelenburg and lateral tilt. Mobile base with manually operated locks remote control facility for table movement operation. Manual removable split leg and head rest sections, pendent hand control, arm board with ball and socket joint adjustment, boay strap polyster fabric. Internal battery power and all accessories for general surgery, kidney bridge, leggings (1pair.) One extra remote control, list of essential spare parts and accessories with rates . 3 years warrenty. Power supply: 220V-230V, Power cord of round 3 pin round type plug top , operating condition:-10-+40 degree centrigrade, RH>80%, altitude>2000m | Set | √ | √ |  |  | where there is ortho surgeon
| LED Operation Theatre light Shadow less, 2 or more than 2 domes. Light handles, (In built cctv camera-optional) Can be used for hours without disruption, heat resistant. Should be fixed to ceiling.'- Back up UPS for 4 Hrs. 210-240 V, 50Hz. with extra 8 set of OT light handle.operating condition:-10-+40 degree centrigrade,RH>80%,altitude>2000m.3 pin round type plug top | Set | √ | √ |  |  | where there is ortho surgeon
| Surgical Diathermy machine with vessel sealing device: should contain 1: Monapolar- Foot switch, connecting cable, Hand piece. 2. Bipolar-Food switch, 10 numbers of bipolar connecting cables, 4 numbers of curved bipolar forceps and two numbers of straight bipolar forceps. 3. Vessel sealing handsets and accessories. | Set | √ | √ | √ |  | where there is ortho surgeon
| Operation Theatre Light handles: compatible to the existing OT light/machine | Piece | √ | √ |  |  | Where there is OT
| OT fumigation machine: Power supply: 220V-230V, Power cord of round 3 pin round type plug top | piece | √ | √ |  |  | Where there is functional OT
| Collin Rectractor (Adult): (24 Cm/10”) (38 X 60 mm) | piece | √ | √
| Collin Rectractor (Paed): (20 Cm/8”) (18 X40 mm) | piece | √ | √
| Cleaning brush, (210 X95 mm) | piece | √ | √
| Cleaning brush, (330 X 130 X50 mm) | piece | √ | √
| Cleaning brush, (400 X 190 X 8/12 mm) | piece | √ | √
Consumables
| Abdominal Sponges (Mops): 12 x 12 inch, thick layer, X-ray detectable | Piece | √ | √ |  |  | where there is functional OT
| Abdominal sponges (mops): 15 x 15 inch, x-ray detectable. | Piece | √ | √ |  |  | where there is functional OT
| Disposable surgical cap (OT caps) | Piece | √ | √ |  |  | where there is functional OT
| Silicon “T” tube: size FG-6 | Pkt. | √ | √ |  |  | where there is surgeon
| Silicon “T” tube: size FG-8 | Pkt. | √ | √ |  |  | where there is surgeon
| Silicon “T” tube: size FG-10 | Piece | √ | √ |  |  | where there is surgeon
| Silicon “T” tube: size FG-12 | Piece | √ | √ |  |  | where there is surgeon
| Medical grade suction tubings wide bore silicon: 4 meter long, with ties in the end for its connection with any type of cannula (Autoclavable) | Piece | √ | √ | √
| Prolene mesh: 15 cm x 15 cm | Piece | √ | √ |  |  | Where there is functional OT
| Prolene mesh: 6 cm x 11 cm | Piece | √ | √ |  |  | Where there is functional OT
| Prolene mesh: 7 cm x 12 cm | Piece | √ | √ |  |  | Where there is functional OT
| Close wound suction unit: size 8 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ |  |  |  | Added for ENT sugery
| Close wound suction unit: Size 10, (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ |  |  |  | Added for ENT sugery
| Close wound suction unit: size 12 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ | √
| Close wound suction unit: size 14 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ |  |  | Where there is surgeon
| Close wound suction unit: size 16 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ | √
| Close wound suction unit: size 18 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ | √
| Scalpel blades: No 10, packet of 100, SS | Pkt. | √ | √ | √ | √
| Scalpel blades: No 11, packet of 100, SS | Pkt. | √ | √ | √ | √
| Scalpel blades: No 15, packet of 100, SS | Pkt. | √ | √ | √
| Scalpel blades: No 20, packet of 100, SS | Pkt. | √ | √ | √
| Haemostatic Gel foam: (standard) 7 x 5 x 1 cm | Piece | √ | √ |  |  | Where there is functional OT
| Linear GI Stapler: (55 mm) | Piece | √ | √
| Linear GI Stapler Cartridge: (55 mm) | Piece | √ | √
| Linear GI Stapler: (75 mm) | Piece | √ | √
| Linear GI Stapler Cartridge: (75 mm) | Piece | √ | √
| Angular GI Stapler | Piece | √ | √
| Curved Intraluminal Stapler : with Adjustable Height Staples 33mm | Piece | √ | √
| Curved Intraluminal Stapler : with Adjustable Height Staples 25mm | Piece | √ | √
| Curved Intraluminal Stapler: with Adjustable Height Staples 21mm | Piece | √ | √
| Re- loadable Linear stapler: 30 mm | Piece | √
| Reload for Linear stapler: 30 mm | Piece | √
| Re- loadable Linear stapler: 60 mm | Piece | √
| Reload for Linear stapler: 60 mm | Piece | √
| Re- loadable Linear cutter: 75 mm | Piece | √
| Reload for Linear cutter: 75 mm | Piece | √
| Re- loadable Linear Cutter: 55 mm | Piece | √
| Reload for Linear cutter: 55 mm | Piece | √
| Round Stapler | Piece | √
| Stapler for skin: disposable, 35 mm | Piece | √ | √ |  |  | Where there is functional OT
| Skin stapler: 35 mm clip remover | Piece | √ | √ | √
| Suction nozzle Yangkuar plastic | Piece | √
| Oxidized Regenerated Cellulose: (Absorbable haemostat) 5 cm x 7.5 cm | Pkt. | √ | √ |  |  | Where there is surgeon
| Tube drain: size - 8, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 10, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 12, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 14, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 16, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 18, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 20, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Tube drain: size - 22, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
| Theatre Incise Drape: 56 cm x 84 cm | Piece | √ | √ |  |  | do
| Theatre Incise Drape: 45 cm x 28 cm | Piece | √ | √ |  |  | do
| Theatre Incise Drape: 40 cm x 42 cm | Piece | √ | √ |  |  | do
| Theatre Incise Drape: 30 cm x 28 cm | Piece | √ | √ | √
| Disposable Suprapubic catheter set: 16 Fr. | Set | √ | √ |  |  | General
| Disposable Suprapubic catheter set: 14 Fr. | Set | √ | √ | √ |  | General
| Disposable Suprapubic catheter set: 8 Fr. | Set | √ | √ | √ |  | General
| Carbon cautry tip cleaner: with x-ray detectable | Packet | √ | √ | √
| Disposable surgical gowns, small | Piece | √ | √ |  |  | Where there is functional OT
| Disposable surgical gowns, medium | Piece | √ | √ |  |  | do
| Disposable surgical gowns, large | Piece | √ | √ |  |  | do
| Bard Parker Blade handle No. 3, SS | Piece | √ | √ | √ | √
| Bard Parker Blade handle No. 4, SS | Piece | √ | √ | √
| Bard Parker Blade handle No. 7, SS | Piece | √ | √ | √
| Gloves sterile, Non-powdered, Non latex Size 6.5 | Pair | √ | √
| Gloves sterile, Non-powdered, Non latex Size 7 | Pair | √ | √
| Goves sterile, Non-powdered, Non latex Size 7.5 | Pair | √ | √
| Gloves sterile, Non-powdered, Non latex Size 8 | Pair | √ | √
| Patient Positioning Pads (Visco elastic polymer): a. Donut head pad with centre dish-adult, size: 20cmx7.5cmx4.5cm, pediatric size: 14cmx5.7cmx3.2 cm. b. Horseshoe head pad-adult, size: 20.5x7.5x7.5 cm, pediatric size: 14x5.7x3.2 cm. c. Dome shaped positioned semi circular design with flat bottom configuration 36x15x7.6 cm. d. Contoured chest roll, small: 30.5x16.13 cm, medium: 40.6x16x13 cm, large: 50.8x16x13 cm. | set | √
| Patient Positioning Pads (Visco elastic polymer): b. Horseshoe head pad-adult, size: 20.5x7.5x7.5 cm, pediatric size: 14x5.7x3.2 cm. | set | √
| Patient Positioning Pads (Visco elastic polymer): c. Dome shaped positioned semi circular design with flat bottom configuration 36x15x7.6 cm. | set | √
| Patient Positioning Pads (Visco elastic polymer): d. Contoured chest roll, small: 30.5x16.13 cm, medium: 40.6x16x13 cm, large: 50.8x16x13 cm. | set | √
5.3 | Instrument
1 | Cooley vascular clamp: Atraumatic grip, 17 cm, SS | Piece | √ | √
| Dissecting right angle forcep baby adson: size - 16 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
| Dissecting right angle forcep baby adson: size - 18 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
| Dissecting right angle forcep baby adson: size - 20 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
| Fistula Probe: 12.5 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
| Forceps, artery: curved, size - 12.5 cm, SS | Piece | √ | √ | √ | √
| Forceps, artery: curved, size - 15 cm, SS | Piece | √ | √ | √ | √
| Forceps, artery: curved, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, artery: mosquito curved, size - 10 cm , SS | Piece | √ | √ | √
| Forceps, artery: mosquito curved, size - 12.5 cm , SS | Piece | √ | √ | √
| Forceps, artery: mosquito curved, size - 10 cm , SS | Piece | √ | √ | √
| Forceps, artery: mosquito curved, size - 12.5 cm , SS | Piece | √ | √ | √
| Forceps, artery: straight, size - 15 cm, SS | Piece | √ | √ | √
| Forceps, artery: straight, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, artery: straight, size - 25 cm, SS | Piece | √ | √
| Forceps, dissecting: non-tooth, size - 10 cm, SS | Piece | √ | √ | √ | √
| Forceps, dissecting: non-tooth, size - 12.5 cm, SS | Piece | √ | √ | √ | √
| Forceps, dissecting: non-tooth, size - 15 cm, SS | Piece | √ | √ | √
| Forceps, dissecting: non-tooth, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, dissecting: toothed, size - 10 cm SS | Piece | √ | √ | √
| Forceps, dissecting: toothed, size - 12.5 cm SS | Piece | √ | √ | √
| Forceps, dissecting: toothed, size - 15 cm SS | Piece | √ | √ | √
| Forceps, dissecting: toothed, size - 20 cm SS | Piece | √ | √ | √
| Forceps, dissecting: toothed, size - 25 cm SS | Piece | √ | √ | √
| Forceps, Kocher's: curved, size - 15 cm, SS | Piece | √ | √ | √
| Forceps, Kocher's: curved, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, Kocher's: straight, size - 15 cm, SS | Piece | √ | √ | √
| Forceps, Kocher's: straight, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, sinus, size - 15 cm, SS | Piece | √ | √ | √
| Forceps, sinus, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, sponge holding, size - 15 cm, SS | Piece | √ | √ | √ | √
| Forceps, sponge holding, size - 20 cm, SS | Piece | √ | √ | √
| Forceps, sponge holding, size - 25cm, SS | Piece | √ | √ | √
| Forceps, tissue, Allis, size - 5 inch, SS | Piece | √ | √ | √
| Forceps, tissue, Allis, size -6 inch, SS | Piece | √ | √ | √
| Forceps, tissue, Allis, size - 7 inch, SS | Piece | √ | √ | √
| Forceps, tissue, Babcock, size - 5 inch, SS | Piece | √ | √ |  |  | where there is surgeon
| Forceps, tissue, Babcock, size - 6 inch, SS | Piece | √ | √ |  |  | do
| Forceps, tissue, Babcock, size - 7 inch, SS | Piece | √ | √ |  |  | do
| Intestinal clamp, crushing curved (payr), size - 8 inch, SS | Piece | √ | √ |  |  | where there is functional OT
| Intestinal clamp, crushing straight (payr), size - 8 inch, SS | Piece | √ | √ |  |  | do
| Intestinal clamp, non-crushing curved with non- traumatic serration, size 21.5 cm/8.5 inch, SS | Piece | √ | √ |  |  | do
| Intestinal clamp, non-crushing straight with non- traumatic serration, size: 21.5 cm/ 8.5 inch, SS | Piece | √ | √ |  |  | do
| Kidney stone forceps Randall (Sets of 4) | Set | √ | √
| Micro-adson (delicate dissecting forceps tooth), size: 12 cm, 4.75 inch, 1 x 2 teeth, SS | Piece | √ | √
| Micro-adson (non-tooth), size: 12 cm, 4 .75 inch, SS | Piece | √ | √
| Needle holder, size - 10 inch , SS | Piece | √ | √ | √
| Needle holder, size - 10 inch, Tungsten carbide | Piece | √ | √
| Needle holder, size - 4 inch, SS | Piece | √ | √ | √
| Needle holder, size - 4 inch, Tungsten carbide | Piece | √ | √
| Needle holder, size - 5 inch, SS | Piece | √ | √ | √
| Needle holder, size - 5 inch, Tungsten carbide | Piece | √ | √
| Needle holder, size - 6 inch, SS | Piece | √ | √ | √
| Needle holder, size - 6 inch, Tungsten carbide | Piece | √ | √
| Needle holder, size - 8 inch, SS | Piece | √ | √ | √ | √
| Needle holder, size - 8 inch, Tungsten carbide | Piece | √ | √
| Non traumatic bull dog clamp DeBakey, 3 sm, SS | Piece | √ | √
| Non traumatic bull dog clamp DeBakey, 8 cm, 3.125 inch, SS | Piece | √ | √
| Non traumatic bull dog clamps (curved), 6 cm, SS | Piece | √ | √
| Non traumatic bull dog clamps (curved), 6.7 cm, SS | Piece | √ | √
| Non traumatic bull dog clamps (curved), 8.5 cm, SS | Piece | √ | √
| Non traumatic bull dog clamps (straight), 6 cm, SS | Piece | √ | √
| Non traumatic bull dog clamps (straight), 6.7 cm, SS | Piece | √ | √
| Non traumatic bull dog clamps (straight), 8.5 sm, SS | Piece | √ | √
| Pelvis Retractor St. Marks, 29cm, 193 x 60 mm, SS | Piece | √ | √
| Retractor Deaver, 100 x 25mm (AxB), 30cm, 12 inch, SS | Piece | √ | √ |  |  | where there is functional OT
| Retractor Deaver, 100 x 50mm (AxB), 30cm, 12 inch, SS | Piece | √ | √ | √
| Retractor, langenback, 14cm, SS | Piece | √ | √ | √
| Retractor, langenback, 16cm, SS | Piece | √ | √ | √
| Retractor, langenback, 21cm, SS | Piece | √ | √ | √
| Retractor, Liver, SS | Piece | √ | √ | √
| Retractor, Morris , SS | Piece | √ | √ | √
| Safety pin (Instrument holder), SS, 13.5 cm for holding ring handle instruments | Piece | √ | √ | √
| Ring retractor set, Kirschner, size - 340 x 240 mm, SS | Piece | √ | √
| Scissors curved metzenbaum, 4 inch, SS | Piece | √ | √ | √
| Scissors curved metzenbaum, 5 inch carbide coated | Piece | √ | √ | √
| Scissors curved metzenbaum - 6 inch, SS | Piece | √ | √ | √ | √
| Scissors curved metzenbaum - 8 inch carbide coated | Piece | √ | √ | √
| Scissors curved metzenbaum - 10 inch carbide coated | Piece | √ | √ | √
| Scissors Mayo, Curved - 10 inch, SS | Piece | √ | √ | √
| Scissors Mayo, Curved - 6 inch, SS | Piece | √ | √ | √
| Scissors Straight metzenbaum -4 inch, SS | Piece | √ | √ | √
| Scissors Straight metzenbaum - 5 inch, SS | Piece | √ | √ | √
| Scissors Straight metzenbaum -6 inch, SS | Piece | √ | √ | √ | √
| Scissors Straight metzenbaum - 8 inch, SS | Piece | √ | √ |  |  | where there is surgeon
| Suture cutting scissors, 23 cm | Piece | √
| Dissecting scissors curved, 23 cm | Piece | √
| Self-retaining abdominal Retractor, Balfour - Adult, SS (100 x 35 mm) | Piece | √ | √ | √
| Self-retaining abdominal Retractor, Balfour - Child, SS (24 x 20 mm) | Piece | √ | √ | √
| Bookwalter retractor with instrument case (Consisting of accessories that can be used for advanced stomach, thoractomy, liver, gallbladder, spleen, Gynecology and other abdominal organs surgery) Horizontal Bar Post Coupling Medium Oval Ring 10.5 x 14.5" I.D. (26.7 x 36.8cm) Malleable Retractor Blade (2ea) 2'' x 6'' (51mm x 152mm) Malleable Retractor Blade 1.5'' x 6'' (38mm x 152mm) Harrington Retractor Blade Malleable Retractor Blade 3'' x 6'' (76mm x 152mm) Kelly Retractor Blade (2ea) 2'' x 3'' , 51mm x 76mm Kelly Retractor Blade 2'' x 4'' 51mm x 101mm Kelly Retractor Blade 2'' x 5'' 51mm x 127mm Kelly Retractor Blade 2'' x 6'' 51mm x 152mm Liver Retractor (Detail list of accessories should be clearly mentioned by the supplier during quotation) | set | √ |  |  |  | Added for onco sugery
| Self retaining laminectomy retractors, Beckman-Eaton, 32 cm long | Piece | √ | √ |  |  | do
| Self retaining retractor, GELPI , 17.5 cm long | Piece | √ | √ |  |  | do
| Self retaining retractor, GELPI, 18.5 cm long | Piece | √ | √ |  |  | do
| Skin HooKs - Large, SS | Piece | √ | √ | √
| Skin HooKs - Small, SS | Piece | √ | √ | √
| Thoracotomy Lung Rectractor - Adult, SS | Piece | √ | √
| Thoracotomy Lung Rectractor - Child, SS | Piece | √ | √
| Thoracotomy Rib Rectractor - Adult, SS | Piece | √ | √
| Thoracotomy Rib Rectractor - Child, SS | Piece | √ | √
| Towel Clamps, Backhaus with locking device, size - 11.5 cm, SS | Piece | √ | √ | √
| Towel Clamps, Backhaus, size – 11 cm | Piece | √ | √ | √
| Towel Clamps, Jones/Jones-Martin, 8.5 cm | Piece | √ | √ | √
| Varicose vein stripper with probes, SS | Set | √ | √ |  |  | where there is surgeon
| Humby's skin grafting handle | Piece | √ | √ | √
| Humby's skin grafting blade | Pkt. of 10 | √ | √ | √
| Self-retaining abdominal Retractor, Balfour - Adult, SS | Piece | √ | √ |  |  | where there is surgeon
| Self-retaining abdominal Retractor, Balfour - Child, SS | Piece | √ | √ |  |  | do
| Self-retaining abdominal Retractor, Balfour - Adult, SS | Piece | √ | √ |  |  | do
| Self-retaining abdominal Retractor, Balfour - Child, SS | Piece | √ | √ |  |  | do
| Self-retaining abdominal Retractor, Balfour - Adult, SS | Piece | √ | √ |  |  | do
| Self-retaining abdominal Retractor, Balfour - Child, SS | Piece | √ | √ |  |  | do
Accessories
| Disposable Bipolar hand piece: with connecting cable compatible with Cautery machine (Martin) | Piece | √ | √
| Disposable Bipolar hand piece: with connecting cable compatible with Cautery machine (ERRBE) | Piece | √ | √
| Disposable Monopolar hand piece: with connecting cable compatible with Cautery machine (Martin) | Piece | √ | √
| Disposable Diathermy plate: with connecting cable for children (compatible with Martin machine) | Piece | √ | √ | √
| Disposable Diathermy plate: with connecting cable for children (compatible with ERRBE machine) | Piece | √ | √ | √
| Disposable Diathermy plate: with connecting cable for adult (compatible with Martin machine) | Piece | √ | √ | √
| Disposable Diathermy plate: with connecting cable for adult (compatible with ERRBE machine) | Piece | √ | √ | √
| Long tip for the monopolar cautry hand piece: (compatible with ERBEE Cautry machine) | Piece | √ | √ | √
| Long tip for the monpolar cautry hand piece: (compatible with Martin Cautry machine) | Piece | √ | √ | √
New Items
| Povidone –Iodine 10% w/v (500 ml), for surgical skin preparation | 500mls | √
| Surface and equipment disinfectant spray. Aldehyde free (250 ml-500 ml with dispenser)For Electronic device, Laparoscopic screens, C-Arms, Monitors, emergency trolley’s | 250-500mls | √
| Disposable diathermy plate: For OR/ICU/Wards (H2O2 10% and silver nitrate 0.01%) | 500mls | √
| Surgical hand scrub solution: Chlorhexidine 5% V/V, Ethyl alcohol 61% W/V, with emollient and moisturizer) with dispenser | 500mls | √
| Surgical Hand disinfectant gel with dispenser: Bactericidal, tuberculocidal, Virucidal | 500mls | √
| Microfiber cloth (green): For cleaning operating equipments | piece | √
| Microfiber clot(red): For cleaning operating Table | piece | √
| Operation theatre cleaning cart: Single bag capacity, above 8 inch ball bearing wheel, sufficient storage space, interchangeable handles, lightweight, rubber tool grips and hooks, secure cabinet doors, water and chemical resistant), length 48-49 inch, width 21-22 inch, height 43-45 inch | piece | √
| Surgical gloves powdered: Sterile, size-6 | pair | √
| Surgical gloves powdered: Sterile, size-6.5 | pair | √
| Surgical gloves powdered: Sterile, size-7 | pair | √
| Surgical gloves powdered: Sterile, size 7.5 | pair | √
| Surgical gloves powdered: Sterile, size-8 | pair | √
| Disposable diathermy plate: For Force-Triad (Covidone) with connecting cable | piece | √
| Disposable diathermy plate: For Schiller diathermy Machine with connecting cables | piece | √
| Para formaldehyde tablets: For disinfection of endoscopic cables and accessories | piece | √
| Glutaraldehyde 2% with OPA: For high level disinfection | 5 ltrs | √
| Automatic Liquid Dispenser ( For Soap/disinfectant solution): Stainless steel or plastic | 250-500mls | √
Opthalmology
Equipments
| SnellenTrail Set - Non Illuminated, Power supply 220-240v, 50Hz.Power cord with round 3 pin top. | Piece | √ | √ | √
| Cabinet Roller Test Type Drum for 3 Meter Distance. Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √ | √
| Internally Illuminated Rotating Drum Distance Vision | Piece | √ | √ | √
| Near Vision Eye Testing Drum, Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √ | √
| Binomag Loupe With Adjustable Headband | Piece | √ | √ | √
| Slit Lamp - Haag Streight Slit Lamp, Power supply 220-240v, 50Hz. | Piece | √ | √
| Slit Lamp - Topcon SL D2, Voltage 100-120 / 200-240, Power 160VA, | Piece | √ | √ | √
| Lensometer | Piece | √ | √ | √
| Bipolar Cautery, Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √
| Ophthalmoscope Direct - 3.5V Coaxial 11720 | Piece | √ | √ | √
| Retinoscope - Streak - 18240 | Piece | √ | √ | √
| Indirect Ophthalmoscope - 12500 | Piece | √ | √
| Tonometer - Schiotz | Piece | √ | √ | √
| Tonopen Reichert | Piece | √ | √
| Perkins Tonometer | Piece | √ | √
| Plantech Refraction Unit with LCD screen and remote control | Set | √ | √
| AB Scan, Power supply 220-240v, 50Hz. | Set | √ | √
| Operating Microscope - portable - OPMI - 1 -FR, Power supply 220-240v, 50Hz. | Set | √ | √
| Operating Microscope ( CL-995D), The main scope adopts the technology of 5-step,3.5,5.4,8.5,13.5,21X, optical routeway and motorized zooming. The Image is clear, good stereoscopic sense and wide view field. The working distance is 200mm. Bi-oculars microscope has the angel between the main and bi-oculars is 90°. coaxial Cold light source illumination, and brilliance of the light spot:≥50000LX. Optional: Light distributor, cameral interface, single-ocular showing scope. | Set | √ | √
| Automated Humphrey Perimeter | Set | √ | √
| Digital Retinal Camera | Set | √ | √
| Clinical Specular Microscopy Non Robo - NSD 9900 | Set | √
| Eye Bank Kerato analyser Model No. EKA - 04 | Set | √
| YAG LASER | Set | √ | √
| Iridex ARGON Red LASER | Set | √ | √
| LASER Eye Safety Filter | Piece | √ | √
| Vitrectomy Machine, Power supply 220-240v, 50Hz. | set | √
| OCT Machine, Power supply 220-240v, 50Hz. | set | √
| Green Diode LASER For Photocoagulation, Power supply 220-240v, 50Hz. | set | √
| Lameller Flow Hood | Piece | √
| Pedriatric Trial Frame | Piece | √ | √ | √
| Keeler Trial Frame (Pediatric) | Piece | √ | √
| Paediatric opthalmoscope with alkaline battery | Piece | √
| Autoclave machine | Piece | √ | √ | √
| WelchAllyn wall mount Retinoscope set 3.5V | set | √ | √ | √
| LASER Saftety Glasses | Piece | √ | √
| Mainster Focal / Grid Retina Argon / Diode Laser Lens | Piece | √
| Oaiy Abraham Iridectomy Yag Leser Lens | Piece | √
| Oaiy Abraham Capsulotomy Yag Leser Lens | Piece | √
| VOLK Super Quad 160 Fundus Laser Lens | Piece | √
| VOLK Pan Retinal 2.2 Laser Lens (AE031590) | Piece | √
| Hertal Exophthalmometer | Piece | √ | √
| Vitrectomy irrigation lens, plano convex | Piece | √
| Vitrectomy irrigation lens, plano concave | Piece | √
| Ellman Cautery Radiofrency with Electroepilation Tips | Piece | √
| Eye Model for Teaching | Piece | √ | √ | √
| I Care Tonometer | Piece | √ | √
| Haag Streit Applanation Tonometry | Piece | √ | √
| Recgargeable battery 3.5Volts Catalog No.72200 made in Japan WELCH ALLYN | Piece | √ | √ | √
| Recgargeable battery 3.5Volts Catalog No.72300 made in Japan WELCH ALLYN | Piece | √ | √ | √
Vitreo-retinal Sub-speciality Clinic - Equipments, consumables & accessories
| Vitrectomy - Constellation Vision System 1.Online UPS, 2.Vitrectomy cassette, 3.Tubing accessories, 3.Tubing accessories, 3.Tubing accessories | Piece | √
| Optical Coherence Tomography - Cirrus HD 1.Online UPS, 2.Coulur Printer, 3.Printing Papers | Piece | √
| Fundus Camera 1.Colour Printer, 2.Vitrectomy cassette | Piece | √
| YAG 532 nm Laser Machine YAG 532 nm Laser Machine, 2.Squad lens - 150D, 3.Coupling media, 4.Eye safety spectacles, 5.Fibre optic pipe | Piece | √
| Argon Laser Machine 1.Smart key, 2.Eye safety spectacles | Piece | √
6 | Phaco-machine, Universal II 1.Tubing accessories, 2.Foldable IOLs, 3.IOL injector, 3.IOL injector | Piece | √
7 | Vitrectomy Machine | Piece | √
8 | Laser Machine. For Endolaser in operation theatre | Piece | √
9 | Laser Machine: Retinal laser therapies in OPD | Piece | √
10 | Laser safety glasses: Use for safety purpose during laser procedures | Piece | √
11 | OCT Machine: High defination scanning of retinal structures | Piece | √
12 | Digital Auto Lensmeter 7"Big touch screen Ophthal Display 7" TFT LCD 758x480 pixels PD UV with printer | Piece | √ | √
13 | Hand Held Keratometer Macro KM 500 with infra-red tansmission to printer and with charger | Set | √ | √
14 | Non-Mydriatic Auto Fundus Cemera AF-330 with color printer | Set | √ | √
15 | Digital Vision Chart, LCD 18.5 inch, Wall mount,Remote controlled, 3 meters, Display Resolution 250+-30cd/m2 | Piece | √ | √
16 | IOL Master 700 / Optical Biometry with printer | set | √ | √
17 | Slit Lamp Digital System SL-7D with color printer Slit Lamp Digital System Viewlight SL-7 Background Illumination adapter Beam Splitter Digital Adapter Camera Canon Rebel XSi (SLR 12.2 mpx) Software for Photo Management Canon Power adapter for Canon Rebel Memory Card 8 GB Electric Table Camera parameters are setup, so you don't have to figure out the correct settings. The electric table is shipped with the setup for the Slit Lamp SL-7. | set | √ | √
PHACO MACHINE
| Phaco machine (Alcon) Universal II | Piece | √
| Floridable Intra Ocular Lens | Piece | √
| Phaco accessories | Piece | √
| Nugent forcep, Scleral bucking | Piece | √
| Visulas Yag III Combi | Piece | √
| Volk Lens 90 Diopter | Piece | √ | √
| Volk Lens 20 Diopter | Piece | √ | √
| 3 Angle Mirrors | Piece | √ | √
INSTRUMENTS
1 | Graefe Muscle Hook | Piece | √ | √
2 | Alfonso Eye Speculum | Piece | √ | √
3 | Clark Eye Speclum - Adult size | Piece | √ | √ | √
4 | Clark Eye Speclum - Child size | Piece | √ | √ | √
5 | Sauer Eye Speculum Solid Blades - Infant size | Piece | √ | √
6 | Cottle elevator | Piece | √ | √
7 | Kay's scissors | Piece | √ | √
8 | Kimura Spatula | Piece | √ | √ | √
9 | Sauer Eye Speculum Solid Blades - Child size | Piece | √ | √
10 | Maumenee-Park eye speculum | Piece | √ | √
11 | Thudicum Nasal Speculum - Infant, | Piece | √ | √
12 | Thudicum Nasal Speculum - Child | Piece | √ | √
13 | Thudicum Nasal Speculum - Adult | Piece | √ | √
14 | Barraquer Eye Speculum (heavy) - Adult | Piece | √ | √ | √
15 | Barraquer Eye Speculum (heavy) - Child | Piece | √ | √ | √
16 | Desmarres Lid Retractor Infant - 7.5mm Wide | Piece | √ | √ | √
17 | Desmarres Lid Retractor - Adult - 13mm Wide | Piece | √ | √ | √
18 | Dastoor Iris Retractor | Piece | √ | √
19 | Knapp Lacrimal Sac Retractor | Piece | √ | √
20 | Killner Retractor for Occuloplasty | Piece | √ | √
21 | Jaeger - Lid Plate | Piece | √ | √
22 | Stallard Ptosis Plate | Piece | √ | √
23 | Tooks Corneal Knife | Piece | √
24 | Jameson Strabismus Hook | Piece | √
25 | Gass Retinal Detachment Hook - 13mm Flattened with 1.5mm Oval Hole at Tip | Piece | √
26 | Indo Gem Retinal Detachment Hook with 2mm Bulbos Tip with Groove Suture | Piece | √
27 | Sinskey Lens Manipulator | Piece | √ | √
28 | Dastoor Spatula for Pupil & Iris Repositor | Piece | √ | √
29 | Lang Lacrimal Sac Dissector | Piece | √ | √
30 | Kirby Lens Loop & Expressor | Piece | √ | √
31 | Wells Enucleation Spoon | Piece | √ | √
32 | Mule Evisceration Spoon | Piece | √ | √
33 | Meyerhoefer Chalazion Curette 2.0mm Cup | Piece | √ | √ | √
34 | Flieringa Scleral Fixation Ring - 12mm, | Piece | √
35 | Flieringa Scleral Fixation Ring - 14mm, | Piece | √
36 | Flieringa Scleral Fixation Ring - 16mm, | Piece | √
37 | Flieringa Scleral Fixation Ring - 18mm | Piece | √
38 | Hesshurg Barron Vacuum corneal trephine - 6mm | Piece | √
39 | Hesshurg Barron Vacuum corneal trephine - 6.50mm | Piece | √
40 | Hesshurg Barron Vacuum corneal trephine - 7.00mm | Piece | √
41 | Hesshurg Barron Vacuum corneal trephine - 7.50mm | Piece | √
42 | Hesshurg Barron Vacuum corneal trephine - 8.00mm | Piece | √
43 | Hesshurg Barron Vacuum corneal trephine - 8.50mm | Piece | √
44 | Hesshurg Barron Vacuum corneal trephine - 9.00mm | Piece | √
45 | Hesshurg Barron Vacuum corneal Punch - 6.00mm | Piece | √
46 | Hesshurg Barron Vacuum corneal Punch - 6.50mm | Piece | √
47 | Hesshurg Barron Vacuum corneal Punch - 7.00mm | Piece | √
48 | Hesshurg Barron Vacuum corneal Punch - 7.50mm | Piece | √
49 | Hesshurg Barron Vacuum corneal Punch - 8.00mm | Piece | √
50 | Hesshurg Barron Vacuum corneal Punch - 8.50mm | Piece | √
51 | Hesshurg Barron Vacuum corneal Punch - 9.00mm | Piece | √
52 | Hesshurg Barron Vacuum corneal Punch - 9.5 mm | Piece | √
53 | Corneal Transplant Trephine | Piece | √
54 | Flexible Ruler with Inches & mm Marking, Stainless Steel | Piece | √
55 | Castroveijo Caliper Straight | Piece | √ | √
56 | Bowman Lacrimal Probe Cylinder Stainless Steel, 00 & 0 | Piece | √ | √ | √
57 | Bowman Lacrimal Probe Cylinder Stainless Steel, 1 & 2, | Piece | √ | √ | √
58 | Bowman Lacrimal Probe Cylinder Stainless Steel, 3 & 4 | Piece | √ | √ | √
59 | Catalano Lacrimal Intubation Set, Stainless Steel Malleable Probes With Silicone Tubing | Piece | √ | √
60 | Nettle Ship Punctum Dilator | Piece | √ | √ | √
61 | West Bone Chisel | Piece | √ | √
62 | West Bone Gauge | Piece | √ | √
63 | Freer Periosteal Twin Elevator Double Ended | Piece | √ | √
64 | Kerrison Bone Nibbling Ronguer | Piece | √ | √
65 | Graefee Fixation Forceps-0.12mm | Piece | √ | √
66 | Elschnig Fixation Forceps | Piece | √ | √
67 | Dastoor Superior Rectus Forceps | Piece | √ | √
68 | Jaffe Tying Forceps - Straight, | Piece | √ | √
69 | Mcpherson Tying Forceps - Curved | Piece | √ | √
70 | Mcpherson Corneal Forceps | Piece | √ | √
71 | Utrta Capsulorexhis forceps | Piece | √ | √
72 | Lims Corneal Forceps | Piece | √ | √
73 | Moorfield Conjunctival Suture Forceps | Piece | √ | √
74 | Kuglen Hook [Iris Hook & Lens Manipulator - Push & Pull Model] | Piece | √ | √
75 | Barraquer Cilia Forceps | Piece | √ | √ | √
76 | Berke Ptosis Clamp | Piece | √ | √
77 | Snellen Entropion Clamp - Right | Piece | √ | √
78 | Snellen Entropion Clamp - Left | Piece | √ | √
79 | Lambert Chalazion Forcep -10mm, | Piece | √ | √ | √
80 | Lambert Chalazion Forcep -15mm | Piece | √ | √ | √
81 | Watzke Sleeve Spreading Forceps | Piece | √ | √
82 | Scleral Depressor | Piece | √ | √
83 | Metallic Retractor | Piece | √ | √
84 | Hartman Haemostatic Mosquito Forceps - Straight | Piece | √ | √
85 | Hartman Haemostatic Mosquito Forceps - Curved | Piece | √ | √
86 | Schaedel - Towel Clips | Piece | √ | √
87 | Diefenbach Serrefine Clamps | Piece | √ | √
88 | Allis Tissue Forceps | Piece | √ | √ | √
89 | Bonn Iris Scissor | Piece | √ | √
90 | Knapp Strabismus Scissor | Piece | √ | √
91 | Vannas Scissor Straight | Piece | √ | √
92 | Gills Welsh Vannas 11mm Angled Blade | Piece | √ | √
93 | Westcott Tenotomy Scissor Blunt Tips | Piece | √ | √
94 | Kalt Needle Holder Straight | Piece | √ | √
95 | Castroviejo Needle Holder Without Lock | Piece | √ | √
96 | Mcpherson Needle Holder Micro Straight 11mm Jaws | Piece | √ | √
97 | Barrquer Blade Braker & Holder | Piece | √ | √
98 | Lacrimal Cannula 25 Guage 1/2 Curved, SS | Piece | √ | √
99 | Simcoe Irrigating & Aspirating Cannula, SS | Piece | √ | √
100 | Drews Irrigating Vectis | Piece | √ | √
101 | Stainless Steel Sterilizing Case - 210x125x20mm | Piece | √ | √
102 | Stainless Steel Sterilizing Case - 330x220x20mm | Piece | √ | √
103 | Titanium Forceps Straight Tying | Piece | √ | √
104 | Titanium Forceps Straight Tooth | Piece | √ | √
105 | Titanium Mcpherson Forceps Angled | Piece | √ | √
106 | Titanium Needle Holder Barquer Straight | Piece | √ | √
107 | Castroviejo CorneoScleral Punch 1.5mm | Piece | √ | √
108 | Skin hooks - double prong | Piece | √ | √
109 | Paufique forceps | Piece | √ | √
110 | Stevens scissors - Curved | Piece | √ | √
111 | Jewellers forceps | Piece | √ | √
112 | Castro viejo needle holder with Lock | Piece | √ | √
113 | Westcott scissors - Sharp Tip | Piece | √ | √
114 | Westcott scissors - Blunt Tips | Piece | √ | √
115 | Malleable retractors sizes - 6X200mm, | Piece | √ | √
116 | Malleable retractors sizes - 8X200mm, | Piece | √ | √
117 | Malleable retractors sizes - 10X200mm, | Piece | √ | √
118 | Malleable retractors sizes - 12X200mm | Piece | √ | √
119 | Sewell's retractors | Piece | √ | √
120 | Sayer retractors | Piece | √ | √
121 | Stainless Stell Corneal Schields | Piece | √ | √
122 | Ball Implant Sizer | Piece | √ | √
123 | Ball Implant Introducer | Piece | √ | √
124 | Colorado tips for the cautery machine | Piece | √ | √
125 | Skin hooks- single and double prong | Piece | √ | √
126 | Cage & Jar (Moist Chamber For Enucleated Eye Storage) | Piece | √ | √
127 | Corneal Burrs Battery Operated | Piece | √ | √
128 | Groove Director For Lacrimal Intubation Tubes/Mickey Mouse | Piece | √ | √
129 | Implant Introducer - 20mm | Piece | √ | √
130 | Implant Introducer - 18mm | Piece | √ | √
131 | Implant Introducer - 16mm | Piece | √ | √
132 | Prism Bar- Horizontal + Vertical | Piece | √ | √
133 | Loose Prisms | Piece | √ | √
134 | Jackson Cross Cylinder 0.50D & 1.00D | Piece | √ | √ | √
135 | Lens Bar | Piece | √ | √
136 | Stand Magnifiers and Reading Magnifiers | Piece | √
137 | Illuminated Stand Magnifiers - 2.8X, | Piece | √
138 | Illuminated Stand Magnifiers - 3.0X, | Piece | √
139 | Illuminated Stand Magnifiers - 4.0X, | Piece | √
140 | Illuminated Stand Magnifiers - 5.0, | Piece | √
141 | Illuminated Stand Magnifiers - 6.0X, | Piece | √
142 | Illuminated Stand Magnifiers - 10.0X | Piece | √
143 | Illuminated Stand Magnifiers - 12.5X | Piece | √
144 | Hgher Additions (Prismatic Bino Comfort) - 1X, | Piece | √
145 | Hgher Additions (Prismatic Bino Comfort) - 1.25X, | Piece | √
146 | Hgher Additions (Prismatic Bino Comfort) - 1.5X, | Piece | √
147 | Hgher Additions (Prismatic Bino Comfort) - 2X, | Piece | √
148 | Hgher Additions (Prismatic Bino Comfort) - 2.5X | Piece | √
149 | Noves Bino - 1X, | Piece | √
150 | Noves Bino - 1.25X, | Piece | √
151 | Noves Bino - 1.5X, | Piece | √
152 | Noves Bino - 2.0X, | Piece | √
153 | Noves Bino - 2.5X | Piece | √
154 | Hand Held Magnifyer And Reading Magnifiers | Piece | √
155 | Illuminated Pocket magnifiers - 3X, | Piece | √
156 | Illuminated Pocket magnifiers - 3.5X, | Piece | √
157 | Illuminated Pocket magnifiers - 4X, | Piece | √
158 | Illuminated Pocket magnifiers - 5X, | Piece | √
159 | Illuminated Pocket magnifiers - 7X | Piece | √
160 | Illuminated Pocket magnifiers - 10X, | Piece | √
161 | Illuminated Pocket magnifiers - 12.5X | Piece | √
162 | Aplanatic Reading Magnifiers - 2.4X, | Piece | √
163 | Aplanatic Reading Magnifiers - 2.7X, | Piece | √
164 | Aplanatic Reading Magnifiers - 2.9X, | Piece | √
165 | Aplanatic Reading Magnifiers - 3.8X, | Piece | √
166 | Aplanatic Reading Magnifiers - 4.5X | Piece | √
167 | Aspheric Reading Magnifiers - 2.8X, | Piece | √
168 | Aspheric Reading Magnifiers - 3.0X, | Piece | √
169 | Aspheric Reading Magnifiers - 3.5X, | Piece | √
170 | Aspheric Reading Magnifiers - 4.0X, | Piece | √
171 | Aspheric Reading Magnifiers - 5.0X, | Piece | √
172 | Aspheric Reading Magnifiers - 6.0X | Piece | √
173 | End-Gripping Forceps - 1286 - B, 20G | Piece | √
174 | Serrated Jaw Forceps - 1286 - C, 20G | Piece | √
Microscissors
| Vertical Scissors - 1286 - E, 20G | Piece | √
| Straight Scissors - 1286 - J, 20G | Piece | √
ILM Forceps
| Eckardt End-Gripping Forceps - 1286 - I | Piece | √
| FB Forceps |  | √
| Avci FB Forceps - 2286 - I, 20G/0 | Piece | √
| Avci FB Forceps - 2286 - H, 17G/1 | Piece | √
Vitrectomy Lenses
| Flat Lens - 1284 - DW | Piece | √
| Biconcave Lens (For FAE) - 1284 - EW | Piece | √
| Wide Field Lens (For Central PO & Viterous) Sterior - 1284 - GW | Piece | √
| 20 Degree Prism Lens (For Posterior Periphery) - 1284 - HW | Piece | √
| 30 Degree Prism Lens (For The Posterior Periphery Beyond The Equator) - 1284 - IW | Piece | √
| Suturable vitrectomy lens ring (adult size) - 1284 - R | Piece | √
| Suturable vitrectomy lens ring (Pedriatric size) - 1284 - R | Piece | √
| Magnifying Lens - 1284 F | Piece | √
| Prismatic Biconcave Lens - 1284 - L | Piece | √
Reusable Vitrectomy Accessories
| Scleral Plug Forceps - 1277 | 3.5/4.0mm | √
Reusable Microforceps
| End-Gripping Forceps - 1286 - B06 | Piece | √
| Eckardt End-Gripping Forceps - 1286 - W06 | Piece | √
| Serrated Jaw Forceps - 1286 - C06 | Piece | √
| Vertical Scissors - 1286 - E06 | Piece | √
| Straight Scissors - 1286 - J06 | Piece | √
23G Backflush Instruments
| Backflush With Blunt Needle With Passive Aspiration - 1281 - AD06 | Piece | √
| Backflush With Blunt Needle With Active Aspiration - 1281 - BTD06 | Piece | √
ILM Procedures
| Diamond Dusted Sweeper - 1290 - DSS06, 23G/0 | Piece | √
| Soft Tipped Canula With Leur Lock Connector - 1272 - DS23, 23G/0 | Piece | √
| Spaide Soft Tipped Diamond Dusted Canula For Trimcinolone Removal - 1272 - SDT23, 23G/0 | Piece | √
| Retinal Magnetic Forceps, 23G/0 | Piece | √
| Kerrison Rongeur 1.50mm | Set | √ | √
| Kerrison Rongeur 2.00mm | Piece | √ | √
| Kerrison Rongeur 2.50mm | Piece | √ | √
| Kerrison Rongeur 3.00mm | Piece | √ | √
Consumables
1 | Sterile Microsurgical Blade (Cresent) 2mm Wide, Tip Angled | Piece | √ | √
2 | Sterile Microsurgical Blade (Keratome) 3.2mm, Tip Angled | Piece | √ | √
3 | Sterile Microsurgical Blade (Keratome) 2.8m,m Tip Angled | Piece | √ | √
4 | Sterile 15 Degree Knife Side Port | Piece | √ | √
5 | Injection Methyl Cellulose Intra Ocular (Visco Elastic) In Prefilled Syringes 2ml | Piece | √ | √
6 | Injection Hyalorinadase 1500iu (vials) | Piece | √ | √
7 | Optisol GS (Corneal Storage Media) | Piece | √
8 | MK Media [McCarey-Kaufman] | Piece | √
9 | Aurosling - L527 | Piece | √
10 | Aurolac 0275D Frontalis Sling | Piece | √
11 | Trypan Blue 0.06% | Piece | √ | √
12 | Carbachol Intra Ocular Solution 0.01% | Piece | √ | √
13 | Bandage Contact Lens | Piece | √ | √
14 | Punctum Plugs | Piece | √ | √
15 | Cyanocrilate Glue | Piece | √ | √
16 | Eye shields - Small, Medium, Large | Piece | √ | √
17 | Eye Drape - SMS Polypropylene fabric water repellant Sheet size: 100x70 cm‚ Incise area: 7x9 cm‚ Drainage pouch of transparent PE film | Piece | √ | √
18 | Incise Drape Small D1020 | Piece | √
19 | Silicone Band 240 (2.5X0.6X125.0mm) | Piece | √
20 | Silicone Tyre 276 | Piece | √
21 | Silicone Sponge (3.0X5.0X100mm) | Piece | √
22 | Octane (5ml) Perfluoro-n-octane | Piece | √
23 | Polydimethyl Siloxane (10ml) | mL | √
24 | Sulphur Hexafluride (SF6) | Piece | √
25 | Perfluoropropane (C3F8) | Piece | √
26 | 30GX1/2" (0.3X13mm) Needle | Piece | √
27 | Ringer Lactate (Bottle With Rubber Cork) | Piece | √
28 | Injection Bevacizumab | Piece | √
29 | Injection Fluorescein for FFA | Piece | √
30 | Injection Triamcinolone Acetonide. | Piece | √ | √
31 | Schimer's Test Strip | Piece | √ | √ | √
32 | Thermal Paper URK-700 Auto | Piece | √ | √
33 | Thermal Paper Ref No. 112 60 FDT, Zeiss Humphrey Systems | Piece | √ | √
34 | Thermal paper (AB Scan) 110mmX18m | Piece | √ | √
35 | Plastic Universal Eye Schield | Piece | √ | √
36 | Snellen Distance Vision E Chart | Piece | √ | √ | √
37 | Snellen Distance Vision A Chart | Piece | √ | √ | √
38 | Visual Acuity ETDRS Chart | Piece | √ | √
39 | Visual Acuity LogMAR Chart | Piece | √ | √
40 | Near Vision Chart | Piece | √ | √ | √
41 | Isihara Color Vision Chart | Piece | √ | √ | √
42 | Universal Conformers - 20mm, | Piece | √
43 | Universal Conformers - 22mm, | Piece | √
44 | Universal Conformers - 24mm | Piece | √
45 | Artificial Eyes Small sizes (Right Eye) | Piece | √ | √ | √
46 | Artificial Eyes Small sizes (Left Eye) | Piece | √ | √ | √
47 | Artificial Eyes Medium Sizes (Right Eye) | Piece | √ | √ | √
48 | Artificial Eyes Medium Sizes (left Eye) | Piece | √ | √ | √
49 | Artificial Eyes Big Sizes (Right Eye) | Piece | √ | √ | √
50 | Artificial Eyes Big Sizes (Left Eye) | Piece | √ | √ | √
51 | Eye Spehere Solid Type - 13mm, | Piece | √
52 | Eye Spehere Solid Type - 15mm, | Piece | √
53 | Eye Spehere Solid Type - 17mm, | Piece | √
54 | Castroveijo Implant Solid Plastic - 13mm, | Piece | √
55 | Castroveijo Implant Solid Plastic - 15mm, | Piece | √
56 | Castroveijo Implant Solid Plastic - 17mm | Piece | √
57 | Medpor imlants- sheet and ball implant | Piece | √
58 | Eye Spere Solid Plaque - 19mm, | Piece | √
59 | Eye Spere Solid Plaque - 17mm, | Piece | √
60 | Eye Spere Solid Plaque - 15mm, | Piece | √
61 | Eye Spere Solid Plaque - 13mm | Piece | √
62 | Suction Cannula | Piece | √ | √
63 | Cambridge Crowding Cards | Piece | √ | √
64 | Cambridge Single cards | Piece | √ | √
65 | TNO Test For Stereoscopic Vision | Piece | √ | √
66 | RAF Ruler | Piece | √ | √
67 | Portable Low Vision Chart | Piece | √ | √
68 | Armsler grid | Piece | √ | √
69 | Cardiff Picture Chart (Pedriatric) | Piece | √ | √
70 | Portable Love & Bailey Low Vision Chart | Piece | √ | √
71 | Infusion Canula 5.mm | Piece | √
72 | Infusion Canula 7.mm | Piece | √
73 | Infusion Canula | Piece | √
74 | 23G Endolaser Pipes, 23G/0 | Piece | √
75 | Trochar Set Disposable | Piece | √
76 | Bulbs - Retinoscope Bulb 08200u | Piece | √ | √ | √
77 | Bulbs - Ophthalmoscope 11730 - 04900u | Piece | √ | √ | √
78 | Bulbs - Indirect Ophthalmoscope Ophthalmoscope - 1200u | Piece | √ | √
79 | Bulbs - Pocket Ophthalmoscope - 3900u | Piece | √ | √ | √
80 | Bulbs - Penlite - 07600u | Piece | √ | √
81 | Bulbs for Topcon Slitlamp - SL-3F | Piece | √ | √
82 | Bulbs for Topcon Slitlamp - SL-D2 | Piece | √ | √
83 | Bulbs for Slitlamp Haag Streit - BQ-900 | Piece | √ | √
84 | Bulbs for Slitlamp Haag Streit - BM-901 | Piece | √ | √
85 | Air Cannula (RYCROFT)27Gx22mm,angle 45' ,4mm bend | Piece | √
86 | Constellation Vitrotectomy Cassette 25 Gauze | Piece | √
87 | Constellation Vitrotectomy Cassette 23 Gauze | Piece | √
88 | Constellation Vitrotectomy Cassette 20 Gauze | Piece | √
89 | Perfluorocarbon Liquids (PFCL) (heavy liquid) | Piece | √
90 | I Care probe for Tonometery | Piece | √ | √
PCIOL
1 | PC IOL-Single piece 13mm Length - +13.00D | Piece | √ | √
2 | PC IOL-Single piece 13mm Length - +14.00D | Piece | √ | √
3 | PC IOL-Single piece 13mm Length - +15.00D | Piece | √ | √
4 | PC IOL-Single piece 13mm Length - +16.00D | Piece | √ | √
5 | PC IOL-Single piece 13mm Length - +16.50D | Piece | √ | √
6 | PC IOL-Single piece 13mm Length - +17.00D | Piece | √ | √
7 | PC IOL-Single piece 13mm Length - +17.50D | Piece | √ | √
8 | PC IOL-Single piece 13mm Length - +18.00D | Piece | √ | √
9 | PC IOL-Single piece 13mm Length - +18.50D | Piece | √ | √
10 | PC IOL-Single piece 13mm Length-+19.00D | Piece | √ | √
11 | PC IOL-Single piece 13mm Length-+19.50D | Piece | √ | √
12 | PC IOL-Single piece 13mm Length-+20.00D | Piece | √ | √
13 | PC IOL-Single piece 13mm Length-+20.50D | Piece | √ | √
14 | PC IOL-Single piece 13mm Length-+21.00D | Piece | √ | √
15 | PC IOL-Single piece 13mm Length-+21.50D | Piece | √ | √
16 | PC IOL-Single piece 13mm Length-+22.00D | Piece | √ | √
17 | PC IOL-Single piece 13mm Length-+22.50D | Piece | √ | √
18 | PC IOL-Single piece 13mm Length-+23.00D | Piece | √ | √
19 | PC IOL-Single piece 13mm Length-+23.50D | Piece | √ | √
20 | PC IOL-Single piece 13mm Length-+24.00D | Piece | √ | √
21 | PC IOL-Single piece 13mm Length-+24.50D | Piece | √ | √
22 | PC IOL-Single piece 13mm Length-+25.00D | Piece | √ | √
23 | PC IOL-Single piece 13mm Length-+25.50D | Piece | √ | √
24 | PC IOL-Single piece 13mm Length-+26.00D | Piece | √ | √
25 | PC IOL-Single piece 13mm Length-+27.00D | Piece | √ | √
26 | PC IOL-Single piece 13mm Length-+28.00D | Piece | √ | √
27 | PC IOL-Single piece 13mm Length-+29.00D | Piece | √ | √
28 | PC IOL-Single piece 13mm Length-+30.00D | Piece | √ | √
29 | PC IOL-Single piece 13mm Length-+31.00D | Piece | √ | √
30 | PC IOL-Single piece 13mm Length-+32.00D | Piece | √ | √
31 | PC IOL-Single piece 13mm Length-+33.00D | Piece | √ | √
32 | PC IOL-Single piece 13mm Length-+34.00D | Piece | √ | √
33 | PC IOL-Single piece 13mm Length-+35.00D | Piece | √ | √
AC IOL
| AC IOL-Single piece 13mm Length - +13.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +14.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +15.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +16.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +17.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +18.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +19.00D | Piece | √ | √
| AC IOL-Single piece 13mm Length - +20.00D | Piece | √ | √
PC Foldable Len
| Foldable IOL dioptre size 19.00D | Piece | √
| Foldable IOL dioptre size 19.50D | Piece | √
| Foldable IOL dioptre size 20.00D | Piece | √
| Foldable IOL dioptre size 20.50D | Piece | √
| Foldable IOL dioptre size 21.00D | Piece | √
| Foldable IOL dioptre size 21.50D | Piece | √
| Foldable IOL dioptre size 22.00D | Piece | √
| Foldable IOL dioptre size 22.50D | Piece | √
| Foldable IOL dioptre size 23.00D | Piece | √
| Foldable IOL dioptre size 23.50D | Piece | √
| Foldable IOL dioptre size 24.00D | Piece | √
Orthopedic Department
Basic Orthopaedic Supplies:
Plaster of Paris
1 | Paster of Paris: Size 7.5cmx 2.7m | Roll | √ | √ | √
2 | Paster of Paris: Size 10cm x 2.7m | Roll | √ | √ | √
3 | Paster of Paris: Size 15cm x 2.7m | Roll | √ | √ | √
4 | Fiber glass cast: Size- 10cm x 2.7m | Roll | √ | √
5 | Fiber glass cast: Size- 7.5cm x 2.7m | Roll | √ | √
Cotton cast padding (soft rolls) :
6 | Cotton cast padding (soft rolls) : Size 10cm x 3m | Roll | √ | √ | √
7 | Cotton cast padding (soft rolls) : Size 15cm x3m | Roll | √ | √ | √
Stokinette Rolls:
8 | Stokinette Rolls: Size 5cmx 10m | Roll | √ | √
9 | Stokinette Rolls: Size 7.5cm x 10m | Roll | √ | √
10 | Stokinette Rolls: Size 10cm x 10m | Roll | √ | √
Dressing Materials for Advanced Wound Mangt.
11 | Knited elastic fixation bandage (Tensofix): Size: 10 x 2.4m unstretched and 4m stretched | Roll | √ | √
| Adhesive incise drape, eg. OPSITE Incise, 30 x 28cm | 10/Pack | √ | √
| Adhesive incise drape, eg. OPSITE Incise, 45 x 55cm | 10/Pack | √ | √
| Adhesive film dressing with absorbent pad, eg. OPSITE Post-Op, 15.5 x 8.5cm | 20/Box | √ | √
| Adhesive film dressing with absorbent pad, eg. OPSITE Post-Op, 25 x 10cm | 20/Box | √ | √
| Absorbent dressing with adhesive border, eg. PRIMAPORE, 15 x 8cm | 20/Box | √ | √
| Absorbent dressing with adhesive border, eg. PRIMAPORE, 25x10cm | 20/Box | √ | √
| Hydrocellularantimicrbial foam dressings, eg. Allevyn Ag Gentle Border, 12.5 x 12.5cm | 10/Pack | √ | √
| Hydrocellularantimicrbial foam dressings, eg. Allevyn Ag Gentle Border, 17.5 x 17.5cm | 10/Pack | √ | √
| Hydrofiber Dressing, eg. Aquacel Ag Surgical, 9 x 15cm | 10/Pack
| Hydrofiber Dressing, eg. Aquacel Ag Surgical, 9 x 25cm | 10/Pack
| Hydrofiber Dressing, eg. Aquacel Foam Adhesive, 17.5 x 17.5cm | 10/Pack
| Hydrofiber Dressing, eg. Aquacel Foam Sacral, 20 x 16.9cm | 5/Pack
Negative Pressure Wound Therapy (NPWT) System
| Closed Cyclic NPWT PU Wound Management device | No | √ | √
Disposable supplies:
| CCNPWT canister (500ml) with Bacteria filter, Y Connector and Drainage tube- | No | √ | √
| CCNPWT PU Wound Drainage Kit- Small/7.5 x 10 x 3cm | No | √ | √
| CCNPWT PU Wound Drainage Kit- Medium/12.5 x 18 x 3cm | No | √ | √
| CCNPWT PU Wound Drainage Kit- Large/15 x 26 x 3cm | No | √ | √
| Bacteriostatic & Fungistatic Dressings- small/medium/large | No | √ | √
Abdominal Wound Kit
| Splints, Braces & Accessories | No | √ | √
| Universal Buck's skin traction splint | No | √ | √ | √
| Leg Traction Brace with velcro closure: Size: small | No | √ | √ | √ | √
| Leg Traction Brace with velcro closure: Size: medium | No | √ | √ | √ | √
| Leg Traction Brace with velcro closure: Size: large | No | √ | √ | √ | √
| Thomas splint adult size | No |  |  | √ | √
| Thomas splint child size | No |  |  | √ | √
| Traction Splints: Hare / Sagar traction splint: universal size | No | √ | √
| Crammer Wire Splints: Size - 24''x2" | No |  |  | √ | √
| Crammer Wire Splints: Size - 36"x3" | No |  |  | √ | √
| Crammer Wire Splints: Size - 36"x4" | No |  |  | √ | √
| Cervical orthosis (Philadelphia): small | No | √ | √
| Cervical orthosis (Philadelphia): medium | No | √ | √
| Cervical orthosis (Philadelphia): large | No | √ | √
| Spine Board and Strapping Systems-eg BAXSTRAP TM | No | √ | √
| Jewitt's Brace or Ash Brace: medium | No | √ | √
| Jewitt's Brace or Ash Brace: Large | No | √ | √
| Taylor's Brace: small | No | √ | √
| Taylor's Brace: medium | No | √ | √
| Taylor's Brace: Large | No | √ | √
| Knee Brace long- medium | No | √ | √
| Knee Brace long- large | No | √ | √
Steinmann Pins:
| Steinmann Pins, with trocar tip, 3.5mm dia., length 150mm | No | √ | √
| Steinmann Pins, with trocar tip, 3.5mm dia., length 200mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size - 4.0mm x 175mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size - 4.0mm x 200mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size - 4.0mm x 225mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size -4.5mmx250mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size- 5.0mm x 250mm | No | √ | √
| Bone Hand Drill : Bunnell- Hand drill with three-jaw chuck & key for shanks up to 4 mm diameter,cannulated | No | √ | √ | √ |  | Where there are ortho technicians
| Spica Cast Table (Paediatric) | No | √ | √
Basic Equipment for Orthopaedic Surgery
| Oscillating Plaster Saw, 220V/50Hz with the following parts: a. Two Spanners to exchange the saw blades; b. Spare saw blade (dia. 50mm); c. Spare saw blade (dia. 65mm); c. Spare saw blade (dia. 65mm); d. Spare saw blade, special hardness, (dia. 50mm); e. Spare saw blade, special hardness, (dia. 65mm);f. Plaster Spreader, 22. 5cm, 9" | Set | √ | √ | √ |  | For district hospitals wherever there are orthopaedic Technicians
| Plaster Spreader, 22. 5cm, 9" | No | √ | √ | √ |  | Same as above
| Plaster-of- Paris Shears (Plaster cutting scissors): Bruns, 24 cm, 9 1/2" | No | √ | √ | √ |  | Same as above
| Arm Machine : C-arm high frequency mobile X-ray Machine Suitable for the operating room and emergency room for real time guidance of orthopaedic operations. | Set | √ | √
| Electric / battery Powered Bone Drills | [5 empty] | Will be needed in Gelephu in 2017
| Hand Pieces: Duel Trigger Rotatory Hand Piece | No | √ | √
| Sagittal Saw Hand Piece | No | √ | √
Drilling and Reaming Attachments consisting of:
| Small Synthes Quick Connect | No | √ | √
| Stryker 1/4' Adjustable Jacob's Chuck ( keyed) | No | √ | √
| Synthes' Reaming Attachment (Large) | No | √ | √
| Synthes' DHS Reaming Attachment | No | √ | √
| Wire Collet (0.7mm -2.8mm) | No | √ | √
| Pin Collect (2.0mm-4.0mm) | No | √ | √
Accessories consisting of:
| Battery Charger | No | √ | √
| Batteries (3 pairs) | No | √ | √
| Aseptic Battery Kit | No | √ | √
| Sterilization Case for 3 hand Pieces | No | √ | √
| Saw Blades all sizes (3 pairs) | No | √ | √
Pneumatic Tourniquet System:
| Zimmer or Stryker equivalent doul channel tourniquet | Set | √ | √
| Tourniquet Cuff :Paediatric size | No | √ | √
| Tourniquet Cuff : size 12 inches | No | √ | √
| Tourniquet Cuff : size 18 inches | No | √ | √
| Tourniquet Cuff : size 24 inches | No | √ | √
| Tourniquet Cuff : size 30 inches | No | √ | √
| Tourniquet Cuff : size 34 inches | No | √ | √
Eshmarch :
| Eshmarch : 5m long and 7.5cm wide | No | √ | √
| Eshmarch : 5m long 10cm wide | No | √ | √
| Eshmarch : 5m long and 15cm wide | No | √ | √
| Size:LxWxH= 480mmx260mmx5mm | No | √ | √
General Orthopaedic Instruments
Bone Cutters:
| Ruskin Rongeurs, Straight, 7.5 inches | No | √ | √
| Ruskin Rongeurs, Curved, 7.5 inches | No | √ | √
| Ruskin Rongeurs, Straight, 9.0 inches | No | √ | √
| Ruskin Rongeurs, Curved, 9.0 inches | No | √ | √
| Ruskin-Liston Bone Cutting Forceps, Straight | No | √ | √
| Ruskin-Liston Bone Cutting Forceps, Angled | No | √ | √
Bone Mallets:
| Bone Mallets, 250gms | No | √ | √
| Bone Mallets, 400gms | No | √ | √
Osteotomes, Chisels and Gouges:
| Mini Osteotomes, 6 mm | No | √ | √
| Mini Osteotomes, 8 mm | No | √ | √
| Mini Osteotomes, 10 mm | No | √ | √
Bone Curettes:
| Endoscopic Curettes, Ring Curette | No | √ | √
| Endoscopic Curettes, Oval Cup Curette | No | √ | √
Bone Holding Instruments:
| Lane Bone Holding Forceps, 250mm, with ratchet | No | √ | √
| Lane Bone Holding Forceps, 300mm, with ratchet | No | √ | √
| Pelvic Reduction Forceps, short tips, angular, short pointed-ball tip | No | √ | √
| Pelvic Reduction Forceps, long tips, angular, long pointed-ball tips | No | √ | √
| Bone Reduction Forceps, with rachet, serrated, length 140mm | No | √ | √
| Bone Reduction Forceps, with rachet, serrated, length 170mm | No | √ | √
| Bone Reduction Forceps, with rachet, serrated, length 200mm | No | √ | √
| Bone Reduction Forceps, with rachet, pointed, length 140mm | No | √ | √
| Bone Reduction Forceps, with rachet, pointed, length 170mm | No | √ | √
| Bone Reduction Forceps, with rachet, pointed, length 200mm | No | √ | √
| Plate and Bone holding forceps, one pointed tip and foot plate | No | √ | √
Elevators, Awls and Bone Hooks:
| Periosteal Elevators, 6mm straight edge, straight blade | No | √ | √
| Periosteal Elevators, 13mm straight edge, straight blade | No | √ | √
Retractors:
| Homann Retractors, 8mm blade width, length 8.5 inches | No | √ | √
| Blount Retractors, double prong retractor | No | √ | √
| Humeral Head Retractor | No | √ | √
| Volkman Retractors, 2 prongs, sharp | No | √ | √
| Volkman Retractors, 3 prongs, sharp | No | √ | √
| Volkman Retractors, 4 prongs, sharp | No | √ | √
| Volkman Retractors, 2 prongs, blungt | No | √ | √
| Volkman Retractors, 3 prongs, blunt | No | √ | √
| Volkman Retractors, 4 prongs, blunt | No | √ | √
| Skin Hooks, single pronged | No | √ | √
| Skin Hooks, two pronged | No | √ | √
| Operating Scissors: Kelly 16cm, 6 1/4" | No | √ | √
| Operating Scissors: Mayo 14.5cm, 5 3/4" (straight) | No | √ | √
| Operating Scissors: Mayo 14.5cm, 5 3/4" (curved) | No | √ | √
| Operating Scissors: Mayo 17cm, 6 3/4" | No | √ | √
| Operating Scissors: Metzenbaum 11.5cm, 4 1/2" (straight) | No | √ | √
| Operating Scissors: Metzenbaum 11.5cm, 4 1/2" (curved) | No | √ | √
| Operating Scissors: Metzenbaum 14.5cm, 5 3/4" (curved) | No | √ | √
| Operating Scissors: Metzenbaum 18cm, 7" (curved) | No | √ | √
| Needle Holders : Halsey 13cm, 5 1/8" | No | √ | √
| Needle Holders : Crile-Wood 15cm, 6" | No | √ | √
| Needle Holders : Crile-Wood 23cm, 9" | No | √ | √
| Needle Holders : Mayo Hegar 16cm, 6 1/4" | No | √ | √
| Needle Holders : Mayo-Hegar 18.5cm, 7" | No | √ | √
Needle Holders : Mayo-Hegar 27cm, 10 3/4"
| Needle Holders : Kilner 13cm, 5 1/6" | No | √ | √
| Thumb Tissue Forceps: Adson, toothed, 12cm | No | √ | √
| Thumb Tissue Forceps: Adson, toothed, 15cm | No | √ | √
| Thumb Tissue Forceps: Adson-Brown, toothed, 12cm | No | √ | √
| Bone Files : Size- 22cm | No | √ | √
| Wire Instruments along with a Sterilizing tray for wire instruments consisting of : a. Wire passer, 70mm bending diameter; b. Wire holding forceps for circlage wires, L 170mm, c. Vice Grip, self retaining, L 180mm, d. Paralel pliers, flat nosed, L 160mm, e. Wire bending pliers, L 155mm , f.Wire cutter, small, L 160mm, g. Wire cutter, large, L 220mm, h.Wire Cutters, 48cm for hard wires/pins up to 6mm diam | No | √ | √
| Circlage Wires in coils of length 10m: Diameter 0.6 mm | No | √ | √
| Circlage Wires in coils of length 10m: Diameter 0.8mm | No | √ | √
| Circlage Wires in coils of length 10m: Diameter 1.0mm | No | √ | √
| Kirscher Wires with trocar tip: 1.0 mm dia. Lengths 150mm | No | √ | √
| Kirscher Wires with trocar tip: 1.25mm dia., lengths 150mm | No | √ | √
| Kirscher Wires with trocar tip: 1.6mm dia., lengths 150mm | No | √ | √
| Kirscher Wires with trocar tip: 2.0mm dia., lengths 150mm | No | √ | √
Kirscher Wires with trocar tip:2.5 mm dia., lengths 230mm
| Kirscher Wires with trocar tip:2.5 mm dia., lengths 280m | No | √ | √
| Kirscher Wires with threaded tip: 1.25mm dia., length 150mm | No | √ | √
| Kirscher Wires with threaded tip: 2.0mm dia., length 150mm | No | √ | √
| Kirscher Wires with threaded tip: 2.5mm dia., length 230mm | No | √ | √
| Kirscher Wires with threaded tip: 3.0mm dia., length 280mm | No | √ | √
TRAUMATOLOGY & GENERAL OSTEOSYNTHESIS (Synthes/Zimmer or Equivalent quality products)
Small Fragment Instrument & Implant Set- LC-DCP
| Small Fragment Instrument & Implant Set- LC-DCP Graphic Case | No | √ | √
| Small Fragment Instrument Set, st. steel (Synthes/Zimmer or equivalent) | Complete Set | √ | √
Non-Locking screws, 3.5mm diameters
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 14mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 16mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 18mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 20mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 22mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 14mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 16mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 18mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 20mm | No | √ | √
DCP & LC- Dynamic Compression Plate 3.5 mm
| LC- Dynamic Compression Plate 3.5 mm, Ti. 5 Holes | No | √ | √
| LC- Dynamic Compression Plate 3.5 mm, Ti. 6 Holes | No | √ | √
| LC- Dynamic Compression Plate 3.5 mm, Ti. 7 Holes | No | √ | √
| DCP 3.5mm, st. steel, 5 holes | No | √ | √
| DCP 3.5mm, st. steel, 6 holes | No | √ | √
| DCP 3.5mm, st. steel, 7 holes | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 5 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 6 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 7 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 8 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 9 holes, Ti: | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 6 holes, st. steel | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 7 holes, st. steel | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 8 holes, st. steel | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 9 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 5 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 6 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 7 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 8 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 10 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 6 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 7 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 8 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 10 holes, st. steel | No | √ | √
Cancellous screws 4.0 mm, fully threaded, self taping:
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 14mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 16mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 18mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 20mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 22mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 24mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 26mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 28mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 30mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 35mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 45mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 50mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 55mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 60mm | No | √ | √
Cancellous bone screw 4.0mm, partially threaded:
| Cancellous bone screw 4.0mm, partially threaded: Screw length 14mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 16mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 18mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 20mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 22mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 24mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 26mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 28mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 30mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 35mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 40mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 45mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 50mm | No | √ | √
| Washer, 7.0mm for small screws | No | √ | √
Large Fragment Basic Instrument Set for LC-DCP and DCP
| Basic Instrument Set LC-DCP and DCP Graphic Case, perforated, autoclavable, complete with cover | No | √ | √
| Basic Instruments for LC-DCP and DCP, Synthes/Zimmer or Equivalent | Complete Set | √ | √
Large Fragment Instrument Implants
| Large Fragment Screws (4.5 mm): | No | √ | √
| 4.5mm cortex screws, self taping, length 30mm | No | √ | √
| 4.5mm cortex screws, self taping, length 32mm | No | √ | √
| 4.5mm cortex screws, self taping, length 34mm | No | √ | √
| 4.5mm cortex screws, self taping, length 36mm | No | √ | √
| 4.5mm cortex screws, self taping, length 38mm | No | √ | √
| 4.5mm cortex screws, self taping, length 40mm | No | √ | √
Cannulated Cancellous Screw 4.0 mm System
| Instrument Set for Cannulated Screw 4.0 mm, st. steel (Synthes/Zimmer or Equivalent) | Set | √ | √
| Contents consisting of: | Qty
| Guide wire,1.2mm dia,with threaded tip | 5 | √ | √
| Drill bit,dia.2.7mm,cannulated | 2 | √ | √
| Countersink,cannulated for 4.0mm cannulated screws | 1 | √ | √
| Tap,cannulated for 4.0mm cannulated screws | 2 | √ | √
| Screwdriver,hexagonal,cannulated sw 2.5 with holding sleeve | 1 | √ | √
| Screw holding forceps | 1 | √ | √
| Protection sleeve,8.0/3.0 | 1 | √ | √
| Guide sleeve,3.0/1.2 | 1 | √ | √
| Stylet,dia.1.2mm | 2 | √ | √
| Direct measuring device | 1 | √ | √
| Screw Extraction device | 1 | √ | √
| Screwdriver,sw2.5,solid | 1 | √ | √
| Basic tray for 4.0mm cannulated screw instrumentation | 2 | √ | √
| Top tray for 4.0mm cannulated screw instrumentation | 2 | √ | √
| Lid for 4.0mm cannulated screws instrumentation | 2 | √ | √
| 4.0mm Cannulated Screw Box | 2 | √ | √
Implants (4.0mm Cannulated Screws, Ti):
| 4.0mm cannulated screws, total length/thread length : Screw length 20/7mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw lengths 22/7mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 24/8mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 26/9mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 28/9mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 30/10mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw lengths 32/11mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 34/11mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 36/12mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 38/12mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 40/13mm | No | √ | √
| Washer 7/0/3.6 mm for screws 2.7 to 4.0 mm, st. steel | No | √ | √
| Instrument Set for Cannulated Screw 6.5 mm, st. steel (Synthes/Zimmer or Equivalent) | Set | √ | √
| Contents consisting of: | Qty
| Basic tray for 6.5mm cannulated screws instrumentation | 1 | √ | √
| Top tray for 6.5mm cannulated screws instrumentation | 1 | √ | √
| Lid for 6.5mm cannulated screw instrumentation tray | 1 | √ | √
| Guide wire,2.0mm dia with threaded tip | 5 | √ | √
| Drill bit,dia.4.5mm,cannulated | 2 | √ | √
| Countersink, cannulated for 6.5mm cannulated screws | 1 | √ | √
| Tap ,cannulated for 6.5mm cannulated screws | 1 | √ | √
| Parallel wire guide | 1 | √ | √
| Screwdriver, hexagonal, cannulated for 6.5mm cannulated screws | 1 | √ | √
| Depth gauze - 01no | 1 | √ | √
| Drill sleeve,8.0/5.0 | 1 | √ | √
| Protection sleeve, 11.0/8.0 | 1 | √ | √
| Stylet,dia.2.0mm | 2 | √ | √
| Drill sleeve,5.0/2.0 | 1 | √ | √
| Trocar | 1 | √ | √
| Screwdriver, SW 3.5,solid | 1 | √ | √
| Parallel wire guide, adjustable | 1 | √ | √
| Spanner | 1 | √ | √
Angled Blade Plate Ststem:
| Angled Blade Plate Instrument Set | Set | √ | √
| Contents consisting of: | Qty
| Grapic Case for Angled Blade plate instrument set without content | 1 | √ | √
| Kirschner wire with trocar tip, 2.0mm/150mm, SS | 5 | √ | √
| Drill Bit, 2.0mm dia, length 100/75 mm for quick coupling | 3 | √ | √
| Double Drill Sleeve 4.5/3.2 | 1 | √ | √
| Chisel Guide | 1 | √ | √
| Seating chisel, length 320mm | 1 | √ | √
| Inserter/Extractor with adjustable clamps length, 270mm | 1 | √ | √
| Slotted Hammer, length 270mm | 1 | √ | √
| Impactor | 1 | √ | √
| Postioning plate, triangular, 90°/50°/40° | 1 | √ | √
| Postioning plate, triangular, 80°/70°/30° | 1 | √ | √
| Postioning plate, triangular, 100°/60°/20° | 1 | √ | √
| Condylar plate guide | 1 | √ | √
Angled Blade Plate Implant:
| Angled Blade Plate 95 degree, 60 mm blade 5 holes | No | √ | √
| Angled Blade Plate 95 degree, 60 mm blade 6 holes | No | √ | √
| Angled Blade Plate 95 degree, 60 mm blade 7 holes | No | √ | √
| Angled Blade Plate 95 degree, 60 mm blade 8 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 5 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 6 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 7 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 8 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 5 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 6 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 7 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 8 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 5 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 6 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 7 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 8 holes | No | √ | √
Proximal Femoral Nail (PFN) System: PFN Instrument Set (Contents) consisting of:
| Proctecting sleeve | 1 | √
| Guide wire sleeve | 1 | √
| Proximal cannulated drill | 1 | √
| Quick coupling handle,T | 1 | √
| Screwdriver,T | 1 | √
| Tissue proctector | 1 | √
| Insertion handle | 1 | √
| Connecting screw | 2 | √
| Connector | 1 | √
| Adaptor | 1 | √
| Sliding hammer | 1 | √
| Combination wrench | 1 | √
| Proximal device bolt | 1 | √
| Guide sleeve,8.2/2.0 | 1 | √
| Trocar,4.0 | 1 | √
| Guide pin,threaded,3.2/400 | 3 | √
| Guide wire,threaded,2.0/340 | 2 | √
| Guide wire,2.0/340 | 2 | √
| Guide wire,2.5 / 340 | 2 | √
| Guide wire,threaded,2.5 / 320 | 2 | √
| Depth gauze for proximal screw | 1 | √
| Step cannulated reamer,4.5/6.3 | 1 | √
| Step cannulated reamer,6.3 / 9 | 1
| Stoper for step reamer | 2 | √
| Depth gauze | 1 | √
| Drill sleeve, 8.2 / 4.0 | 1 | √
| Drill sleeve, 8.2 / 11 | 2 | √
| Drill sleeve, 9.2 / 11 | 1 | √
| Guide sleeve, 9.2 / 2.5 | 1 | √
| Trocar, 2.5 | 1 | √
| Trocar, 2.0 | 1 | √
| Screwdriver , sw 3.5 | 1 | √
| Wrench for conical bolt | 1 | √
| Impactor-extractor rod | 1 | √
| Proximal locking device Radiolucent | 1 | √
| Cannulated Awl | 1 | √
| Wrench, sw 3 | 1 | √
| Drill bit, Stop, 4.0 | 1 | √
| Drill bit, dia.4.0 | 2 | √
| Screwdriver,cannulated,sw 5.0 | 1 | √
| Locking screw holder | 1 | √
| Screw inserter T type for 9.0mm screw | 1 | √
| Screwdriver sleeve | 1 | √
PFN Implants:
| Proximal femoral nail, 130˚, short, Dia. 9mm, Length 240mm, Titanium | No. | √
| Proximal femoral nail, 130˚, short, Dia. 10mm, Length 240mm, Titanium | No. | √
| Proximal femoral nail, 130˚, short, Dia. 11mm, Length 240mm, Titanium | No. | √
| Proximal femoral nail, 130˚, short, Dia. 12mm, Length 240mm, Titanium | No. | √
| End cap for Proximal femoral nail, without extension, Titanium | No. | √
| End cap for Proximal femoral nail, 5mm extension, Titanium | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 80mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 85mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 90mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 95mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 100mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 70mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 75mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 80mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 85mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 90mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 95mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 30mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 32mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 34mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 36mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 38mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 40mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 42mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 44mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 46mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 48mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 50mm | No. | √
Rush Rods
| Rush Rod 3mm diameter of various lengths | No. | √ | √
| Rush Rod 5mm diameter of various lengths | No. | √ | √
| Rush Rod 6mm diameter of various lengths | No. | √ | √
External Fixation System: Small External Fixator System: Small External Fixator Graphic Case with graphic outlines, complete with cover
Small External Fixator Instruments, consisting of:
| Combination Wrench, 7 mm width across flats | 1 | √ | √
| Socket Wrench, 7 mm width across flats | 1 | √ | √
| Small External Fixator Drill Sleeve System | 1 | √ | √
Small External Fixator Implants:
| 1.6 mm Kirscher Wire, 150 mm, trocar point | No. | √ | √
| 2.0 mm Kirscher Wire, 150 mm, trocar point | No. | √ | √
| 2.5 mm Kirscher Wire, 150 mm, trocar point | No. | √ | √
Fixation components
| Closed clamps, 4.0 mm/2.5 mm | No. | √ | √
| Closed clamps, 4.0 mm/4.0 mm | No. | √ | √
| Open clamps, 4.0 mm/2.5 mm | No. | √ | √
| Open clamps, 4.0 mm/2.5 mm | No. | √ | √
| Spring-Loaded Nut | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 60mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 80mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 100mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 120mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 140mm, | No. | √ | √
| Connecting Bars, 4.0mm dia. carbon fiber, length 160mm, | No. | √ | √
| Connecting Bars, 4.0mm dia. carbon fiber, length 180mm, | No. | √ | √
| Connecting Bars, 4.0mm dia. carbon fiber, length 200mm, | No. | √ | √
| Schanz Screws for small & large external fixator, thread length 20mm | No. | √ | √
| Schanz Screw 4.0/3.0 self drilling, total length 80mm | No. | √ | √
| Schanz Screw 4.0/3.0 self drilling, total length 100mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 60mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 80mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 100mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 120mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 140mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 160mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 175mm | No. | √ | √
| Schanz Screw, 4.5mm, self drilling, total length 150mm | No | √ | √
| Schanz Screw, 4.5mm, self drilling, total length 175mm | No | √ | √
| Schanz Screw, 5.0mm, self drilling, total length 150mm | No | √ | √
| Schanz Screw, 5.0mm, self drilling, total length 150mm | No | √ | √
Large External Fixator System: Large External Fixator Instruments, consisting of:
| Combination Wrench, 11 mm width across flats | 1 | √ | √
| Socket Wrench, 11 mm width across flats | 1 | √ | √
| Large External Fixator Drill Sleeve System | 1 | √ | √
Large External Fixator Components:
| Adjustable clamps | No | √ | √
| Open clamps | No | √ | √
| Universal Joint for two tubes | No | √ | √
| Tube-to-tube clamps | No | √ | √
| Transverse clamp | No | √ | √
| Combination clamp | No | √ | √
| Universal clamp | No | √ | √
Connecting Rods, 11.0mm dia. carbon fiber
| Connecting Rods, 11.0mm dia. carbon fiber, length 100mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 150mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 200mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 250mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 300mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 350mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 400mm, | No | √ | √
Reduction wires (olive wire) for Illizarov Fixator
| Reduction wire (olive), 1.8mm | No | √ | √
| Reduction wire (olive), 1.8mm | No | √ | √
Arthroplasty System
Prosthesis for Hemiarthroplasty:
| Austin Moore Endoprosthesis, Head 41.0 mm, Stem 140mm | No | √ | √
| Austin Moore Endoprosthesis, Head 43.0 mm, Stem 140mm | No | √ | √
| Austin Moore Endoprosthesis, Head 45.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 47.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 49.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 51.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 53.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 55.0 mm, Stem 160mm | No | √ | √
Bipolar Prosthesis:
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 37.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 39.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 41.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 43.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 45.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 47.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 49.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 51.0 mm, Stem length 155 mm | No | √ | √
Spine Implants
Pedicle screw ( Polyaxial): Synthes/Zimmer or Equivalent
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 30 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 32 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 35 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 38 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 40 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 30 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 35 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 40 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 45 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 35 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 40 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 45 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 50 mm | No. | √
Pedicle screw( Monoaxial): Synthes/Zimmer or Equivalent:
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 30 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 32 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 35 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 38 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 40 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 30 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 35 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 40 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 45 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 35 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 40 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 45 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 50 mm | No. | √
| Nut for Pedicle screw | No. | √
Hard Rods:
| 6.0mm hard rods: Length 250mm | No. | √
| 6.0mm hard rods: Length 350mm | No. | √
| 6.0mm hard rods: Length 450mm | No. | √
Spare Parts for Existing Instruments
Synthes DHS/DCS Instrument Spare Parts:
| 338.00, 2.5mm Threaded Guide Wires | No | 10 Nos
| 338.20, DHS/DCS Coupling Screw, Short | No | 3 Nos
| 338.21, DHS/DCS Guide Shaft | No | 3 Nos
| 338.010, Drill Bit | No | 1 No
| 338.47, DCS Reaming Head | No | 1 No
Karl Storz Arthroscopy Equipment Spare Parts:
| 495 ND Fiber Optic Cable, 3.5mm, length 400cm | No | 1 No
Power Drill Spare Parts:
| ConMed Power Pro, PRO200, Linvatec Cable Ref MC5057 | No | 1 No
| Rodar Power Control Unit, Cable Part number 15.01 | No | 1 No
Arthroscopic Surgery Accessory Supplies:
| HOPKINS II Telecope 30 degrees, 4mm, 18cm | Piece | 1
| Take-apart Scissors | Piece | 1
Pathology Department: Microbiology
Culture Media
| Agarose | 100 gm | √
| Alkaline peptone water | 100gm | √ | √ | √
| Bile esculin agar | 100gm | √ | √ | √
| Bismuth sulphate agar | 500gm | √ | √ | √
| Bolton broth | 500gm | √
| Brain heart infusion agar | 500gm | √ | √
| Brain heart infusion broth | 500gm | √ | √ | √
| Cary-blair medium | 500gm | √ | √ | √
| Columbia blood agar base | 500gm | √ | √ | √
| Cooked meat medium | 500gm | √
| Corn meal agar | 500gm | √
| Dnase test agar with toludine blue | 100gm | √ | √ | √
| Thioglycollate Broth | 100gm | √ | √
| Gc agar Base | 500gm | √ | √ | √
| Hippurate hydrolysis broth | 100gm | √ | √ | √
| Luria-Bertani broth | 100gm | √
| Nutrient broth | 500gm | √ | √ | √
| MacConkey agar with CV & Bile salt | 500gm | √ | √ | √
| MacConkey agar without salt | 500gm | √ | √
| Haemophilus test medium | 500gm | √ | √ | √
| Malonate broth | 100gm | √ | √ | √
| Modified CCDA agar | 500gm | √
| MR-VP medium | gm | √ | √ | √
| Mueller hinton agar | 500gm | √ | √ | √
| Nitrate broth | 500gm | √ | √ | √
| Nutrient agar | 500gm | √ | √
| Nutrient gelatin | 100gm | √ | √ | √
| Oxidative Fermentative Basal medium | 100gm | √ | √ | √
| Ornithine decarboxylase broth | 100gm | √ | √ | √
| Lysin decarboxylase broth | 100gm | √ | √ | √
| Arginine decarboxylase broth | 100gm | √ | √ | √
| Rice extract agar | 500gm | √
| Sabrouad dextrose agar | 500gm | √ | √
| Sabrouad dextrose agar with antibiotic supplements | 500gm | √ | √
| Selenite broth | 500gm | √ | √ | √
| SIM medium | 500gm | √ | √ | √
| Simmons citrate agar | 100gm | √ | √ | √
| Soyabean casein digest medium with sps | 500gm | √ | √
| TCBS cholera medium | 100gm | √ | √
| Triple sugar iron agar | 100gm | √ | √ | √
| Urea broth base | 100gm | √ | √ | √
| XLD agar | 500gm | √ | √ | √
| Cystein lactose electrolite deficient agar (CLED) | 500gm | √ | √ | √
| MR ( Methyl Red) broth - peptone - dextrose - potassium- l napthol | 500 gm pack each | √
Growth Supplements
| Haemoglobin powder (to be used with GC Agar) | 100gm | √ | √ | √
| Urea 40% solution | 5ml | √ | √ | √
| VCNT supplement | gm | √ | √ | √
| Vitamin growth supplement | gm | √ | √ | √
| Antibiotics Disc |  | √ | √ | √
| Amikacin, 30mcg | 50 Disc/ vial | √ | √ | √
| Amphicilin, 10mcg | 50 Disc/ vial | √ | √ | √
| Aztreonam (30 mcg) | Disc | √ | √ | √
| Carbenicillin 100mcg | Disc | √ | √
| Cefoparazone (30 mcg) | Disc | √ | √
| Cefazolin 30mcg | 50 Disc/ vial | √ | √ | √
| Cefepime (30mcg) | Disc | √ | √ | √
| Cefixime (5mcg) | Disc | √ | √ | √
| Cefotoxime/Calvunalic acid (30/10 mcg) | 50 Disc/ vial | √ | √ | √
| Cefoxitin disc (10mcg) | 50 Disc/ vial | √ | √ | √
| Cefpirome (30 mcg) | Disc | √ | √
| Cefpodoxime - (30mcg) | Disc | √ | √
| Ceftazidime 30mcg | 50 Disc/ vial | √ | √ | √
| Ceftazidime/Calvunalic acid (30/10 mcg) | 50 Disc/ vial | √ | √ | √
| Ceftriaxone, 30mcg | 50 Disc/ vial | √ | √ | √
| Cefuroxime (30mcg) | Disc | √ | √ | √
| Cephalexin, 30mcg | 50 Disc/ vial | √ | √ | √
| Cephotaxime, 30mcg | 50 Disc/ vial | √ | √ | √
| Chloramphenicol, 30mcg | 50 Disc/ vial | √ | √ | √
| Ciprofloxacin, 5mcg | 50 Disc/ vial | √ | √ | √
| Clarithromycin 15mcg | 50 Disc/ vial | √ | √ | √
| Clindamycin 2mcg | 51 Disc/ vial | √ | √ | √
| Cotri-moxazole, 25mcg | 50 Disc/ vial | √ | √ | √
| Doxycycline Hcl, 30mcg | 50 Disc/ vial | √ | √ | √
| Erythromycin, 15mcg | 50 Disc/ vial | √ | √ | √
| Gentamicin, 10mcg | 50 Disc/ vial | √ | √ | √
| Gentamicin, 120mcg | 51 Disc/ vial | √ | √ | √
| Imipenam (10mcg) | 50 Disc/ vial | √ | √ | √
| Meropenem (10mcg) | 51 Disc/ vial | √ | √ | √
| Nalidixic acid, 10mcg | 50 Disc/ vial | √ | √ | √
| Nitrofurantoin, 300mcg | 50 Disc/ vial | √ | √ | √
| Norfloxacin, 10mcg | 50 Disc/ vial | √ | √ | √
| Novobiocin Disc (5mcg ) | 50 Disc/ vial | √ | √ | √
| Oxacillin 1mcg | 50 Disc/ vial | √ | √ | √
| Penicillin G, 10units | 50 Disc/ vial | √ | √ | √
| Pipracillin, 100mcg | 50 Disc/ vial | √ | √ | √
| Polymixin B 300mcg | 50 Disc/ vial | √ | √ | √
| Tobramycin, 10mcg | 50 Disc/ vial | √ | √ | √
| Vancomycin, 30mcg | 50 Disc/ vial | √ | √ | √
| Agumentin | 50 disc/vial | √
| Amphicillin | 50 disc/vial | √
| E-test Cefotoxime (NCCLS) | 100 Strips | √ | √ | √
| E-test Ceftriaxone (NCCLS) | 100 Strips | √ | √ | √
| E-test Penicillin (NCCLS) | 100 Strips | √ | √ | √
| E-test Meropenem (NCCLS) | 100 Strips | √ | √ | √
| E-test Vancomycin (NCCLS) | 100 Strips | √ | √ | √
| E test Oxacillin | 100 strips | √
| Idenfication disc |  | √ | √ | √
| Bacitracin disc | 50 Disc/ vial | √ | √ | √
| Optochin disc | 50 Disc/ vial | √ | √ | √
| Oxidase disc | 50 Disc/ vial | √ | √ | √
| Nitrocefin beta - lactamase | 50 Disc/ vial | √ | √ | √
| V Factor disc | 50 Disc/ vial | √ | √ | √
| X factor disc | 50 Disc/ vial | √ | √ | √
| X+V Factor disc | 50 Disc/ vial | √ | √ | √
| Stains & Chemicals |  | √ | √ | √
| Acetone, ExcelaR | 500 ml | √ | √ | √
| Laboratory disinfectent solution | 5Ltrs | √ | √ | √
| Alberts stain A | 125 ml | √ | √
| Alberts stain B | 125 ml | √ | √
| Alcohol based antiseptic handrub | 500 ml | √ | √ | √
| Amyl alcohol | 500 ml | √ | √
| Anaerobic gas generating kit | pkt | √
| Auramine O | 25 gm | √
| Basic fuchsin | 25 gm | √ | √ | √
| Grams crystal violet | 125 ml | √ | √ | √
| Ethidium Bromide | gm | √
| Glycerol (reagent grade) | 500 ml | √ | √
| Grams iodine | 125ml | √ | √ | √
| Lactophenol cotton blue | 125ml | √ | √
| Methylene blue powder | 25gm | √ | √ | √
| N-Acethyl-l-cysteine | 100gm | √ | √
| N-naphthylethylenediamine dihydrochloride | 25 gm | √ | √
| Phenol crystals | 500 gm | √ | √ | √
| Potassium Iodate | 100 gm | √ | √ | √
| Potassium permanganate | 500 gm | √
| Safranin, 0.5% | 125 ml | √ | √ | √
| Sodium Nitrate | 100gm | √ | √
| Spore strips | Strip | √ | √ | √
| Sulfanilamide | 100gm | √ | √
| Plastic wares/ Other |  | √ | √ | √
| Gas Regulator for Co2 | piece | √
| Gas Regulator for mix Gas cyclinder | piece | √
| Antibiotic Zone Scale TM-C | piece | √ | √ | √
| Test tube caps 12/13mm - blue | piece | √ | √ | √
| Test tube caps 12/13mm - silver | piece | √ | √ | √
| Test tube caps 12/13mm - red | piece | √ | √ | √
| Test tube caps 15/16mm - silver | piece | √ | √ | √
| Bottle w/pouring ring & screw cap (Max. temp. 140˚C) 1000ml | piece | √ | √
| Bottle w/pouring ring & screw cap (Max. temp. 140˚C) 500ml | piece | √ | √
| Conical Tubes (Falcon,polypropyline), 15ml | piece | √ | √
| Conical Tubes (Falcon, polypropyline), 50ml | piece | √ | √
| Cryovials (2ml) | 1000/ pkt | √ | √ | √
| McCartney vial w/aluminium caps w/black rubber seals 14 ml | piece | √ | √
| Micropipette tips: 0.5 - 20 µl | piece | √ | √
| Metal loop TM. CH, changeable | piece | √ | √ | √
| Nichrome loop-D-2 | piece | √ | √ | √
| Nichrome loop-D-1 | piece | √ | √ | √
| Petri plates sterile 90mm diameter. | piece | √ | √ | √
| Sterile disposable pipette, 5ml (Individually wrapped) | 100/pkt | √ | √ | √
| Sterile swab, general purpose with screw cap | 100/ pkt | √ | √ | √
| Sterile swab with wooden stick | 100/ pkt | √ | √ | √
| Sterile screw capped stool container with attached scoop (20ml) | 100/pkt | √ | √ | √
| Sterile, single use filter unit 0.20 µm | piece | √ | √
| Sterile screw caped urine container (50ml) | 500/ pkt | √ | √ | √
| Tips rack, resuable & autoclavable for tips 0.1 - 10 µl, 34 mm | piece | √
| Tips rack, resuable & autoclavable for tips: 0.5 - 20 µl, 46 mm | piece | √
| Tube racks, 24 place capacity for 1.5/2.0 ml | piece | √ | √ | √
| Tube, safe lock for PCR, 1.5 ml | piece | √
| Tube, safe lock for PCR, 2.0 ml | piece | √
| Transparent autoclavable plastic bag (1 ltr) | piece | √ | √
| Glasswares |  | √ | √ | √
| Aspirator bottle with tap, Polythene bottle with tubulation, Capacity: 20litres | piece | √ | √ | √
| Burner, Bunsen (ordinary) for LPG | piece | √ | √ | √
| Gel Electrophoresis Chamber Set with power supply | piece | √
| Glass slide with single cavity | piece | √ | √ | √
| Mix gas (5 % oxygen 15% Co2 85% Nitrogen) with regulator | piece | √
| PCR Gel casting & running tray, with combs | piece | √
| Petri plate Rack | piece | √ | √
| Slide, slope-rack | piece | √ | √
| UV face shield | piece | √
| UV transparent Gel scoops for 10 x 7cm | piece | √
| Test kit |  | √ | √ | √
| API 20 C | 25 test | √
| API 20 strept | 25 test | √
| API 20A | 25 test | √
| API 20E with complete reagent kit | 25 test | √ | √
| API 20 GN with complete reagent kit | 25 test | √ | √
| API Campy with complete reagent kit | 25 test | √
| API NH with complete reagent kit | 12 test | √ | √
| Rapid 1D 32 E (API) | 25 tests/pkt | √
| Bacterial antigen test kit | 50 test | √ | √
| Helicobacter pylori Elisa test kit(IgM) | 96 test/Kit | √
| Helicobacter pylori Elisa test kit(IgG) | 96 test/Kit | √
| Mineral oil | 100 ml | √ | √
| Rota-Virus, latex test kit | 25 test | √ | √
| Streptococcal grouping test kit | 50 test | √ | √ | √
| DNA extraction kit (bacterial) | test | √
| dNTP | ml | √
| 50 milli mole Magnesium Chloride | gm | √
| TAE buffer (10x) | ml | √
| Taq polymerase with buffer and MgCl2 500unit | piece | √
| TBE buffer (10x) | ml | √
| Kala Azar RK39 Rapid test kit | 25 test/ kit | √ | √
| Dengue Rapid test kit(IgM,IgG) | 25 test/ kit | √ | √ | √
| Dengue Elisa test Kit(IgG) | 96 wells | √ | √
| Scrub typhus Rapid test kit(IgM,IgG) | 30 test/ kit | √ | √ | √
| Anti Sera |  | √ | √ | √
| Escherichia coli Agglutinating Sera poly 2 | 2ml | √
| Escherichia coli Agglutinating Sera poly 2 - 3 | 2ml | √
| Escherichia coli Agglutinating Sera poly 2 - 4 | 2ml | √
| Escherichia coli type 0111:K58(B4) | 2ml | √
| Escherichia coli type 0112:K66(B11) | 2ml | √
| Escherichia coli type 0114:K90(B) | 2ml | √
| Escherichia coli type 0119:K69(B14) | 2ml | √
| Escherichia coli type 0124:K72(B17) | 2ml | √
| Escherichia coli type 0125:K70(B15) | 2ml | √
| Escherichia coli type 0126:K71(B16) | 2ml | √
| Escherichia coli type 0127:K63(B8) | 2ml | √
| Escherichia coli type 0128:K67(B12) | 2ml | √
| Escherichia coli type 0142:K86(B) | 2ml | √
| Escherichia coli type 018C:K77(B21) | 2ml | √
| Escherichia coli type 026:K60(B6) | 2ml | √
| Escherichia coli type 044:K74(L) | 2ml | √
| Escherichia coli type 055:K59(B5) | 2ml | √
| Escherichia coli type 086:K61(B7) | 2ml | √
| Salmonella Agglutinating Sera phase 2 | 2ml | √ | √
| Salmonella Agglutinating Sera polyvalent H, phase 1 & 2 | 2ml | √ | √
| Salmonella Agglutinating Sera (1,2 - H) | 2ml | √
| Salmonella Agglutinating Sera (1,5 - H) | 2ml | √
| Salmonella Agglutinating Sera (1,7 - H) | 2ml | √
| Salmonella Agglutinating Sera (a - H) | 2ml | √ | √
| Salmonella Agglutinating Sera ( b - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (c - H ) | 2ml | √ | √
| Salmonella Agglutinating Sera (d - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (e, n, z 15 - H) | 2ml | √
| Salmonella Agglutinating Sera (eh - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (fg - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (gm - H) | 2ml | √ | √
| Salmonella Agglutinating Sera ( i - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (iv - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (k - H) | 2ml | √ | √
| Salmonella Agglutinating Sera (mt - H | 2ml | √ | √
| Salmonella Agglutinating Sera (r - H) | 2ml | √ | √
| Salmonella Agglutinating Sera polyvalent O, Groups A - G | 2ml | √ | √
| Salmonella Agglutinating Sera (2-0) | 2ml | √ | √
| Salmonella Agglutinating Sera (4-0) | 2ml | √ | √
| Salmonella Agglutinating Sera ( 6, 7-0) | 2ml | √ | √
| Salmonella Agglutinating Sera ( 8-0) | 2ml | √ | √
| Salmonella Agglutinating Sera ( 9-0) | 2ml | √ | √
| Salmonella Agglutinating Sera polyvalent O (Vi) | 2ml | √ | √
| Shigella Flexneri polyvalent 1 - 6, X & Y | 2ml | √ | √
| Shigella Sonnei phases 1 & 2 | 2ml | √ | √
| Shigella Agglutinating Sera dysentry Polyvalent 1-10 | 2ml | √ | √
| Shigella Boydii polyvalent 1 (1 - 6) | 2ml | √ | √
| Shigella Boydii polyvalent 2 (7-11) | 2ml | √ | √
| Shigella Boydii polyvalent 3 (12-15) | 2ml | √ | √
| Vibrio cholera polyvalent O-1 | 2ml | √ | √
| Vibrio cholera polyvalent Inaba | 2ml | √ | √
| Vibrio cholera polyvalent Ogawa | 2ml | √ | √
Equipment specific Reagents
| FA (Fan Aerobic) for Bact/Alert 3D | 100 bottles/ box | √
| PF (Pediatric Fan) for Bact/Alert 3D | 100 bottles/ box | √
| FA (Fan Aerobic) for BD 9050 | 50 bottles/ box | √ | √ | √
| PF (Pediatric Fan) for BD 9050 | 50 bottles/ box | √ | √ | √
| Flash Disc Bact Alert Blood culture system | piece | √
| Reflectance Standards for Bact/Alert system (A & B) | piece | √
Control strains
| Escherichia coli ATCC 25922 | loop | √ | √ | √
| Pseudomonas aeruginosa ATCC 27853 | loop | √ | √ | √
| Staphylococcus aureus ATCC 25923 | loop | √ | √ | √
| Klebsiella pneumoniae ATCC BAA 1705 | loop | √
| Klebsiella pneumoniae ATCC BAA1706 | loop | √
Infectious Disease Serology Reagent
| Anti-HBc for Vitrous ECIQ System | ml | √ | √
| Anti-HBs for Vitrous ECIQ System | ml | √ | √
| Anti-HCV for Vitrous ECIQ System | ml | √ | √
| Anti-HIV for Vitrous ECIQ System | ml | √ | √
| CMV IgG for Vitrous ECIQ System | ml | √ | √
| CMV IgM for Vitrous ECIQ System | ml | √ | √
| HBeAg for Vitrous ECIQ System | ml | √ | √
| HBsAg for Vitrous ECIQ System | ml | √ | √
| Rubella IgG for Vitrous ECIQ System | ml | √ | √
| Rubella IgM for Vitrous ECIQ System | ml | √ | √
| Toxo IgM for Vitrous ECIQ System | ml | √ | √
| Toxo Plasma IgG for Vitrous ECIQ System | ml | √ | √
| Anti HBC | ml | √ | √
| Anti HBs calibrator | ml | √ | √
| Anti HIV calibrator | ml | √ | √
| CMV Calibrator | ml | √ | √
| CMV Control | ml | √ | √
| HBc IgM calibrator | ml | √ | √
| HBe calibrator | ml | √ | √
| HBeAg calibrator | ml | √ | √
| HBsAg calibrator | ml | √ | √
| HCV calibrator | ml | √ | √
| Rubella IgG Calibrator | ml | √ | √
| Rubella IgG control | ml | √ | √
| Toxo IgM calibrator | ml | √ | √
| Toxo Plasma IgM controls | ml | √ | √
| Anti - HBC IgM Controls | ml | √ | √
| Anti HBe Controls | ml | √ | √
| Anti HBs Controls | ml | √ | √
| Anti HIV Controls | ml | √ | √
| HBsAg Controls | ml | √ | √
| HCV Controls | ml | √ | √
| HBc lgM ( 52 well ) for Vitrous ECIQ System | ml | √ | √
| HBe for Vitrous ECIQ System | ml | √ | √
| Diluent A | ml | √ | √
| Diluent B | ml | √ | √
| Signal Reagent | ml | √ | √
Rapid/Elisa test Kit
| HBsAg Rapid Test Kit, sensitivity 100%, Specificity >98%, Reference Lab evaluated. Assay type: Rapid Diagnostic Test, Immunochromatography (lateral-flow tests or immunofiltration) Antibody type: monoclonal Solid phase: Nitrocellulose membrane strip Specimen type: Serum/plasma from venous whole blood, Detection type: HBsAg, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
| HCV Rapid Test Kit, sensitivity 100%, Specificity >98%, Reference Lab evaluated. Assay type: Rapid Diagnostic Test, Immunochromatography (lateral-flow testsorimmunofiltration) Antigen type: Recombinant proteins synthetic peptides core and NS3 and NS 5 antigens Solid phase: membrane strip Specimen type: Serum/plasma /blood Detection type: HCV antibodies, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
| HIV Rapid Test Kit: WHO prequalified,Reference Lab evaluated. sensitivity 100%, specificity >98%. Assay type: Rapid Diagnostic Test, Immunochromatography (lateral-flow tests or immunofiltration), Antigen type: Recombinant and synthetic peptides for HIV-1 gp41, HIV-1 gp120, and HIV-2 gp36, Solid phase: Nitrocellulose membrane strip, Specimen type: Serum/plasma / whole blood, Detection type: Combined detection of HIV-1/2 antibodies Subtype detection: M N O sub types, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
| RPR floculation test kits. Assay type: Macroscopic flocculation test Antibody type: Monoclonal antibodies Specimen type: Serum/plasma Detection type: Anti- lipoidal antibodies, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
| TPHA Rapid Test Kit, Assay type: Hemagglutination assay/card test/rapid test Antibody type: Monoclonal antibodies Specimen type: Serum/plasma Detection type: Antibodies against the causative agent of syphilis, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
| Anti-HIV 1 &2, ELISA, WHO prequalified, Reference Lab evaluated, sensitivity 100%, Specificity >99.4%, Analytical sensitivity: must be less than 60 picograms/ml and /or less than 0.050 IU/ml as per different standardsAssay type: 3rd generation 'Sandwich' ELISA, Antigen type:Recombinant and synthetic peptides for HIV-1 gp41, HIV-1 0 gp41, and HIV-2 gp36, Solid phase:12 x 8 microwell strips, Specimen type: Serum/plasma, Detection type:Combined detection of HIV-1/2 antibodies, Subtype detection: M N O sub types, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life : After receiving at central depot should be 12 months at 2-8 C. | 96 test/192 tests/ Kit | √ | √
| HBsAg ELISA, Reference Lab evaluated, 100%, Specificity >98%, Analytical sensitivity: Must detect less than 60 picograms/mland /or less than 0.050 IU/ml as per different standards, Assay type: 3rd generation 'Sandwich' ELISA, Antibody type: Monoclonal antibodies, Solid phase: 12 x 8 microwell strips, Specimen type: Serum/plasma Detection type: detection of HBsAg, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life: After receiving at central depot should be 12 months at 2-8 C. | 96 test/ Kit | √ | √
| Anti-HCV ELISA,sensitivity 100%, Specificity >98%, Reference Lab evaluated, Assay type: 3rd generation 'Sandwich' ELISA, Analytical sensitivity: must be less than 60 picograms/ml and /or less than 0.050 IU/ml as per different standards, Antigen type: Encoded with recombinant antigens to Core, NS3, NS4and NS 5 antigens, Solid phase: 12 x 8 microwell strips, Specimen type: Serum/plasma Detection type: Detection of HCV antibodies. Must detect genotypes 1,3 and 6 distributed in Asian population and subtypes of HCV, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life: After receiving at central depot should be 12 months at 2-8 C. | 96 test/ Kit | √ | √
| HIV-1/2 is an in vitro diagnostic test kit for the detection of antibodies to Human Immunodeficiency Viruses type1 and/or type2 (HIV-1/HIV-2) in human serum/plasma. sensitivity 100%, Specificity >98%, Reference Lab evaluated, WHO prequalified, Shelf life : After receiving at central depot should be 12 months at 2-8 C. | 96/ test/Kit
| Infectious mononucleosis test | 20 test/ Kit | √ | √
| Typhidot test for Typhoid | 25 test/ kit | √ | √ | √
| Tuberculin PPD (dil) 5TU/0.1ml | 5ml | √ | √ | √
TORCH PANEL
| Toxoplasma IgG ELISA | 96 wells | √
| Toxoplasma IgM ELISA | 96 Wells | √
| Rubella IgG ELISA | 96 wells | √
| Rubella IgM ELISA | 96 wells | √
| CMV IgG ELISA | 96 Wells | √
| CMV IgM ELISA | 96 Wells | √
| Herpes IgG ELISA | 96 Wells | √
| Herpes IgM ELISA | 96 Wells | √
| HBeAg ELISA | 96 wells | √
| Ciprofloxacin(1ug) | 50 Disc/ vial | √ | √ | √
| Ceftriaxone(0.5ug) | 50 Disc/ vial | √ | √ | √
| Penicillin G(0.5ug) | 50 Disc/ vial | √ | √ | √
| Tetracycline(10ug) | 50 Disc/ vial | √ | √ | √
| Cefpodoxime(10ug) | 50 Disc/ vial | √ | √ | √
| Azithromycin | 50 Disc/ vial | √ | √ | √
Reagents for CD4 counter
| BD FACS Count Control | 25 test/ pkt | √ | √ | √
| BD FACS Count Reagent | 50 test/ pkt | √ | √ | √
| BD FACS Flow Solution | 20Ltrs | √ | √ | √
| Thermal paper Printer roll for BD FACS count | rolls | √ | √ | √
| CD4 count, Pima Bead Std | Pkt | √ | √ | √
| CD4 count, Pima Cartridge | 100test/ kit | √ | √ | √
| Echinococcus ELISA | 96test/ Pkt | √
| Kala azar rapid test kits | 96test/ Pkt | √ | √ | √
| Anaerobic indicator | 50 pcs/ pkt | √ | √
| Catalyst | 10 pcs/ pkt | √
| Anaerobic Mixed gas (5 % H2 5% Co2 90% Nitrogen) with regulator | Pcs | √
Clinical Parasitology
| IUI- sperm washing media/ sperm rinse ( Bicarbonate and HEPES buffered medium containing human serum albumin) 21 mM HEPES buffer 4 mM Sodium Bicarbonate 5 mg/ml (0.5%) Human Serum Albumin Storage 2-8°C Shelf Life-2 years from date of manufacture | 12 x12 mL, 100 mL. | √
Any equipments selected, the supplier should be able to install the equipment , give training to the staff and the Biomedical unit in the hospital, only then the equipment will be selected.
Equipments
| Balance: 0.0001g Readability. Maximum capacity: 210g, Readability: 0.0001g. Settling time (typical): 10…16s, Adjustment with internal weights: proFACT, fully automatic time- and temperature-controlled adjustment. Adjustment with external weights: custom weight. Linearization: automatic linearization with each internal adjustment. Sensitivity (temperature Sampling pump drift): ± 1.0 ppm/°C. Sensitivity (long-term stability): ± 0.0001%. Free height above weighing pan: 55mm. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
| Media Dispenser: Range: 0.5ml - >5 liters. Accuracy,%: ± at 10 ml. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
| Cellulose acetate sterile syringe filter: Filter pore size 0.45 µm, 25 mm diameter membrane, surfactant free-cellulose acetate (SFCA). sample volume: 0.5 ml to 100 ml. | Set | √ | √
| Micropipette: adjustable, volume range: 0.1 - 2.5 µl | Piece | √
| Micropipette: adjustable, volume range: 2 - 25 µl | Piece | √
| 8 Channel Micropipette: adjustable, volume range: 50 - 200 µl | piece | √ | √
| Micro-centrifuge: Microprocessor controlled. Max. RPM: 15,000 RPM. Max. RCF: 16,853 x g. Max. Capacity: 1.5ml x 24 tubes. 200 µl PCR tube adaptor. Display: back light LCD. Over speed protector, door locking, unbalance detector. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
| Anaerobic Work Station: Pressure controll: Automatic. Temp. Range: Ambient + 5°C to 70°C. Temp. Accuracy: +/-0.2°C at 37°C. Temp. Controller: Digital P.I.D. controller, 4Zone PID. Display: Dual LED display, Circulation fan: 5" convection fan. Standard device: Vacuum pump, Glove, Catalyst Pack. Circulator, Drilite Pack, Vacuum gauge, Pressure gauge, Glass shelf. Power supply: 220-240V, 50/60Hz. It should be complimented with: Filled different gas cylinder 6 Nos each. Gas regulator 6 Nos. Tubings & Voltage stabilizer. | Set | P
| Anaerobic jar: Capacity: 2.5 liters Stainless steel/ polypropylene | Piece | √ | √
| Biological safety cabinet-Class II type A1. Main control: Full digital microprocessor control with LCD digital display. Airflow: should be 70% recirculation 30% exhausted. Filters: Should have 0.3 micron 99.9% efficiency and HEPA should be protected by a perforated metal diffuser covering the top of the work zone. Work access opening should be 8 inches (203mm) high. Cabinet should be easily fumigated employing an established procedure. Sound level shall be no more than 67 dbA measured 15 inches (381mm) above the work tray and 12 inches (305mm) in front of viewing window. Illumination: Mounted Fluorescent lighting with UVG lamp. Utility: Two gas valve/service couplings, three or four (230VAC) internal electrical circuits. Cabinet shall be easily converted to a free-standing console model with the addition of the optional base support stand. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √ | √
| Biological safety cabinet-Class II type II B1: Main control: Full digital microprocessor control with LCD digital display. Airflow: should be 30% recirculation 70% exhausted. Filters: Should have 0.3 micron 99.9% efficiency and HEPA should be protected by a perforated metal diffuser covering the top of the work zone. Work access opening should be 8 inches (203mm) high. Cabinet should be easily fumigated employing an established procedure. Sound level shall be no more than 67 dbA measured 15 inches (381mm) above the work tray and 12 inches (305mm) in front of viewing window. Utility: Two gas valve/service couplings, three or four (230VAC) internal electrical circuits. Cabinet exhaust duct connection shall be 12-inch (305mm) diameter. Cabinet shall be easily converted to a free-standing console model with the addition of the optional base support stand. Power: 220V, 50/60Hz. Accessory: Air vilocity measuring device, Storage Pull-Out Trays. Stainless Steel Arm Rest and elbow rests. | Set | P | P
| Blood culture system (It should have minimum 120 cells): An automated modular system with fluorocesence/coloremetric. Technology for culturing, of bacteria and yeast fungi. The system should be able to manage the bottle handling functions such as loading and unloading bottles and regulates the temperature of the system. It should have minimum 120 cells in separate or one drawer module. The system should be configured independently to test blood culture and body fluid cultures. The system should provide basic data management function. It should have an interface to the Laboratory Information System (LIS) to eliminate manual data entry. Reporting can be customized and scheduled to print at a pre-determined time. Should be supplied with latest computer with OS/ (Windows XP), analysis soft ware standard keyboard, mouse,17" LCD monitor and Laser jet printer Should have latest external computer and a Laser jet printer and bar code reader. It should be supplied with the reflectance standard & UPS. supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √
| Blood culture system (It should have minimum 60 cells): An automated modular system with fluorocesence/coloremetric. Technology for culturing, of bacteria and yeast fungi. The system should be able to manage the bottle handling functions such as loading and unloading bottles and regulates the temperature of the system. It should have minimum 60 cells in separate or one drawer module. The system should be configured independently to test blood culture and body fluid cultures. The system should provide basic data management function. It should have an interface to the Laboratory Information System (LIS) to eliminate manual data entry. Reporting can be customized and scheduled to print at a pre-determined time. Should be supplied with latest computer with OS/ (Windows XP), analysis soft ware standard keyboard, mouse,17" LCD monitor and Laser jet printer Should have latest external computer and a Laser jet printer and bar code reader. It should be supplied with the reflectance standard & UPS. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √
| Bunsen burner: 3 mode system (continuous/sensor/footswitch) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √
| Electrophoresis power supply: For gel electrophoresis 100 hour up/down timer. Voltage adjustable from 40 to 120V. Precise digital control. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √
| Microscope, binocular: phase contrast, darkfield & bright field Illumination: with objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top . | Set | √
| Gel photo documentation: Gel imaging system, PC, includes darkroom, camera, cables, image lab software; samples trays with complete accessories, printers and thermal papers. | set | √
| Auto-steam sterilizer (Vertical). With microprocessor digital control and temperature display. Capacity: 80 - 100 liters. Operating temperature range: 105 - 142°C. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √ | √
| Hot plate & magnetic stirrer: Speed Range: 30 - 1500 RPM. Temperature Range: Ambient to 350°C. Temp. Controller: Hydraulic Thermostat. Heater: Interchangeable plate heater. Plate Area: 180 x 180 mm. Top Plate: Ceramic coated top/plate. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √ | √ | √
| Ice maker: Production: 250 liters/day. Storage Capacity: 125 liters. Material (IN): Stainless steel. (OUT): Steel plate with powder coating. Door: Front sliding door. Safety Device: Power breaker, Water level monitor, Inventory level monitor. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
| Incubator CO2: Main control: VFD Displayed Microprocessor Controller. Temperature range: Ambient + 5ºC to 60ºC. Co2 range: 0 to 20%, +/-0.1% at 5%. Co2 detection: Infra-red sensor. Humidity: Up to 96% RH. Programmability: 10 steps / 10 Memory. Display: VFD Digital, 80 character display. Capacity: 165 liters. Safety device: Over/Under Temp. Protector, Gas high/low monitor, Water level monitor, Power breaker. Power: 220V, 50/60Hz. It should be supplied with the following: Filled Co2 Cylinder 6 Nos. Co2 gas regulator 2 Nos. Tubings, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | set | √
| Thermal cycler: Sample capacity: 96 x 0.2 ml PCR tubes or one 96 PCR plate. Block temperature range: 4 - 99°C. Degree range of gradient, maximum: 1-24°C. Temperature control range of gradient: 30 - 99°C. Temperature of lid: 105°C. Block homogenicity: 20-72 ± 0.3°C, 90°C ± 0.4°C. Control accuracy: ± 0.2°C. Heating speed: approx. 6°C/s. Cooling speed: approx. 4.5°C/s. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. Should be supplied with compatible computer and operating systems. | set | √
| Turbidometer: For measuring 0.5 Macfarland turbidity standard.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
| Microplate strip washer: Microplate: 96 wells plate washer. Wash programs: Up to 40. Wash cycles Up to 9. Fluid flow rates; variable, 150 - 1000uL/well/s. Dispense volumes: 50 to 450µl in 1 µl increments. Dispensing accuracy: <5% at 300µl. Display: LCD display with touch screen. Automatic internal four buffer switch. Washing-manifold type: 1x8 manifold. Soaking time programmable in minutes and up to 60 seconds. Washing speed: 3 aspirate and dispense cycles >300 µl/well of 96 wells 8 tube manifold till final aspirate time: ≤130 seconds. Operating temperature: -10°C to +40°C. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √ | √*
| Microplate reader: Detection method: Absorbance.Read method: End point, kinetic, linear scanning. Microplate type: 96 well plate. Reading speed: 8 single wavelengths, 18 double wavelength. Measurement system: 8-channel optical system. Wavelength range: 300 to 700nm. Filters nm: 340, 405, 450, 492, 630. Filter capacity: 6-filter capacity wheel. Light source: Tungsten halogen. Shaking: Low, Medium, High and variable speeds. Reading range: Absorbance range. Abs; 0.000 - 4.000. Dynamic range: 0 to 4.0 OD. Optical Density (OD) accuracy: <1% at 2.5 OD typical. Reading speed: 8 seconds interval. Display: large LCD display with touch screen. External printer compatible to reader. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √ | √ | √*
| Serological analyzer (Infectious diseases): An automated modular system based on the principle of Chemiluminescence. Reaction technology OR enhanced Electrochemiluminescence. Along with hormone and other parameters the analyzer should be able to do. Hepatitis including HBsAg, HCV serology, HIV serology and TORCH panels. Throughput: 60 - 80 test or more/Hour. Must have continuous sample loading. There should be bar codes reader for reagents, controls & calibrators. Analyzer should be capable of using primary sample tubes itself & be able to detect clots in sample. STAT sample positioning should be available. Facility for on board cooling should be available. It should have external printer, UPS. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
| Shaker: (Rotary mixer), Programmable Speed Range: 0 - 300 RPM. Speed controller: Electronic controller. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √ | √
| Incubator (Shaking): Microprocessor controlled PID type. Temperature range: 35˚C ±2˚C. Capacity: 45-55 litres. Over temperature cut-out, programmable shaking frequency up to 250 rotations per hour. Power supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top | Set | √
| Microscope, binocular, fluorescent: Illumination: with objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X Complete with power cord. Power supply: 220-240V, 50/60HZ., | Set | √
| FULLY AUTOMATED URINE SEDIMENTANALYZER: Should be able to identify cells, casts and crystals. Other Test Items should include Urobilinogen, Bilirubin, Ketone, Blood, Protein, Nitrate, glucose, Specific Gravity, Ascorbic acid, pH Wavelength: 557, 620, 656nm Principle: Reflectance photometer Strips: 11 items Throughput: 60 to120 tests per hour, Display: Large LCD display Memory: about 1000 results Printer: Built in thermal printer Interface: Standard RS 232 Working environment: Temperature: -10 to +40°C, Humidity: 0 to 95%; with barcode reader to link with LIS. Altitude: > 2000 meter. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
| High Speed Refrigerated Centrifuge (for TB Lab) varible high speed (35,000 to 50,000 RPM 3000g (rotor speed), brushless induction motor powered, centrifuge with micro-processor controls and alphanumeric display. Maximum capacity: 1 to 2 litres. Self-diagnosis on strat up. Programmable speed regulators(independent RPM sensor) amd break. one touch selection for PRM and RCF. Digital display of all the parameters. Automatic rotor detcetion and program loading (approximately up to 30 program memory for storing protocols) Audio and visual alarm for deviation from set parameter and on program complection. Swinging bucket rotor with holder for 15 ml and 50 ml centrifuge tube with adapter. Hard and soft accelaration and breaking essenrial spears and accessories. Power supply: 220 - 230 V. should be supplied with voltage stablizer. | Pcs | √ | √
| A fully automated ELISA processor ,ELISA Plate Processing.The instrument needs to process the key steps of an enzyme immunoassay (EIA), from sample preparation to interpretation of optical density (OD) results. Components include:Pipettor used for diluting and/or dispensing samples and reagents. Incubation slots preset at 5°C, room temperature, or 37°C. 8-channel manifold washer. Microplate reader. Positive identification of samples, reagents, and controls is performed via the internal barcode reader.4 units 96 well microplates (pipetting,incubator,reading).1 unit independent washer &reader( two 8 nozzle manifolds).Liquid levels are checked and presence of possible clots detected. can be connected to Two-way connectivity to the laboratory information System (LIS). Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Pcs | √ | √
| 1000µl (Filtered pipette tips are used for protection against contamination. Application to PCR, sequencing and other techniques involving the use of vaporous, radioactive, biohazardous or corrosive materials.Virgin polyethylene filters.Ensure that all aerosols and liquids are prevented from entering the pipette.Packaged in rack and pre-sterilized.The filter pipette tip is best price and great quality. DNase/RNase free.Autoclavable. | 96pcs/rack
| 50µl (Filtered pipette tips are used for protection against contamination. Application to PCR, sequencing and other techniques involving the use of vaporous, radioactive, biohazardous or corrosive materials.Virgin polyethylene filters.Ensure that all aerosols and liquids are prevented from entering the pipette.Packaged in rack and pre-sterilized.The filter pipette tip is best price and great quality. DNase/RNase free.Autoclavable.) | 96pcs/rack
| 100µl (Filtered pipette tips are used for protection against contamination. Application to PCR, sequencing and other techniques involving the use of vaporous, radioactive, biohazardous or corrosive materials.Virgin polyethylene filters.Ensure that all aerosols and liquids are prevented from entering the pipette.Packaged in rack and pre-sterilized.The filter pipette tip is best price and great quality. DNase/RNase free.Autoclavable.) | 96pcs/rack
| HIV Antigen-antibody combo ELISA, 4th generation invitro diagnostic kit for the detection of p24 antigen and antibody in human serum or plasma.WHO prequalified, Reference Lab evaluated, sensitivity 100%, Specificity >99.4%, Analytical sensitivity: must be less than 60 picograms/ml and /or less than 0.050 IU/ml as per different standards. Recombinant and synthetic peptides for HIV-1 gp41, HIV-1 0 gp41, and HIV-2 gp36, Solid phase:12 x 8 microwell strips, Specimen type: Serum/plasma, Detection type:Combined detection of HIV-1/2 antibodies and antigen, Subtype detection: M N O sub types, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life : After receiving at central depot should be 12 months at 2-8 C. | 96tests/pack
Pathology Department: Biochemistry and Immunology
Reagent & Consumables for ELECSYS 2010
| AFP calset for ELECSYS 2010 | 4x1ml | √
| AFP for ELECSYS 2010 | 100test/ pkt | √
| Assay Cup for ELECSYS 2010 | 60x60 | √
| CEA calset for ELECSYS 2010 | 4xlml | √
| Clean cell for ELECSYS 2010 | 6x380 mL | √
| Cortisol calset for ELECSYS 2010 | 4x1ml | √
| Cortisol for ELECSYS 2010 | 100test/pkt | √
| Estradiol calset for ELECSYS 2010 | 4x1ml | √
| Estradiol for ELECSYS 2010 | 100test/ pkt | √
| Ferritin Calset for ELECSYS 2010 | 4x1ml | √
| Ferritin for ELECSYS 2010 | 100test/pkt | √
| Free PSA II calset for ELECSYS 2010 | 4x1ml | √
| Free PSA II for ELECSYS 2010 | 100test/ pkt | √
| FSH calset for ELECSYS 2010 | 4x1ml | √
| FSH for ELECSYS 2010 | 100test/pkt | √
| HCG -1 Beta calset for ELECSYS 2010 | 4x1ml | √
| HCG -1 Beta for ELECSYS 2010 | 100test/ Pkt | √
| LH calset for ELECSYS 2010 | 4x1ml | √
| LH for ELECSYS 2010 | 100test/Pkt | √
| Maintenance kit for ELECSYS 2010 | Pkt | √
| Preci control - U for ELECSYS 2010 | 4x3 ml | √
| Procell for ELECSYS 2010 | 6x380 mL | √
| Progesterone calset for ELECSYS 2010 | 4x1ml | √
| Progesterone for ELECSYS 2010 | 100test/ pkt | √
| Prolactin calset for ELECSYS 2010 | 4x1ml | √
| Prolactin for ELECSYS 2010 | 100test/ pkt | √
| PSA II calset for ELECSYS 2010 | 4x1ml | √
| Serum CEA for ELECSYS 2010 | 100test/ pkt | √
| SYS clean for ELECSYS 2010 | 5x100ml | √
| SYS Wash for ELECSYS 2010 | 500ml | √
| T3 calset for ELECSYS 2010 | 4x1ml | √
| T3 for ELECSYS 2010 | 200test/ pkt | √
| T4 calset for ELECSYS 2010 | 4x1ml | √
| T4 for ELECSYS 2010 | 200test/ pkt | √
| Testosterone calset for ELECSYS 2010 | 4x1ml | √
| Testosterone for ELECSYS 2010 | 100test/ pkt | √
| Total PSA II for ELECSYS 2010 | 100test/ pkt | √
| TSH calset for ELECSYS 2010 | 4x1ml | √
| TSH for ELECSYS 2010 | 200test/ pkt | √
| Assay tips | 30x120/ tips | √
| CA-125 calset | 4x1ml | √
| CA-125 Reagent | 100test/pkt | √
| CA 19.9 Calset | 4x1ml | √
| CA 19.9 Reagent | 100test/pkt | √
| CA 15.3 Calset | 4x1ml | √
| CA 15.3 Reagent | 100test/pkt | √
| Oncology Control Normal/ Abnormal | 4x3 m l | √
| Diluent Universal | 2x16ml | √
| Vit B12 Calset | pkt | √
| Vit B12 Reagent | pkt | √
| Folate Calset | pkt | √
| Folate Reagent | pkt | √
| Preci control Varia, Vit.B12, Ferritin, Folate | 4x3 mL | √
Reagent & Consumables for Electrolyte Analyzer (Easylyte)
| Chloride Electrode for Easylyte | pkt | √ | √ | √
| Daily cleaning solution for Easylyte | 6x15ml | √ | √ | √
| Internal filling solution for Easylyte | 125ml | √ | √ | √
| Membrane Assembly for Easylyte | pkt | √ | √ | √
| Na/K/Cl for Easylyte | Pkt | √ | √ | √
| Potassium Electrode for Easylyte | pkt | √ | √ | √
| Reference Electrode for Easylyte | pkt | √ | √ | √
| Sample detector for Easylyte | pkt | √ | √ | √
| Sample probe for Easylyte | pkt | √ | √ | √
| Sodium Electrode for Easylyte | pkt | √ | √ | √
| Solution Valve for Easylyte | pkt | √ | √ | √
| Tubing Kit for Easylyte | pkt | √ | √ | √
| Wash solution for Easylyte | 50ml | √ | √ | √
Reagent & Consumables for ELECTROLYTE Analayzer (DIAMOND)
| Chloride Electrode for DIAMOND | pkt | √ | √ | √
| Daily cleaning solution for DIAMOND | 6x15ml | √ | √ | √
| Internal filling solution for DIAMOND | 125ml | √ | √ | √
| Membrane Assembly for DIAMOND | pkt | √ | √ | √
| Na/K/Cl for DIAMOND | Pkt | √ | √ | √
| Potassium Electrode for DIAMOND | pkt | √ | √ | √
| Reference Electrode for DIAMOND | pkt | √ | √ | √
| Sample detector for DIAMOND | pkt | √ | √ | √
| Sample probe for DIAMOND | pkt | √ | √ | √
| Sodium Electrode for DIAMOND | pkt | √ | √ | √
| Solution Valve for DIAMOND | pkt | √ | √ | √
| Tubing Kit for DIAMOND | pkt | √ | √ | √
| Wash solution for DIAMOND | 50ml | √ | √ | √
Reagent & Consumables for Immunology Analyzer (ECIQ)
| Maintenance pack for ECiQ | Pkt |  | √
| Thyroid control for ECiQ | Pkt |  | √
| β-HCG II calibrator for ECiQ | Pkt |  | √
| β-HCG II for ECiQ | Pkt |  | √
| TSH calibrator for ECiQ | Pkt |  | √
| TSH for ECiQ | Pkt |  | √
| TT3 calibrator for ECiQ | Pkt |  | √
| TT3 for ECiQ | Pkt |  | √
| TT4 calibrator for ECiQ | Pkt |  | √
| TT4 for ECiQ | Pkt |  | √
| Universal Wash for ECiQ | Pkt |  | √
| Versa Tips for ECiQ | Pkt |  | √
| Micro sample cup for ECIQ | Pkt |  | √
| Signal Reagent for ECiQ | Pkt |  | √
Reagent for Photometer
| Albumin for Photometer | Pkt | √ | √ | √ | √
| Alkaline Phosphatase for Photometer | pkt | √ | √ | √ | √
| ALT (SGPT) for Photometer | pkt | √ | √ | √ | √
| AST (SGOT) for Photometer | pkt | √ | √ | √ | √
| Bilirubin Total & Direct for Photometer | pkt | √ | √ | √ | √
| Cholesterol for Photometer | Pkt | √ | √ | √ | √
| Creatinine for Photometer | Pkt | √ | √ | √ | √
| Gamma GTP for Photometer | pkt | √ | √ | √ | √
| Glucose for Photometer | pkt | √ | √ | √ | √
| Total protein for Photometer | Pkt | √ | √ | √ | √
| Triglycerides for Photometer | pkt | √ | √ | √ | √
| Urea for Photometer | pkt | √ | √ | √ | √
| Uric Acid for Photometer | pkt | √ | √ | √ | √
| Thermal paper roll for Photometer | Roll | √ | √ | √ | √
| Halogen lamp 20W 12V for Photometer | No. | √ | √ | √ | √
| Control normal | Pkt | √ | √ | √ | √
| Control high | Pkt | √ | √ | √ | √
| LE (SLE) slide test | 50test/ pkt | √ | √ | √ | √
| ASO Rapid Test Kit | Pkt | √ | √ | √ | √
| CRP Rapid Test Kit | pkt | √ | √ | √ | √
| RF Rapid Test Kit | pkt | √ | √ | √ | √
| Widal Slide Test | 4x5ml | √ | √ | √ | √
Vacuttainer Tubes
| Vacuette tube with glycolytic inhibition - Plastic | 100/pkt | √ | √ | √ | √
| Vacuette tube with gel,10ml capacity - Plastic | 100/pkt | √
| Vacuette tube with serum clot activator 10ml capacity - Plastic | 100/pkt | √
| Vacuette tube with serum clot activator 5ml capacity - Plastic | 100/pkt | √ | √ | √ | √
OTHERS
| Bar code label thermal paper Size 2"x1" (500 labels/roll) | Roll | √ | √
| Glucose (powder) | 250gm | √
| Urea (powder) | 250gm | √
| Creatinine (powder) | 250gm | √
| Aspartate Transaminase (lyopholized) | 100gm | √
| Alanine Pyruvate Tramsaminase (lyopholized) | 100gm | √
| Alkaline phosphatase (lyopholyzed) | 100gm | √
| Bilirubin (powder) | 250gm | √
| Albumin (Powder) | 250gm | √
| Lactate Dehydrogenase (lyopholyzed) | 100gm | √
| Micro Haematocrite Capillary Tube, Hepainized (75mm length) | 100 tube/pkt | √ | √ | √
| Capillary sealing compound (Haematrocrit tube): Sealing compound provides perfect seal. Capillary Tubes are put into the sealing compound with a twesting motion. | 10trays/ pkt | √ | √ | √ | √
| Almunium Crimp Neck auto Sampler vial cap | 100/pkt | √
| Amber Glass bottle 10ml with rubber cap | 500/pkt | √
Equipments
| ELECTROLYTE ANALYZER: Na, K, Cl, Li+, TCO2 CLIA Classification: moderate complexity Sample: Whole blood, Serum, Plasma, Urine Body fluids Sample Size: Approxmately 100 uL Performance: Na+, K+, Cl+, Li+, TCO2 Accuracy: 0.99 Analysis Time: Less than 60 sec (blood) 90 sec (urine) Data Storage: Patient result and QC results Normal, Low and High Calibration: Automatic or On - Demand Output: Printing of result, RS-232 optional, Autosampler tray optional Ambient Conditions: 15 - 32 °C, humidity: <85% Power: Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS. | Set | √
| ELECTROLYTE ANALYZER: Na, K, Cl, Classification: moderate complexity Sample: Whole blood, Serum, Plasma, Urine Body fluids Sample Size: Approximately 100 uL Performance: Na+, K+, Cl Accuracy: 0.99 Analysis Time: Less than 60 sec (blood) 90 sec (urine) Data Storage: Patient result and QC results Normal, Low and High Calibration: Automatic or On - Demand Output: Printing of result, RS-232 optional, Auto sampler tray optional Ambient Conditions: 15 - 32 °C, humidity: <85% Power: 220 - 240V, 50. power cord with round 3 pin top and appropriate UPS. | Set | √ | √ | √
| Chemistry Analyzer System: Discrete, multi-channel, fully automatic and computerized machine, truly random access, open system with onboard washing of reaction cuvettes. Test modes: stat, batch, random access. Assay modes: endpoint, kinetic, prozone check, rate A, B with blank Channels: on board chemistry and ions. Throughput: 400 – 600 tests per hour with ISE Serum indices: lipemic, Icteric, hemolysed Reagent Volume: minimum (R1- 150 – 450 uL, R2 – 20 – 450 uL ) Sample Volume: 2 – 45 uL, step by 0.1 uL Sample types: serum/plasma/urine/CSF & other fluids like peritoneal/pleural/pericardial etc. Reaction temperature: 37°C Stat positions: about 20 positions Cuvettes: Automatic washing cuvettes preferably non-disposable Printer: inbuilt with use of non thermal paper Interface: mono and bi-directional interface essential Cooling: Cool compartment for reagent 4 – 8 °C Method types: performs routine tests, enzymes, drugs assays, immunoturbidimetric assays, Auto dilution and auto rerun during the run Water Supply: On line water supply plant Wavelength: 340, 380, 412, 450, 505, 546, 570, 605, 660, 700, 740, 800 Absorbance: 0 – 3 Abs. Sample probe: Liquid level detection, clot detection and collision protection Reagent probe: Liquid level detection, collision protection and inventory checking Temperature: 15 – 30 °C Humidity: 35 – 80 % Power supplies: 200 – 240 V, 50/60Hz, 1200 VA with UPS Operation System: Windows XP, Interface RS – 232 Accessories: 1. Adaptors should be provided for different sizes of reagent bottles. 2. Start up reagent kits should be provided free of cost during the installation 3. List of spare parts and consumables should be submitted. 4. Warranty 5. Deionizer should be supplied with the analyzer 6. User training must be provided. | Set | √
| Chemistry Analyzer System: Discrete, multi-channel, fully automatic and computerized machines, truly random access, open system with onboard washing of reaction cuvettes. Test modes: stat, batch, random access. Assay modes: endpoint, kinetic, prozone check, rate A, B with blank Channels: on board chemistry and ions. Throughput: 200 – 400 tests per hour with ISE Serum indices: lipemic, Icteric, hemolysed Reagent Volume: minimum ( R1- 150 – 450 uL, R2 – 20 – 450 uL ) Sample Volume: 2 – 45 uL, step by 0.1 uL Sample types: serum/plasma/urine/CSF & other fluids like peritoneal/pleural/pericardial etc. Reaction temperature: 37°C Stat positions: about 20 positions Cuvettes: Automatic washing cuvettes preferably non-disposable Printer: inbuilt with use of non thermal paperInterface: mono and bi-directional interface essential Cooling: Cool compartment for reagent 4 – 8 °C Method types: performs routine tests, enzymes, drugs assays, immunoturbidimetric assays, Auto dilution and auto rerun during the run Water Supply: On line water supply plant Wavelength: 340, 380, 412, 450, 505, 546, 570, 605, 660, 700, 740, 800 Absorbance: 0 – 3 Abs. Sample probe: Liquid level detection, clot detection and collision protection Reagent probe: Liquid level detection, collision protection and inventory checkingTemperature: 15 – 30 °C Humidity: 35 – 80 % Power supplies: 200 – 240 V, 50/60Hz, 1200 VA with UPS Operation System: Windows XP, Interface RS – 232 Accessories: 1. Adaptors should be provided for different sizes of reagent bottles. 2. Start up kits should be provided free of cost during the installation. 3. List of spare parts and consumables should be submitted. 4. Warranty 5. Deionizer should be supplied with the analyzer. Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS |  |  | √
| PHOTOMETER Light Source: Quartz-halogen lamp 12V/20W Wavelength: 340, 405, 492, 510, 546, 578, 630nm Photometric Range: 0.0000 to 3.0000 ABS, Resolution: 0.0001 ABS Flow cell: Stainless steel body, Measuring volume: 32uL, Aspiration Volume programmable: 200 - 2000uL, linearity: 0-1000 units, Optical path: 10mm Analysis method: End point, kinetic, Fixed Time, Absorbance, Turbidimetry, Linear & non-linear standardization Parameter Setting: Method, Wavelength, Temperature, Reagent blank Y/N, sample blank Y/N, Dalay time, Measuring time, Reaction time, Absorbance limit, Aspiration volume, Standards, Linerity check & Unit for results Temperature Control: 25, 30, 37 °C optional, Precision: +/-0.1°C Display: Large LCD display Printer: Built-in thermal printer Interface: RS 232 serial port Operating Environment: 15 - 30°C, Humidity: 20 - 80% Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS. | Set | √ | √ | √ | √
| FULLY AUTOMATED RANDOM ACESS IMMUNO-ANALYZER: Analyzer should be based on the principle of Chemiluminescence reaction Technology OR enhanced Electrochemiluminescence. 1. Analyzer can be Bench top OR Floor base 2. Reagent Carrousel should have at least 10 positions for keeping reagent 3. Programming by loading 4. On board dilution facility 5. Throughput: 20 – 40 tests or more / hour 6. Assay time should be 9 – 18 minutes maximum 7. Reaction time: 37 °C 8. Temperature: 15 – 30 °C 8. Barcode reader for reagent, control and calibrators 9. STAT sampling position 10. On board cooling facility 11. Interface RS – 232 12. External printing facility 13. Onboard inventory management facility 14. Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS.S. Accessories: 1. Start up kits should be provided free of cost during the installation. 2. List of spare parts and consumables should be submitted. | Set |  | √
| BILIRUBIN METER: Measurement Parameters: 0 - 30 mg/dL (total bilirubin in serum) Light Source: 6V 1.5A tungsten lamp Wavelength: 461, 551 nm Formula: 2 wavelength differential formula Measurement cell: Capillary tube Sample Volume: 50 - 60uL Display: 3-digit LED Power: 220 - 240V, 50 Hz. power cord with round 3 pin top | Set | √ | √ | √
| Barcode Printer with Thermal paper size: 2'' W x 1'' L | Set | √ | √ | √
| Osmometer: Sample volume: 500 µl, Test time: 9 -13 minutes, Single sample capacity, Units: mOsm/kg H2O, Range: 280 - 350 mOsm/kg H2O, Linearity: <2% from a straight line, Temperature: 18 - 35°C, Humidity: 5 – 80 %, Power: 220 - 240V, 50 Hz. power cord with round 3 pin top. | Set | √ | √
| Centrifuge with Swing Out Rotor. Specification: Maximum speed: 8,500 rpm with 841 rotor; 3,900 rpm with 236 rotor. Electronic 30 minute timer and hold mode. Speed Range: 1,000 to 8,500 rpm in 100 rpm increments. Maximum Volume: 400 mL (8 holes x 50 mL ). Sound Level: 52 dBA. Dimensions: 10 3/4" H x 13 1/4" W x 15 3/4" D. | Pcs | √
| Electronic Magnetic Stirrer, Specification: Item Weight: 4.98 pounds. Voltage: 220-240V/60Hz. Scilogex 81111302 MS-Pro, Stainless Steel Plate, Maximum Stirring Speed: 1500 RPM. Overall Height: 3 inches, Overall Length: 11 inches, Overall Width: 6.25 inches, Plate diameter: 5.3", Digital Temperature Range: 0-340 Degrees Celsius (stirrer without heating system is preferred. If heating is present temperature adjustment is required) # different size of magnetic bar should be provided with magnetic stirrer. # any model magnetic stirrer which fulfill the above specification will be accepted. | Pcs | √
| Floor-Model Freez Dryer. Specification: Dimensions (approx): 32'W x 29.25"D x 33.5"H (outside). Electrical: 208 Volts, 60 Hz, 15A. Capacity: 14 Liters for 24 hours with a maximum low condenser temperature at -85C. This freezer dryer also has digital temperature and vacuum display. The FTS Dura-Dryer includes the TD2COT5000 Dura-Top bulk tray dryer. The tray dryer has non adjustable shelves. Feautures: Corrosion resistant construction results in minimal maintenance, Patended coil design prevents moisture carry-over to the vacuum pump. Clear acrylic door gives condenser. Visibility and automatic defrost reduces turnaround time. | Pcs | √
| Glass Vacuum Flask Kit - Heavy Duty. Capacity: 2 Liter. Filter Funnel with glass Frit base 500 ml, SS Funnel Clamp, SS Funnel Cover, Vac House. | Set | √
| Zippette Digital Bottle Dispenser: Specification: Capacity: 0-30ml. Delivery accruracy ± 0.3%, Precission better than 0.1% CV, adjustable volume by turning the top bottom, Chemical resistent can be used for wide range of equipments, Large digital volume display for easy and precise volume setting, display resolution of 100 ul, Autoclavable at 121°C2 bar. Simple user calibration via front panel buttons for GLP/ISO compliance. Safety right-angled spout design. Borosicate glass barrel protected with a transparent olyprophylene sleeve. Bubble free dispensing. | Set | √
| 13mm Flip off Vial Manual Crimper Hand Tool. Specification: Size arrange: 8mm - 32 mm. Gross weight: 1Kg. Color: Silver. Material: 100% stainless steel. Production Capacity: 20 - 30 bottle/min. | Pc | √
Pathology Department: Heamatology
| Cacl2 (0.02mol/L) | 10x15ml | √ | √ |  |  | Machine specific (closed system)
| Cuvette Rings 6 x 10 x 32 | 1920 tests/pkt | √ | √
| Washing Solution | 5 x 15 ml | √ | √
| Sample Cups 4 ml | 2 x 250 pcs | √ | √
| Sample Cups 2 ml | 1 x 500 pcs | √ | √
| Cleaner | 5 x 15 ml | √ | √
| Reagent Vials 5 ml with Caps | 50 pcs/pkt | √ | √
| Teflon Stirrer Bar | 10 pcs/pkt | √ | √
| Hemostat aPTT-EL | 6 x 4 ml | √ | √
| Hemostat D-Dimer | 2 x 1 ml | √ | √
| Hemostat Fibrinogen | 5 x 2ml | √ | √
| Hemostat Thromboplastin-SI | 6 x 10 ml | √ | √
| Hemostat Control Plasma Normal | 6 x 1 ml | √ | √
| Hemostat Control Plasma Abnormal | 6 x 1 ml | √ | √
| Hemostat Calibrator | 4 x 1 ml | √ | √
COAGULATION ANALYZER (Hemostat and Costat)
| Cacl2 (0.02mol/L) | 10x15ml |  |  | √ |  | Open system
| Cuvettes and stirrer for -HEMOSTAT model | 100/Pkt. |  |  | √
| Cuvettes and stirrer for- Costat model | 100/Pkt. |  |  | √
| PT determination reagent | 10x2ml |  |  | √
| APTT determination reagent | 10x4ml |  |  | √
| Normal control plasma for coagulation study | 10x1ml |  |  | √
| Abnormal control plasma for coagulation study | 10x1ml |  |  | √
MS95 & MS95s ANALYZER 5-PART DIFF
| Acti - diff for MS95s | 1.2ltr |  |  | √ |  | Machine specific (closed system)
| Acti - diff for MS95 | 1.2ltr |  |  | √
| Calibrator for MS95s (5-part) | 3ml/pkt |  |  | √
| Clot Filter | 2 pcs/Pkt |  |  | √
| Eo-Diff | 5ltr |  |  | √
| Hemoref | 250ml |  |  | √
| Isoflux | 20ltr |  |  | √
| 102 Control Diff (5-part) | 3ml each/Pkt. |  |  | √
| Transflux | 5ltr |  |  | √
| Cleaning Solution for MS Cell Counter | 250 ml |  |  | √
MS4 ANALYZER 3-PART DIFF
| Actidiff for MS4Se Blood cell counter | 1.2ltr |  |  | √ | √ | Machine specific (closed system)
| Actidiff for MS4e Blood cell counter | 1.2 ltr |  |  | √ | √
| Isoflux for MS4 Blood cell counter | 20ltr |  |  | √ | √
| Transflux for MS4 Blood cell counter | 5ltr |  |  | √ | √
| 102 Control Diff (3-part) | 4ml each for low, normal and high /Pkt |  |  | √ | √
| Cleaning Solution for MS Cell Counter | 250 ml |  |  | √ | √
| Thermal Paper | 1roll/pkt |  |  | √ | √
| Clot Filter | 2pcs/Pkt |  |  | √ | √
SYSMEX Kx 21 ANALYZER
| Cell Clean | 50ml |  |  | √ | √ | Machine specific (closed system)
| Cell Pack | 20ltr |  |  | √ | √
| Eight Check, 3WP, Low, Normal and High Control | 2ml each/Pkt |  |  | √ | √
| Thermal Printer paper roll | 1roll/pkt |  |  | √ | √
| Stromatolyzer - WH | 500ml |  |  | √ | √
AUTOMATED ESR ANALYZER (VESMATIC 20)
| Normal ESR Control | 1x9ml | √
| Abnormal ESR control | 1x9ml | √
| ESR vacutte tube for Vesmatic 20, DIESSE | 2x100 | √
| Paper roll for vesmatic 20 | 1roll/Pkt | √
Hematology general consumables
| Complete blood count QC material (three level-high, low and normal) | 2ml each/Pkt | √
| Rapid Malaria Test kit (ICT) | 50test/ pkt | √ | √ | √ | √
| G6PDH kit (Qualitative) | 20 test/ kit | √
| Reticulocyte stain solution (ready made solution) | 100ml | √ | √ | √ | √
| Safranin solution | 100ml | √
| Perl's stain solution (ready made solution) | 100ml | √
| New methylene Blue powder | 25gm | √ | √ | √ | √
| Sodium phosphate dibasic | 500g | √ | √ | √ | √
| potassium dihydrogen phosphate | 500g | √ | √ | √ | √
| Drabkin solution for Hb estimation | 5ltr | √ | √ | √ | √
| EDTA Dipotassium salt | 50gm | √ | √
| Giemsa stain | 100ml | √ | √ | √ | √
| May-grunwald-giemsa stain | 100ml | √ | √ | √ | √
| Leishman stain | 100ml | √ | √ | √ | √
| Sodium Metabisulphate | 50gm | √
| Trisodium citrate (3.8%) | 500g | √ | √ | √ | √
| Turks solution | 100ml | √ | √ | √
| Potassium ferrocyanide | 500g | √
| DLC counter | 1/pkt | √ | √ | √ | √
| Glass beads | 1kg/pkt | √ | √ | √ | √
| Heamocytometer with improved Neubauer rulings Counting Chamber | 1/pkt | √ | √ | √ | √
| Haemocytometer coverglass | 5/pkt | √ | √ | √ | √
| Sahli hemoglobinometer | 1/pkt | √ |  |  | √
| ESR tube with stand and scale- Westergren tube | 5/Pkt. | √ | √ | √ | √
| ESR vials with cap | 100 /Pkt. | √ | √ | √ | √
| Template for bleeding time (adult)-Depth: 1.0 mm. Length: 5.0 mm. Fully automated incision device for bleeding time determinations. Single use. Latex free. Sterile unless package is damaged or opened. Lot No. 2013-03 | 50/Pkt | √ | √ | √ | √
| Template for bleeding time (infant)-Depth: 0.5 mm. Length: 3-4.0 mm. Fully automated incision device for bleeding time determinations. Single use. Latex free. Sterile unless package is damaged or opened. | 50/Pkt | √ | √ | √ | √
| vacutainer tube coated with K2 EDTA for adults | 100 tube/Pkt | √ | √ | √ | √
| vacutainer tube coated with K2 EDTA for infants | 100 tube/Pkt | √ | √ | √ | √
| Buffered sodium citrate 0.109 M (3.2%) plastic vacutainer tube for adult | 100 tube/Pkt | √ | √ | √
| Buffered sodium citrate 0.109 M (3.2%) plastic vacutainer tube for infant | 101 tube/Pkt | √ | √ | √
| vacutainer tube (sodium/lithium heparin, no gel) | 102 tube/Pkt | √
| Timmer | 1pcs/pkt | √ | √ | √ | √
| Hb color scale | 50 test/pkt |  |  |  | √
Flocytometry
| CD7-FITC | 1 vial each | √
| CD5-FITC | 1 vial each | √
| CD19-PE | 1 vial each | √
| CD10-FITC | 1 vial each | √
| CD33-FITC | 1 vial each | √
| CD117-PE | 1 vial each | √
| CD34-PE | 1 vial each | √
| Thiazole Orange | 1 vial each | √
| BD FACSFlow solution | 20 liters | √
| BD FACS Lysing solution | 50 ml | √
| BD FACS Calibrite 3 Beads | 4 vials/ pkt | √
| BD Pharmlyse solution (10x) | 100 ml | √
| BD Perm II (10x) | 25 ml | √
| Hu CD 19PerCP | 100 test | √
| cCD3PE | 1 vial each | √
| cCD79a PE | 1 vial each | √
| CD45PCP | 1 vial each | √
| MsIg GI FITC | 1 vial each | √
| MSIg GI PE | 1 vial each | √
| CD13 PE | 1 vial each | √
| TDT- FITC | 1 vial each | √
| Kapa FITC | 1 vial each | √
| Lambda PE | 1 vial each | √
EQUIPMENT
| FULLY AUTOMATED COAGULATION ANALYZER: Parameters: Prothrombin Time (PT, PTR, INR), Activated Partial Thromboplastin Time (APTT), Fibrinogen (FIB), Thrombin Time (TT), Protein C, Protein S, ATIII; Coagulation Factors: Extrinsic Factors ( II, V, VII, X), Intrinsic Factors (VIII, IX, XI, XII); Detection Range : 50 - 450 mg/dL; Advanced Photo electronic-Double magnetic Testing; Built - in printer, Open system for reagents, 200 results memory capacity, No interference from abnormal samples (optical dense samples such as lipemic or icteric), Repeatability (CV%) <3%; Sufficient incubation stations for reagents and samples, Testing time 3 to 999 seconds, Inbuilt time for incubation. Should have the option to connect to LIS/HIS. Power supply: AC 220V- 230V, 50HZ; RS232 LIS/HIS connectivity port is mandatory. | Set | √ | √ | √
| FULLY AUTOMATED ESR ANALYZER: Features: Measures ESR according to Westergren Method. Works with EDTA blood. Built-in barcode reader. Sample volume: 1.3 ml whole blood. Analysis time: 30 minutes. The equipment should be complaince with infection control. | Set | √ | √
| ELECTRONIC DIFFERENTIAL COUNTER: – Technical data. Light weight and compact for easy transportation. This counter should have button to count all the important leukocytes, erythroid and megakaryocytic series. There should be short sound coming after every stroke of key and when the totalizer records to 100, 200, 300, etc. a long sound buzzs. The percentage of each type of cell is should be read directly from the window above each key. Function of CLEAR should be present. Function of CORRECTION should be present – To correct the possible counting error. Function of LOCK should be present – To keep the counting results from being touched carelessly. | Set | √ | √
| FULLY AUTOMATED HAEMATOLOGY ANALYZER - 5 part-diff, 21 parameter minimally: Differential counts:WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT, NEU%, LYMPH%, MONO%, EO%, BASO%, NRBC%, NEUT#, LYMPH#, MONO#, EO#, BASO#, NRBC#, RDW-SD, RDW-CV, MPV, P-LCR*, PDW*, PCT*, micro RBX%, macro RBC%, Mode Median, micro THR%, Macro THR%, Able to measure micro & macro platelets and RBC; Sample Volume: ≤20µl Sample Memory: 100-240 results with histogram The analyzer should have In-built printer, Barcode reader, External keyboard and QC soft ware’s for 3 control Throughput of the analyzer should be minimally 60-80 samples/hour. The Shelf life of reagents should be 12 -18 months. Analyzer should have auto Standby mode . Power Supply AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS must be provided. LIS/HIS connection should be there. Other accessories Program upgradable by CD/USB and results backup on CD/USB. Digital blood banks with programmable min/max ranges on all banks. Operating system windows based for user friendly. Cap piercing facilities for infants blood samples, Power supply Accuracy requirements for 5-part and 3-part Automated Hematology Analyzer WBC: Mean ±0.2x103/µl RBC: Mean ±0.04x106/µl PLT: Mean ±10 x103/µl Carry over requirements for 5-part and 3-part Automated Hematology Analyzer WBC: < 3% RBC: < 1.6% HGB: < 1.8% PLT: < 5% Precision requirements for 5-part and 3-part Automated Hematology Analyzer WBC: ≤3.5% RBC: ≤2.0% PLT: ≤6.0% Linearity requirements for 5-part and 3-part Automated Hematology Analyzer WBC: 0-100 x103/µl RBC: 0-7x106/µl HGB: 0.1-25.0 g/dl HCT: 10-67 % PLT: 0-999x103/µl | Set | √ | √ | √
| AUTOMATED HAEMATOLOGY ANALYZER: 3-part diff,18 parameter minimally. Differential counts:WBC, Lymp%, Lymp#, Mon%, Mon#, NEU%, NEU#, Gr#, RBC, Hb, MCV, Hct, MCH, MCHC, PLT, RDW-CV, RDW-SD, PDW, MPV, uRBC%, MacroRBC%, Pct, P-LCR, PDW, Mode, Median, uTHR%, Macro THR%. Sample Volume: ≤20µl Sample Memory: 50-200 results with histogram The analyzer should have In-built printer, Barcode reader, and External keyboard and QC soft wares for 3 controls Throughput of the analyzer should be minimally 40-60 samples/hour. The Shelf life of reagents should be 12 -18 months. Analyzer should have control Standby mode. Power Supply AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS must be provided. LIS/HIS connection should be there. Accuracy requirements for 5-part and 3-part Automated Hematology Analyzer WBC: Mean ±0.2x103/µl RBC: Mean ±0.04x106/µl PLT: Mean ±10 x103/µl Carry over requirements for 5-part and 3-part Automated Hematology Analyzer WBC: < 3% RBC: < 1.6% HGB: < 1.8% PLT: < 5% Precision requirements for 5-part and 3-part Automated Hematology Analyzer WBC: ≤3.5% RBC: ≤2.0% PLT: ≤6.0% Linearity requirements for 5-part and 3-part Automated Hematology Analyzer WBC: 0-100 x103/µl RBC: 0-7x106/µl HGB: 0.1-25.0 g/dl HCT: 10-67 % PLT: 0-999x103/µl |  |  |  | √ | √
| SWING ROLL MIXER: Speed - 0 to 60, Capacity - 50 to 100 tubes. The roller should have adequate vertical bend to ensure homogenous mixing. |  | √ | √ | √ | √
| FLOW CYTOMETRY for diagnosis of Haematological malignancies: At least 5-colored fluorescence, Air-cooled 488nm 15mw laser 10µ x 90µ light beam. Detector Range: 300nm-800nm. Flow-Cell Size: 250µ x 250µ BioSense. Max analysis rate: >3300 cells/sec. Min Cell Size, um: <0.5. Max analysis rate: >3300 cells/sec. Color dot density; contour, isometric, tomogram, overlay, legend and projection plots. PRISM displays; color precedence or color-blend gating; automatic Boolean logic gate creation; 32 gates per file, real-time statistics. Intel Pentium processor and 4 network transporters for digital signal processing. Signal-intensity display, biohazard-safe triple-fluid (sheath/cleanse/waste) system with automated functions and variable flow rate; auto setup; SYSTEM II software with acquisition, list mode, multigraph, reports and utilities. Features Fastest and Cell Tracker. |  | √
| HIGH PERFORMANCE LIQUID CHROMATOGRAPHY (HPLC) for hemoglobin typing- qualitative & quantitative: Injection range: Programmable from 0.1-1.0 μL at 0.1 μL increments, 1 μL to 2.5 mL at 1 μL increments. Precision: Typically < 0.5% RSD of peak areas from 3 μL, linearity > 0.999. Sample syringe size: 50, 100, 250, 500, 1000, 2500 μL. Carryover: < 0.02%. Flush volume: 0-2500 μL. External contacts: Two, start and stop time. Method storage (battery backed): Up to 20 methods including time programming. Stored methods: Editing possible during runs. Injection cycle time 2-3 samples per minute. Valve materials: In contact with Rheodyne® 7725 for analytical – stainless steel, Vespel®, ceramic. PEEK with sample with Rheodyne 9725 for bio applications – PEEK, Tefzel®, ceramic. Display: 7-line by 40-character backlit liquid crystal display. Integrated: For derivatization and dilution options serial communications – standard BCD. Communications interface – standard 5 relay outputs; 2 timed events, 2 inject, 1 start Ambient conditions: 10-30 °C, relative humidity 20-80%. Peltier sample cooling: 4-60 °C, ±1 °C. Power supply: AC220V, 50/60Hz. Should be supplied with UPS |  | √
| Barcode reader (or barcode scanner): The barcode reader should be able to read all the barcode and transfer the laboratory datas( values, histogram and scatter gram) accurately and precisely. | Pcs | √ | √ | √ | √
Pathology Department: Histo/cyto Consumable
| Cover glass 22mmx50mm | 20 pkts of 10gm each | √ | √ | √
| DPX Mountant | 250 ml | √ | √ | √
| Ethanol (99 - 100%) 500 ml | 500ml | √ | √ | √
| Haematoxylin Powder | 25 gm | √ | √
| Slide staining rack, SS. (wxdxh) 100mmx100mmx25mm, use with staining trough (accepts 76x26mm microscope slides) | Piece | √ | √ | √
| Slide staining trough (wxdxh) 112x112x50mm | Piece | √ | √ | √
| Slide tray of 20 slides ( Aluminium ) | Piece | √ | √ | √
| Sodium bi carbonate | 500gm | √ | √ | √
| Xylene extra pure. | 500ml | √ | √ | √
CYTO CONSUMMABLES
| Disposable Filter Paper for Cytocentrifuge (1ml) | 200/ pack | √ | √
| Specimen Chamber (1ml) for Cytocentrifuge | Piece | √ | √
| Specimen Chamber cap (1ml) for Cytocentrifuge | 200/Pkt | √ | √
| Specimen Chamber holder (1ml) for Cytocentrifuge | 200/Pkt | √ | √
| EA - 36 or 50 (Papanicolaous solution) polychromatic solution | 500ml | √ | √ | √
| Harris Haematoxylin Papanicolaus solution 1a | 500ml | √ | √ | √
| Orange Green-6, Papanicolaous solution, 2b Orange II Solution (contains methanol) | 500ml | √ | √ | √
| Polyethylene glycol 200 | 500 ml | √ | √ | √
| Dip quick I | 500gm | √
| Dip quick II | 500gm | √
HISTO CONSUMMABLES
| Alcian blue powder | gm | √ | √
| Aluminium Chloride Crystal | 500gm | √ | √
| Aluminium hydroxide | 500gm | √ | √
| Aluminium Potassium Sulphate | 500gm | √ | √
| Ammonical silver powder | gm | √ | √
| Aniline blue powder | gm | √ | √
| Base mold SS, (reuseable), Size - 25Wx32Dx6Hmm | Piece | √ | √
| Base mold SS, (reuseable), Size - 25Wx38Dx6Hmm | Piece | √ | √
| Base mold SS, (reuseable), Size - 25Wx32Dx13Hmm | Piece | √ | √
| Borax powder (Sodium tetraborate decahydrate) | gm | √ | √
| Carmine Powder | gm | √ | √
| Chromic acid powder | gm | √ | √
| Citric acid | 500 gm | √ | √
| Concentrated Nitric acid | ml | √ | √
| Congo red powder | gm | √ | √
| Chloral hydrate | 500gm | √ | √
| Di Sodium Hydrogen Orthophosphate | 500gm | √ | √
| Embedding Ring | Piece | √ | √
| Eosin Yellow Powder (water soluble) | gm | √ | √
| Ferric Chloride crystal | gm | √ | √
| Ferrocynide powder | gm | √ | √
| Formaldehyde Solution 37-40% | 500ml | √ | √ | √
| Frozen Section Medium ( Neg - 50) | 125 ml | √ | √
| Gold Chloride powder | 1 gm | √ | √
| Gold Nitrate powder | gm | √ | √
| Iron alum powder | gm | √ | √
| Light green powder | gm | √ | √
| Liquid Paraffin | 500 ml | √ | √
| Lithium Carbonate powder | gm | √ | √
| Magnesium Sulphate | 500 gm | √ | √
| Metanil yellow powder | 100gm | √ | √
| Methenamine powder | gm | √ | √
| Methyl violet | ml | √ | √
| Microtome Blade (Low profile disposible, S35) | 50/Pkt | √ | √
| Naphthol powder | gm | √ | √
| N-N-dimethyl formamide solution | ml | √ | √
| N-N-dimethyl, m-phenylenediamine powder | gm | √ | √
| N-N-dimethyl, p-phenylenediamine powder | gm | √ | √
| Nuclear fast red | gm | √ | √
| Oxalic acid powder | gm | √ | √
| Paraffin wax with Ceresin (Congealing point 60 degree C) | 500gm | √ | √
| Pararosaniline (dye) | gm | √ | √
| Periodic acid powder | gm | √ | √
| Perspex box (museum Jar) | Piece | √ | √
| Phloxine powder | 25 gm | √ | √
| Phosphomolybdic acid powder | gm | √ | √
| Phospho Tungstic acid powder | gm | √ | √
| Picric acid powder | gm | √ | √
| Schiff`s reagent | 125 ml | √ | √
| Silver Nitrate | 100gm | √ | √
| Sodium barbital powder | gm | √ | √
| Sodium Di-hydrogen orthophosphate | 500gm | √ | √
| Sodium Iodate | gm | √ | √
| Sodium Thiosulphate | 500gm | √ | √
| Sodium metabisulphite powder | gm | √ | √
| Sudan IV powder (Black) | gm | √ | √
| Tissue cassette ( White colour only ) | 100/Pkt | √ | √
| Tissue grossing knife (Razor) No. 260 Type( L) | Piece | √ | √
| Tri-Sodium Citrate (Na3C6H507.2H2O) | 500gm | √ | √
| Toludine blue powder | gm | √ | √
| Uranium (Uranyl) Nitrate | gm | √ | √
| Whatman, Lens cleaning tissue(only for histopathology),(8X11.5cm). | 20 wallets of 50 sheets | √ | √
| Wide mouth specimen container with cap ( Medical grade PC/PP 125 ml ) | 12 Nos/pack | √ | √
| Wide mouth specimen container wit cap ( Medical grade PC/PP 250 ml ) | 12 Nos/pack | √ | √
| Wide mouth specimen container with cap ( Medical grade PC/PP 500 ml ) | 6 Nos/pack | √ | √
| Wide mouth specimen container with cap ( Medical grade PC/PP 1000ml ) | 6 Nos/pack | √ | √
| Wide mouth specimen container with cap ( Medical grade PC/PP 2500 ml ) | Piece | √ | √
| Wide mouth specimen container with cap ( Medical grade PC/PP 5000 ml ) | Piece | √ | √
IMMUNOCHEMISTRY CONSUMMABLES FOR MANUAL METHOD
| Anti - Alpha-Fetoprotein ( AFP ) C3 , Ready to use | 6ml | √
| Anti - bcl-2 (Oncoprot), Ready to use | 6ml | √
| Anti - Beta HCG ( M94138 ), Ready to use | 6ml | √
| Anti - Calretinin (Polyconal), Ready to use | 6ml | √
| Anti - CD3 (T-Cell ) PS1, Ready to use | 6ml | √
| Anti - CD4(T4),Ready to use | 6ml | √
| Anti - CD5(4C7),Ready to use | 6ml | √
| Anti - CD7(D7-272),Ready to use | 6ml | √
| Anti - CD10(56C6),Ready to use | 6ml | √
| Anti - CD15 ( Tu9 ), Ready to use | 6ml | √
| Anti - CD19(EP169),Ready to use | 6ml | √
| Anti - CD20 (B cell ) L26, Ready to use | 6ml | √
| Anti - CD 30 (ki -1Ag) Ber-H2,Ready to use | 6ml | √
| Anti - CD34, Ready to use | 6ml | √
| Anti - CD45, cocktail LCA88, Ready to use | 6ml | √
| Anti - CD99(MIC2 antigen)-HO36.1.1, Ready to use. | 6ml | √
| Anti - CD117 ( T595 ), Ready to use | 6ml | √
| Anti - CDX-2 ( CDX2-88), Ready to use | 6ml | √
| Anti - CEA (CEA 88), Ready to use | 6ml | √
| Anti - Chromogranin A (LK2H10), Ready to use | 6ml | √
| Anti - Cytokrratin 7, (OV-TL1230), Ready to use | 6ml | √
| Anti - CK-19(RCK108) Ready to use | 6ml | √
| Anti - Cytokeratin 20 ( IT-Ks 20.8), Ready to use | 6ml | √
| Anti - CytoKeratin, Pan (AE1+AE3), Ready to use | 6ml | √
| Anti - Cytokeratin, High MW(34βE12) Ready to use | 6ml | √
| Anti - Desmin(D33),Ready to use | 6ml | √
| Anti - EMA (E29), Ready to use | 6ml | √
| Anti - GFAP(Polyclonal),Ready to use | 6ml | √
| Anti - Her 2/ Neu ( CB11 ), Ready to use | 6ml | √
| Anti - Kappa Light Chain ( K88 ), Ready to use | 6ml | √
| Anti - Lambda Light Chain ( HP6054 ), Ready to use | 6ml | √
| Anti - Melan-A(MART-1) clone A103 Ready to use | 6ml | √
| Anti - Melanoma (HMB45), Ready to use | 6ml | √
| Anti - MIB-1/Ki-67, Ready to use | 6ml | √
| Anti - Myogenin, Ready to use. | 6ml | √
| Anti - Myoglobin ( MG-1 ), Ready to use | 6ml | √
| Anti - Neurofilament ( NE-14, Ready to use | 6ml | √
| Anti - OCT4(EP143),Ready to use | 6ml | √
| Anti - Pax5(ZP007) | 6ml | √
| Anti - PLAP(PL8-F6),Ready to use | 6ml | √
| Anti - PSA (Polyconal) ErPr8, Ready to use | 6ml | √
| Anti - RCC(26),Ready to use | 6ml | √
| Anti - S-100 Protein (15E2E2), Ready to use | 6ml | √
| Anti - SMA (Smooth Muscle Actin) 1A4, Ready to use | 6ml | √
| Anti - TTF-1 (BGX-397A), Ready to use | 6ml | √
| Anti - Vimentin (V9 ), Ready to use | 6ml | √
| Antigen Retrival Citraplus HK081-5K | 250ml | √
| EZ-AR™ Common HK545-XOK | 1000ml | √
| EZ-AR™3 HK543-YOK | 500ml | √
| EZ-AR™4( RTU ) | 1000ml | √
| InSite™ ER Antibody ( ER 1D5) , Ready to use | 6ml | √
| InSite™ PR Antibody (PR88), Ready to use | 6ml | √
| Liquid Blocker super Pap pen | Piece | √
| Polymer HRP (DAB) Code no.QD400-60k sets | set | √
| Protease XXIV 3-Pack | 3x5ml | √
| Special positively charged glass with frosted end | 72/Pkt | √
| Tween® 20 (Polysorbate) for synthesis | 500ml | √
Cytology
| Microscope: 6V 30W halogen bulb. Focusing •Stage height movement by roller guide (rack & pinion). •Stroke per rotation: 36.8 mm. •Full stroke range: 25 mm. •Upper limit stopped by simplified pre-focusing dial. •Tension adjustment on coarse focus adjustment knob Revolving nosepiece Fixed quadruple nosepiece with inward tilt Binocular observation tube Field number 20 Tube inclination 30° Interpupillary distance adjustment range 48-75 mm Stage Size 188 (W) × 134 (D) mm Movement range 76mm X-direction × 50mm Y-direction Specimen holder Double slide holder Rubber grip Equipped as standard Condenser Type Abbe condenser, with built-in daylight filter N.A. 1.25 with oil immersion Aperture iris diaphragm .Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. Right- hand stage control. The microscope should be installed and demonstrated for fuctionallity test. Physically damaged microscope will not be accepted. After sales and services should be provided. | Set | √ | √ | √
| Fume Hood Structure: All steel (1.0mm cold-rolling steel or 1.0mm stainless steel) covered with electrostatic spraying after folding, welding, polishing, acid cleaning, phosphorization, and chemical resistance. Inner liner panel: 6mm compact panel, acid and alkali resistance, temperature: resistance, anti-moisture. Guide plate: reasonable and knock down guide structure. Concentrated and effective discharge air, letting off different poisonous gas. Operation Table board: lab-grade phenolic resins worktop, or epoxy resin worktop, or ceramic worktop and stainless steel as worktop. Window: 6mm thickness steel safe glass, balancing form elevation and subsidence, it can be stopped at any position by pulling. Lighting equipment: 30W antisepsis daylight lamp, lamp house is insulated with gas. Lightness on table not lowers than 400LUX. Operation panel: lighting switch, ON/OFF switch, two sets of 220V/10A versatile electric socket. Ventilation equipment: selecting rust-resistance, low wasting, low noise and famous fan and special wind tube. Noise can be reduced at about 55 DB. Accessories: vertical and single test cock: made by yellow brass, the surface is treated by EPOXY baking paint. Cup slot: acid and alkali resistance, PP water cup.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. It should be installed and demonstrated for functionallity test. After sales and services should be provided. | Set | √ | √
| Cyto Centrifuge Application: Bench top centrifuge for cytological specimens that transfers cells from suspension to a glass slide for microscopic review. Configuration: Bench top design. Process Capacity: Up to 12 disposable chambers per process. Speed: Adjustable from 500 to 2,500 rpm in 100-rpm increments. Programmable Times: Range from 1 to 60 minutes in 1-minute increments. Chamber Sizes: 1, 6, and 12mL disposable chambers. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for fuctionally test. Physically damaged machine will not be accepted. | Set | √ | √
Histology
| Automatic Tissue Processor, (With standard accessories) Specification : Temperature range of wax baths 45 degree to 65 degree celcius, excess temperature cut out 75 degree celcius Computerised freely selectable and storable programming facilities should be available. Easy editing and changing of programmes should be possible, even during a processing run. Infiltration time for each station should be separately programmable. Attached vacuum function device with fume control system. The vacuum device should be attachable both to the whole instrument as well as to individual stations Safely device for protection of drying of specimens in case of power failure. The bucket should go back inside the respective solution when power fails and not hang in mid air LCD display panel, audible alarm warning/error messages. Ergonomic control panel with full protected keyboard and LOC. Machine should also be able to cater to short time (quick) processing. Both immediate and delayed start functions should be available. Delayed start function atleast upto 72 hrs Interrupting an automatic process Carousel type with 12 stations (10 reagent stations, 2 wax baths). Metal containers at each station Capacity of each container 1.8 to 2 litres, Metal tissue baskets - 2 nos. capacity approx 80 - 100 cassettes to be run at one go. The tissue baskets should have less base diameter as compared to upper diameter to avoid basket getting stuck inside the containers. Firm bottom.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for fuctionally test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
| GROSSING STATION: Well ventilated grossing station (max 60inchs Wx32.5inch Dx92 inch H) plenum. Slotted rear of the workstation effectively that pulls fumes away from technician and exhaust to the outside. A rectangular collar mounted on top of workstation for connecting duct. Include ducts and exhaust fans. Fluorescent light and switch (mounted in the angled rear plenum behind a lens to maintain a bright work area. Work area divided between a sloped stainless work surface that drains into a stainless sink. A hand rinse spray nozzle located at the left end of the stainless work surface. A hot/cold faucet mounted behind the sink to rinse waste down the sink Include a garbage disposal can. Station has a lower base cabinet of welded stainless steel. Cabinet has two horizontal sliding doors to access the storage area and leveler. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for fuctionally test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
| ROTATORY MICROTOME: Rotary high precision manual microtome for paraffin sections. Mechanical automated feeding system. Integrated lockable hand wheel capable of being locked in any desired position. Section waste tray. Section thickness ranges 1to 60 micrometers. Trimming range 10 to 30 micrometer. Stable blade holder ( low profile disposable blade, S35) to ensure that no vibrations occur with section is being cut. Desirable to have arrangements to attach a cooling device for the paraffin block. Quick reclamp block holder(Universal block holder). The machine should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
| Tissue Floating bath (Thermostatic) Specification: Thermostatically controlled to 30 to 80˚C maximum Temperature control to +/- 0.5˚C Aluminium bath, polypropylene lid 450 W heating elements (include 2 spare heating elements), Pilot lights indicator for on/off and warn of overheating Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The tissue floating bath should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
| Slide Warmer, Temp. range 20-100˚C: Specification: Flattering table for stretching and drying of section tissue specimen. Temperature adjustable to 85ºC in 1ºC steps. Anodized dull black for good view of specimen. Large heating area accommodate 150-200 microscope glass slides.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer.The machine should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
| Cryostat: Temperature range: 0 to -35˚C, section thickness range 1 to 100 micron meter, Fine section thickness 1 to10 micron meter in 1 micron meter steps, Trimming section Thickness range 5 to 500 micron meter, Maximum specimen size 70x55mm, Total horizontal feed 28000micron meter, vertical cutting stroke 60mm, Sample retraction during return stroke 40micron meter can be deactivated, Power: 220-240volt, 50/60 Hz, Dimensions: (WxDxH) 640(w/o hand wheel) x 760x1200mm.The machine should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
| Tissue Embedding Station with wax Dispenser: Specification: Modular Paraffin Embedding Stations with Large heated work surface warmer with area for standard cable, free forceps, Cryo console with adjustable temperature of at least –50˚C, Isolated cold plate Paraffin dispenser with 3 litres or more paraffin reservoir. Manual and foot pedal controls for adjusting flow rate of paraffin. Should be programmable for preset time for melting paraffin and delay for weekend settings. Temperature setting for dispensing and thermal console. Cassette bath Mold warmer, Good illumination, should be able to operate on 22V. Temperature setting: 55 to 70˚C for cassette bath, mold warmer working area & forceps receptacle. Should supply 12-bulbs, 12- fuses, paraffin scraper-3 with each equipment. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for functionality test. After sales and services should be provided. Physically damaged equipment will not be accepted. | Set | √ | √
| Tissue Block cabinet (vertical): To store at least 20,000 paraffin blocks. Physically damaged tissue block cabinet will not be accepted. | Piece
Histology & Cytology
| Automatic Slide Stainer ( With standard accessories) Specification: Should be suitable for at least 40 slides per batch. It should be designed for continous operation program with the storage capacity of method upto 20 methods in memory. Should be programmable upto 50 steps in each method. Operation condition: Temperature- 10 to 40 degree celsius. Relative humidity- 36% to 85%. Atmospheric pressure-70 to 106 Kpa. Capacity/Fill volumes: Solution reservoirs total (26 numbers), wash reservoirs-(5numbers), Drying stations-( 1 number). Fill level reservoirs-620-650ml. Batch operation- upto 40 slides in two baskets( two baskets=1 basket set). Facillities: Water supply-tap water, water pressure-Dynamic pressure:0.098 to 0.441MPa. Max. Static Presure:0.833MPa. Water temperature: Less than 30 degree celsius. Drain type:Method-gravity drain. Capicity -More than 15 liters. Power supply: 220-240 V, 50/60 Hz. UPS to be provided. The machine should be installed and demonstrated for functionaly test. After sales and services should be provided. Physically damaged machine will not be accepted. | Set | √ | √
| Slide Cabinet: To store atleast 20,000 slides of 75x25mm slide - Horizontal with anodized aluminum slide tray of 350x210x35mm, uniformly slotted rows are numbered 1 to 50 to identify the slides. Physically damaged cabinet will not be accepted. | Piece | √ | √
| Automated Cover slipper: Self-contained, automated instrument for glass cover slipping. Very precise and bubble-free cover slipping Consistent cover slipping quality. High throughput for up to 400 slides per hour. Convenient batch or single slide mode. Flexible mountant type control for volume and flow rate. Automatic detection for broken coverslips. Special histopathology adapter. Full self diagnostic function. Efficient fume control system. Automated rinsing system of DPX by xylene at end of the work..Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for functionaly test. After sales and services should be provided. Physicallity damaged machine will not be accepted. | Set | √
pathology Department: General Lab
General Chemical
| Calcium chloride dihydrate | 500 gm | √ | √
| di-Potassium oxalate purified | 500 gm | √ | √ | √ | √
| Hydrochloric acid (Concentrated) | 500 ml | √ | √ | √ | √
| Immersion Oil | 125 ml | √ | √ | √ | √
| Potassium hydroxide pellets purified | 100 gm | √ | √ | √ | √
| Potassium Permanganate powder | 500 gm | √ | √
| Sodium chloride GR | 500 gm | √ | √ | √
| Sodium hydroxide pellets | 500 gm | √ | √
| Sodium hypochlorite solution GPR(5%) | 1 ltr | √ | √ | √ | √
| Sulphuric acid (concentrated) | 500 ml | √ | √ | √ | √
| Trichloroacetic acid | 500 gm | √
Glass / Plastic Wares
| Bottle-Reagent (glass capacity 250ml) | pcs | √ | √ | √ | √
| Bottle-Reagent (glass capacity 500ml) | pcs | √ | √ | √ | √
| Bottle-Wash, narrow neck 20 - 30mm , capacity 500ml. High chemical resistance and contamination free despensing | pcs | √ | √ | √ | √
| Bottle-aspirator polypropylene. 25 litres, PP tap and screw closure. Wide shoulder for easy carrying and pouring. Supplied with tap closure and leakproof. | pcs | √ | √ | √ | √
| Cover glass 22mmx22mm | 10g/pkt | √ | √ | √ | √
| Cylinder - graduated, glass (1000ml) | pcs | √ | √ | √ | √
| Cylinder - graduated, plastic (1000ml) | pcs | √ | √ | √ | √
| Cylinder - graduated, glass (100ml) | pcs | √ | √ | √ | √
| Cylinder - graduated, plastic (100ml) | pcs | √ | √ | √ | √
| Cylinder - graduated, glass (500ml) | pcs | √ | √ | √ | √
| Cylinder - graduated, plastic (500ml) | pcs | √ | √ | √ | √
| Flask conical, glass (250ml) with screw cap | pcs | √ | √ | √ | √
| Flask conical (500ml) with screw cap | pcs | √ | √ | √ | √
| Flask conical, glass (1000ml) with screw cap | pcs | √ | √ | √ | √
| Flask conical, glass (2000ml) with screw cap | pcs | √ | √ | √ | √
| Flask conical, glass (3000ml) | Pcs | √ | √ | √
| Flask Volumetric, glass 50ml | pcs | √ | √ | √ | √
| Flask Volumetric, glass 100ml | pcs | √ | √ | √ | √
| Flask Volumetric, glass 500ml | pcs | √ | √ | √ | √
| Flask Volumetric, glass 1000ml | pcs | √ | √ | √ | √
| Test tube without rim, glass (10x75mm) | pcs | √ | √ | √ | √
| Test tube without rim, glass (12x100mm) | pcs | √ | √ | √ | √
| Test tube without rim, glass (12x75mm) | pcs | √ | √ | √ | √
| Test tube without rim, glass (15x125mm) | pcs | √ | √ | √ | √
| Test tube with rim, glass (15x150mm) | pcs | √ | √ | √ | √
| Volumetric Glass pipette 1ml, calibrated for delivery | pcs | √ | √ | √ | √
| Volumetric Glass pipette 2ml, calibrated for delivery | pcs | √ | √ | √ | √
| Volumetric Glass pipette 5ml, calibrated for delivery | pcs | √ | √ | √ | √
| Volumetric Glass pipette 10ml, calibrated for delivery | pcs | √ | √ | √ | √
| Test tube holder brass | pcs | √ | √ | √ | √
| Beaker, (polypropylene-transparent, squat form, 1000ml, can be used continuously at 100˚C with spout and graduations) | pcs | √ | √ | √ | √
| Beaker, (glass, squat form, 2000ml, can be used continuously at 100˚C with spout and graduations) | pcs | √ | √
| Tubes, Centrifuge, conical. 15ml tubes with cap, graduated in 0.5mL subdivisions from 2 to 14.5mL, 17x119mm polypropylene. | 50pcs/pkt | √ | √ | √ | √
| Tubes, Centrifuge, conical. 50ml tubes with cap, graduated in 0.5mL subdivisions from 2 to 49.5mL, 30x119mm polypropylene. | 50pcs/pkt | √
| Microscope slide, single frost, thickness 1 - 1.2mm. Dimensions: 75x25mm. The raised opaque coating can easily be marked with a pencil or ink, resistant to every commonly used solvent and it prevents the slides from adhering and becoming scratched. Ready to use | 50slide/pkt | √ | √ | √ | √
| Microscope slide, dimension 76mmx26mm, thickness 1 - 1.2mm, made from extra white glass with finely ground edges.Ready to use | 50slide/pkt | √ | √ | √ | √
| Capillary Tube, Heparinized, hematocrit (75mm length) | 100/pkt | √ | √ | √ | √
| Capillary Tube, non-heparinized, hematocrit (75mm length) | 100/pkt | √ | √ | √ | √
| Capillary sealing compound (Haematrocrit tube): Sealing compound provides perfect seal. Capillary Tubes are put into the sealing compound with a twesting motion. | 10trays/ pkt | √ | √ | √ | √
| Brushes, test tube. Head diameter: 11mm. For tubes diameter: 10 to 13mm | pcs | √ | √ | √ | √
| Brushes, test tube. Head diameter: 29mm. For tubes diameter: 16 to 25mm | pcs | √ | √ | √ | √
| Parafilm, roll (width 2 inches x length 250 ft) | pcs | √ | √ | √ | √
| Pipettes, Pasteur, maximum bulb capacity 3.5ml | 100/pkt | √ | √ | √ | √
| Basket, polished stainless steel for test tube: Dimensions (wxdxh),mm. 150x150x150 | pcs | √ | √ | √ | √
| Basket, polished stainless steel, test tube: Dimensions (wxdxh), mm. 200x200x200 | pcs | √ | √ | √ | √
| Basket, polypropylene, test tube, draining: Dimensions (wxdxh),mm. 400x300x100 | pcs | √ | √ | √ | √
| Test tube rack, polypropylene. Hole dia: 16mm. Nos. of hole: 60. Array: 5x12. Height: 70mm. | pcs | √ | √ | √ | √
| Test tube rack, polypropylene. Hole dia: 13mm. Nos. of hole: 72. Array: 6x12. Height: 57mm. | pcs | √ | √ | √ | √
| Pipettor tips, volume 5 -50 uL, optimal tightness fitting | 1000/ pkt | √
| Pipettor tips, volume 50-200 uL, optimal tightness fitting | 1000/ pkt | √ | √ | √ | √
| Pipettor tips, volume 200 -1000 uL, optimal tightness fitting | 1000/pkt | √ | √ | √ | √
| Pipettor tips, volume 1000 -5000 uL, optimal tightness fitting | 500/ pkt | √ | √
| Pipettor Stand: Holds 5 to 7 pipettors, fits most brands | pcs | √ | √ | √ | √
| Diamond pen | pcs | √ | √ | √ | √
| Filter paper, qualitative: diameter 110mm, provides fine retention and slow flow rate, suitable for general filtration of many types of samples. | 100/pkt | √ | √ | √ | √
| Filter paper sheet 46cmx57cm. provides fine retention and slow flow rate, suitable for general filtration of many types of samples. | 500/pkt | √ | √ | √
| Blotting paper. (Highly absorbent) effectively absorbs liquids including testing or diagnostic test kits, spill control and clean up of toxic and non-toxic sustances. Minimum dimension of 12 inches width x 5 metres length and thickness 0.21mm | rolls | √ | √
| Paper, lens cleaning tissue, high quality lens cleaning tissues, 80mmx100mm | 50sheet/pkt | √ | √ | √ | √
| Spirit lamp, brass | pcs | √ | √ | √ | √
| Digital count down timer.2 count down timers each for up to 23hr 59min 59sec. 12/24hr clock. Extra loud alarm. Bench top standing or wall mounted. Memory function. Powered by AAA battery (supplied) | pcs | √ | √ | √ | √
| Blood Tourniquet. Highly flexible, one hand operation with quick and secure release buckle. | pcs | √ | √ | √ | √
| Sample carrying tray with handle, autoclavable Polypropylene Tray to carry reagents or laboratory apparatus from one place to another, single handedly. | pcs | √ | √ | √ | √
CLINICAL
| Acetic acid glacial | 500ml | √ | √ | √ | √
| Ammonium sulphate | 500g | √ | √
| Barium chloride | 500g | √ | √ | √ | √
| Benedicts reagent | 500ml | √ | √ | √ | √
| Ehrlichs aldehyde reagent | 100ml | √ | √ | √ | √
| Fouchet reagent | 100ml | √ | √ | √
| Liquid ammonium solution | 500ml | √ | √
| Pregnancy test kit (sensitivity as low as 20mIU/ml hCG) | 50strips/ pkt | √ | √ | √ | √
| Sodium nitroprusside | 100g | √ | √ | √
| Sulphosalicylic acid | 100g | √ | √ | √ | √
| Sulphur powder | 500g | √ | √ | √ | √
| Hydrogen peroxide | 500 ml | √ | √ | √ | √
| Urinary strip 10 parameters | 100 strips/ Vial | √ | √ | √
| Urinary strip 2 parameters (albumin and sugar) | 100strips/ Vial | √ | √ | √ | √
| LE (SLE) slide test | 50 test/ pkt | √ | √
| Occult Blood test kit | 50strips/ pkt | √ | √ | √ | √
| Semen diluting fluid | 100ml | √ | √ | √
| pH indicator paper (2-7pH) | 50strips/ pkt | √ | √ | √ | √
| pH indicator paper (8-10.5pH) | 50strips/ pkt | √ | √ | √ | √
| Slide mailer (5 Slides) | Piece | √ | √ | √ | √
| Ayre`s Spatula | pcs | √ | √ | √ | √
| Coplin Jar |  | √ | √ | √ | √
EQUIPMENT (Every equipment supplied must be installed with end user training and demonstrated for functionality)
| Balance: Readability: 0.01g. Max. Load: 610 g. Repeatability: 0.01g. Linearity: 0.02g. Weighing pan: dia160mm. Power supply: AC 220-240V, 50/60Hz. Power cord with round three pin top | Set | √ | √ | √ | √
| Bench top Centrifuge: Swing out rotor with bucket with holders for 12x100mm test tube, and 15 ml tubes. Maximum RPM: 2000 - 5000 RPM. Maximum relative centrifuge force: 4640 x g. Main control: LCD microprocessor controlled. Power supply: AC 220-240V, 50 HZ. Power cord with round three pin top. Operating environmental condition: -10˚C to +40°C, RH >80%. | Set | √ | √ | √ | √
| Incubator: Temperature range: Ambient + 5°C to 70°C. Temperature control: Digital P.I.D. controller. Display: Dual LED display. Capacity: 100 - 300 liters. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √ | √ | √ | √
| Bench top Centrifuge: Fixed angle rotor with holders for 12x100mm test tube, and 15 ml tubes. Maximum RPM: 2000 - 5000 RPM. Maximum relative centrifuge force: 4640 x g. Main control: LCD microprocessor controlled. Power supply: AC 220-240V, 50 HZ. Power cord with round three pin top. Operating environmental condition: -10˚C to +40°C, RH >80%. | Set | √ | √ | √ | √
| Deep Freezer. with microprocessor controlled temperature regulation with digital display. Temperature Range: - 20 to -40ºC. Capacity: 250 -300 Liters. Power: 220-240V, 50/60Hz.Power cord with round three pin top. Should be supplied with appropriate storage racks & Voltage stabilizer. Operating environmental condition: - 10 to +40°C, RH >80%. | Set | √ | √ | √
| Deep Freezer. Vertical type with microprocessor controlled temperature regulation with digital display. Temperature Range: - 70ºC. Capacity: 250 -300 Liters. Power: 220-240V, 50/60Hz.Power cord with round three pin top. Should be supplied with appropriate storage racks with individual storage compartments & Voltage stabilizer. Operating environmental condition: - 10 to +40°C, RH >80%. | Set | √ | √
| Distillation plant: (mono distiller) with all the required accessories and suitable for bench mounting. Thermostatic low water cut-off. Destillate withdrawal through outlet at the condenser. Evaporator easily accessible for cleaning purposes. economical energy consumption by distillation of the heated cooling water. Throughput: 4 liters/hour. Power supply: 220-240V 50/60Hz. | Set | √ | √ | √ | √
| Haematocrit Centrifuge. High speed bench-top centrifuge for haematocrit determinations with brushless drive. Sedimentation is completed at 16,060 x g in a little over 6 minutes. Quick entry foil keypad. Easy to read digital display. Easy entry and display of current parameter values. Metal housing. The lid of the rotor is the evaluation disk and cover in one. The haematocrit value can be read off the lid / evaluation disk after centrifugation. Motor overheating protection. Imbalance switch-off. Timer: 99 minutes max. Maximum speed: 13,000 rpm. (16,060 x g.) in 100 rpm increments. Noise level: 58 dB(A). Operating environmental condition: -10 to +40˚C, RH >80%. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √ | √
| Pipettors, variable volume 5 - 50µl, full volume range, thumb press tip ejector, calibration simple to perform by user. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
| Pipettors, variable volume 50 - 200µl, full volume range, thumb press tip ejector, simple calibration to perform user friendly. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
| Pipettors, variable volume 200-1000µl, full volume range, thumb press tip ejector, simple calibration to perform and user friendly. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
| Pipettors, variable volume 1000 - 5000µl, full volume range thumb press tip ejector, simple calibration to perform and user friendly. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
| Microscope, binocular: (for routine work) Illumination inbuilt,Right-handed stage control. Objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X. Complete with power cord, 220/230 - 50/60HZ. | set | √ | √ | √ | √
| Microscope: Multi head with CCTV system and LCD monitor. Illumination: with objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X Complete with round 3 pin power cord. Power supply: 220V, 50/60HZ. Right-handed stage control. Should be supplied with micophotography attachement and latest computer. | Set | √
| pH Meter: pH Range: 0.00 - 14.00pH. Resolution: 0.01pH. Relative Accuracy: +/- 0.01pH. Millivolt Range: -1999 to 1999mV. Temperature Range: 0.0 - 100.0.Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin round top | Set | √ | √
| Refrigerator with glass door for storage of laboratory reagents: Capacity: 1000 to 1300 liters. Temperature Range: 0 to+10° C. Temp. Accuracy: +/- 0.3 at 4°C. Temp. Controller: Digital P.I.D. controller. Display: Dual LED display. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √
| Refrigerator with glass door for storage of laboratory reagents: Capacity: 350 to 800 Liters. Temperature Range: 0 - 10° C. Temp. Accuracy: +/- 0.3 at 4°C. Temp. Controller: Digital P.I.D. controller. Display: Dual LED display. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √ | √ | √
| Vortex mixer. Speed Range: 0 - 3000 RPM. Operation Mode: Touch on and constant on Head: Rubber Mixing Cup. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top | Set | √ | √
| Water bath: Temperature range: Ambient + 5°C to 99°C. Temp. Accuracy: +/-2˚C at 37˚C. Temp. Controller: Digital P.I.D. controller, 4Zone PID. Display: Dual LED display. Capacity: 20 - 30liters. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √ | √ | √
| Hot air oven: Temp. Range: Ambient + 5°C to 220°C. Temp. Control: Digital P.I.D. controller. Display: Dual LED display. Capacity: 50 - 100 liters. Power cord with round three pin top, Voltage AC 220V-240V. | Set | √ | √ | √ | √
| BILIRUBIN METER: Measurement Parameters: 0 - 30 mg/dL (total bilirubin in serum) Light Source: 6V 1.5A tungsten lamp Wavelength: 461, 551 nm Formula: 2 wavelength differential formula Measurement cell: Capillary tube Sample Volume: 50 - 60uL Display: 3-digit LED Power: 220 - 240V, 50 Hz. power cord with round 3 pin top | set | √ | √ | √
| Hemoglobin color scale | pkt | √ | √ | √ | √
7 | Pediatric Department
| Bedside SpO2 patient monitoring system: Multicolor display screen with a black background for contrast, Backup audible alarm, Multiple-language graphical user interface to support variety of users, 96-hour trend memory with data captured every four seconds,Simple, intuitive operation with a space-saving design Displays lethysmographic waveforms, pulse amplitude and current measured SpO2 and pulse rate (wrap type sensor probe with 9 pins) | Set | √ | √ | √
| CP Monitor (Intensive care monitors) – Pediatric: NIBP, SPO2, ECG, Respiratory rate, heart rate & Temperature, Battery, 15" - 17" colour TFT Display, 6 channels, Modular in design, should have module for measuring, 5 lead ECG with electrocautery filter, multilead ST segment analysis, lethal arrhythmai alarms with 5 lead ECG cable Set, Spo2 measurement with plethysmograph-Neonate and pediatric Spo2 sensors, pitch should vary with saturation, NIBP measurement with pediatric and neonate cuffs, Respiration measurement by impedance method, Provision for invasive pressures ( CVP and IBP) simultaneously with 20 disposable transducers and 2 reusable pessure CVP and IBP pressure cables , Temperature measurement, Co2. | Set | √ | √
| Capnography (ETCO2) monitor, Pediatric : Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. Suitable voltage stabilizer. EtCO2 Measurement Range: 0-100 mmHg (display dependent); 0-13%; 0-12.5kPa. EtCO2 Accuracy 0-40 mmHg ±2 mmHg, 41-70 mmHg ±5%, 71-100 mmHg ±8%. LoFlo 0-40 mmHg ±2 mmHg, 41-70 mmHg ±-5%, 71-100 mmHg ±8%. ±12% for respiration rate above 80 BPM.EtCO2 Resolution: 1 mmHg 0-100 mmHg. EtCO2 Stability: Short Term Drift: Drift over four hours shall not exceed 0.8 mmHg maximum. | Set | √ | √
Long Term Drift: Accuracy specification will be maintained over a 20-hour period after zeroing. EtCO2 Noise: RMS noise of the sensor less than or equal to 0.25 mmHg at 7.5% CO2 EtCO2 Rise Time (10-90%): Mainstream: Less than 60 ms (Adult reusable or single patient use); Less than 50 ms (Infant reusable for single patient use). Sidestream: less than 200 ms. Respiration Rate (RR) Range: 0-150 breaths per minute (BPM). Respiration Rate (RR) Accuracy: ±1 breath. Compensations: Barometric pressure 550 -780 mmHg (EtCO2 Alarm Limits: User-selectable/High 5 to 100 mmHg, Low 0 to 95 mmHg/OFF. Respiration Rate (RR) Alarm Limits: User-selectable, High 5 - 150 respirations per minute, Low 0 - 100 respirations per minute/OFF. Airway Adapter Deadspace: Environmental: Operating Temperature: -10° to 40°C, RH >80%, height 2000m.
| Child ACLS Manikin with Interactive Arrhythmia Simulator : PR; Palpable and visual landmarks; • Fully articulated head, neck and jaw; Airway Management; • Realistic anatomy of the mouth, tongue, oral pharynx, larynx, epiglottis, vocal cords, trachea and esophagus; • Separate left and right lungs for auscultation; • Suctioning capabilities; • Oral, nasal and digital intubation capabilities; • Inflatable stomach bladder to indicate esophageal intubation; I.V. Injection (at arm): • Realistic size of an child’s veins; • Articulated at the biceps for antecubital and dorsal access; • Realistic flashback from a pressurized system; Blood Pressure Measurement: Instructor determines systolic and diastolic levels, heart rate and sound volume; Speaker in arm reproduces real blood pressure sounds; 5 Korotkoff phases can be turned on and off; Auscultatory gap can be turned on and off; | Set | √ | √
Intraosseous Infusion: • Palpable landmarks include the patella, tibia and tibial tuberosity; - Pressurized system allows aspiration of fluid; Femoral Access: • Palpable arterial pulse; • Palpable symphysis pubis and anterior superior iliac spine; • Realistic flashback from a pressurized system; Defibrillation Chest: • Internal load box absorbs full strength of every shock;• Manual, semi-automatic, automatic defibrillation; • Monitor manikin like a real patient at 4 ECG sites and 2 defibrillation sites; • Compatible with all standard brands and types of defibrillators and ECG-monitors; Interactive ECG-Simulator: This external unit allows defibrillator and pacer training directly on the manikin. (A training without a manikin is also possible). Various ECG rhythms can be generated and displayed on an ECG-monitor (not included) for diagnostic and defibrillation training. All arrhythmias can be displayed in adult or paediatric mode so that a total of 34 rhythms are available.
| Digital weighing scale capability to measure 1 milligram to 1000 milligram with flat top | Set | √ | √ | √
| Digital Pediatric NIBP and Spo2 (Portable) monitor, Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. LCD display with heart rate, Spo2 , pulse rate, min. 2 hours battery back-up, with all accessories. | Set | √ | √ | √
| Endotracheal tube pressure cuff inflator | Set | √ | √
| Electro chest physiotherapy device (vibrator/occillator) for Pediatric | Set | √ | √
| ICU Pediatric Bed : Height-adjustable electric pediatric intensive care bed (4 sections, with column motors, all side tilt) | Set | √ | √
| ICU Ventilator (Respiratory Ventilator) Pediatric : 1. Should have facility for Invasive and Non-Invasive ventilation; 2. Microprocessor Control suitable for Pediatric (one onwards) and adult ventilation; 3. Electromagnetic Compatible Hinged arm holder for holding the circuit; 4. Should have built in touch color screen TFT display of minimum 10” or more for display; Of waveforms and Monitored value; 5. Should have inbuilt facility to upgrade with EtcO2; 6. Facility to Measure and display:- a) Status indicator for ventilator mode; b) Battery indication; c) Pressure Vs time Vs volume Vs time, flow Vs time 3 curves/ waveforms; d) Alarm Setting; 7. Automatic compliance and leakage compensation for circuit and ET Tube. 8. Should have facility of log book, for events and alarms with date & time; 9. Should have following Settings; a) Tidal volume (Minimum at least 20 ml, Maximum up to 2000ml); b) Inspiratory Pressure (0 upto 80 cm of H20); c) Respiratory rate 1 to 80 bpm; d) Apnoea back up rate; e) CPAP/PEEP; f) Pressure support; g) Fi02; h) Pause Time; i) Pressure & flow Trigger; j) Inspiratory flow up to 120 Lpm; Temperature of delivered Gas on LED display; c) Temperature should be adjustable; d) Jar should be autoclavable; 17. Quality Certification : Valid CE/BIS/US FDA; 18. Demonstration of the quoted model is must, preferable on site; 19. Nebulization assembly compatible with ventilator and circuit; 20. Should have interface facility; 21. Flow sensor-Should have life more than 1 year; 22. Expiratory Unit- Life should be more than 3yrs; 23. Data storage facility for at least 24hrs; 24. Internal rechargeable battery at least 60 min. backup; 25.; 26. Should have flow sensors having long life and the company shall specify the life cycle and the cost of the flow sensors at the time of quoting the tender; 27. CMC/ AMC for the atleast 5yrs and Cost of consumables spares; 28. Source :Indigenous/Imported; 29. Warranty 2-5 years from the date of installation; 30. Standard Accessories along with: a) Patient breathing circuit of silicone for Paediatric (reusable) 3 Sets; Disposable breathing circuits 100 nos. | Set | √ | √
Peak pressure, mean pressure, CPAP/PEEP, expiratory tidal volume, total frequency, spontaneous frequency, lung function monitoring including C20/C,R,C,Tc, RVR, have CE mark and FDA approval for use on nenonates, Standard scope of supply must include Basic Unit, Trolley, 2 silicone Hose Sets, flow sensors/Exp. Valves - 2 Sets, each unit), O2 cell, oxygen connecting hose, F & P servo control Humidifier with two reusable chamber, Nebuliser circuit support arm, Instruction manual, input power : AC 220-240 volts, suitable voltage stabilizer, operational environmental condition : humidity above 80%, temperature -5 to +40, altitude 2000m, backup source, servo controlled stabilizer to be included. Power supply: 220V-230V, Power cord of round 3 pin type plug top , Suitable UPS.
| Medication cart : Material: Aluminum -Alloy, ABS plastic & stainless steel, Overall size: 850mm*520mm* 1010mm, The whole trolley is including 6 layers of drawers, can be pulled out from both sides, self locked when closed. Totally 50medicine boxes on all drawers. Other components: 2 pcs of garbage bins, 1 pc of transparent file box, 1pc of sliding side board and 1 set of five hooks hanging device | Set | √ | √ | √
| Nasal oxygen high flow system with integrated humidifer with flow generator with regulatable FiO2 with nasal cannula and circuit for Pediatric | Set | √ | √
| Oxygen hood : ideal round shape & flat top enhancing visibility & optimum circulation of oxygen, Polycarbonate, unbreakable, autoclavable & should have oxygen inlet nozzle : Extra Large- 13.0”L x 10.4”W x 8.5”H; Nele opening dimension: 6.3”W x 5.0”H | Piece | √ | √ | √ | √
| Oxygen hood : ideal round shape & flat top enhancing visibility & optimum circulation of oxygen, Polycarbonate, unbreakable, autoclavable & should have oxygen inlet nozzle : Large- 13.0”L x 10.4”W x 8.5”H; Nele opening dimension: 6.3”W x 5.0”H | Piece | √ | √ | √ | √
| Oxygen hood : ideal round shape & flat top enhancing visibility & optimum circulation of oxygen, Polycarbonate, unbreakable, autoclavable & should have oxygen inlet nozzle : Medium- 7.0”L x 7.0”W x 5.4”H; Nele opening dimension: 3.9”W x 2.5”H | Piece | √ | √ | √ | √
| Oxygen analyzer: Range 0-100% oxygen. Display accuracy 0.1%, electrical connector: battery operated, operating temperature 0-50° C, output signal Oxygen in %, sensor type galvanic (electro chemical), response time < 5 sec. to T90, accuracy ± 1 % at full scale, repeatability ±1% of signal, temperature compensation: compensated, maximum storage temperature -20° C to + 60° C, power: alkaline batteries, Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. | Piece | √ | √
| Pediatric Otoscope | Set | √ | √ | √
| Pocket pen tourch for pupil examination with Pupil Gauge | Piece | √ | √ | √
| Stethoscope - Paediatric | Piece | √ | √ | √ | √
| Ultrasonic nebuliser, heavy duty jet nebuliser: heavy duty and provided with support for tubings, fitted with an electronic timer and should operate continuously. The output of drug delivered should be electronically adjustable between 0.1-0.5ml/min. The medium size of particles generated should be 5 microns, facility for automatic stoppage when liquid drug is exhausted. The nebulizer chamber and tubing should be sterilizable. 4 masks (2 each for infants and older children) must be supplied with the unit 2 extra chambers must be supplied with each unit. Input power: AC 220-240 volts, servo controlled stabilizer to be included. | Set | √ | √ | √
| Beneview T5 CP monitor for Spo2 probe : wrap type | Piece | √ | √
| Beneview T5 CP monitor for BP cuff sizes 6-11 cm, , 10- 13 cm, 10-15 cm, 14-21.5 cm | Set | √ | √
| Beneview T5 CP monitor for Spo2 probe: boot type: Paediatric size | Piece | √ | √
| Beneview t5/8 (Mindray) patient monitor reusable NIBP cuff (arm) small infant circumference (cm) 7 to 13, bladder width (cm) 5.8 | Piece | √ | √
| Beneview t5/8 (Mindray) patient monitor reusable NIBP cuff (arm) infant circumference (cm) 10 to 19, bladder width (cm) 9.2 | Piece | √ | √
| Beneview 5/8 (Mindray) patient monitor reusable NIBP cuff (arm) paediatric circumference (cm) 18 to 26, bladder width (cm) 12.2 | Piece | √ | √
| Beneview t5/8 (Mindray) patient monitor wall mount bracket | Piece | √ | √
| Beneview t8 (Mindray) patient paediatric ICP transducer | Piece | √ | √
| CO2 sample line, watertrap, airway adapter for Mindray SynoVent E3 ventilator paediatric | Piece | √ | √
| Disposable breathing tube kid (Child) for Mindray SynoVent E3 ventilator with heating wire | Piece | √ | √
| Disposable filter (round) for Mindray SynoVent E3 ventilator | Piece | √ | √
| Infant reusable comfort NBP cuff 10 to 15 cm for Philip MP 60 monitor | Set | √ | √
| Oxygen sensor probe compatible for Datex Ohmeda | Piece | √ | √
| Pediatric circuit tubings with heated wire, silconised and autoclavable for Paediatric ventilator compatible with puritan bennett 840. | Piece | √ | √
| Paediatric reusable inspiratory bacteria filter (Re/Flex)compatible with Puritan bennett 840 | Piece | √ | √
| Paediatric reusable comfort NBP cuff 14 to 21.5 cm for Philip MP 60 monitor | Piece | √ | √
| Phoenix nCPAP 200 breathing circuit (paediatric) starter kit ( Phoenix nCPAP 200 pressure generator) | Piece | √ | √
| Reusable silicon Breathing tube kit(Child) (115-00 4502-00) for Mindray SynoVent E3 ventilator | Piece | √ | √
| Reusable silicon Breathing tube kit(Infant) (115-00 4503-00) for Mindray SynoVent E3 ventilator | Piece | √ | √
| Renovar CP monitor SpO2 paediatric probe (reusable | Piece | √ | √
| Schillers oxywave Spo2 monitor compatible probe -Paediatric size | Piece | √ | √
| Ultrasonic atom sanitizer 303 spare tubings with mask | Piece | √ | √
| Beneview t5/8 (Mindray CP monitor) reusable SpO2 probe, infant wrap type | piece | √ | √
| Philips MP60 CP montor, reusable SpO2 probe, wrap type. | Piece | √ | √
| Philips MP60 CP montor, reusable SpO2 probe, Boot type. | Piece | √ | √
| Pediatric ophalmoscope | Piece | √ | √
| UPS for ventilators | Nos | √ | √
Accessories for AIRVO2
| Filter 2 pack for AIRVO2 | set | √ | √
| Optiflow junio neonatal circuit | set | √ | √
| Optiflow junior nasal canula (OPT 316) | set | √ | √
| Optiflow junior nasal canula (OPT 318) | set | √ | √
| Optiflow junior wiggle pads (OPT 012) | pieces | √ | √
| Beneview t5/8 (Mindray CP monitor) reusable SpO2 probe, ped clip type | piece | √ | √
| ECG cable Philips MP-60 CP monitor, reusable ECG cable (M1971A) |  | √ | √
| Beneview t5/8 (mindray CP monitor) reusable ECG cable (EL6501-B) |  | √ | √
| Mindray ventilator face mask for CPAP |  | √ | √
Fisher and Paykel Health Care(pediatric size).
General consumable for Pediatric/Neonate
| Atom CPAP Nasal Canula,SS | Piece | √ | √
| Atom CPAP Nasal Canula,S | Piece | √ | √
| Atom CPAP Nasal Canula,M | Piece | √ | √
| Atom CPAP Nasal Canula, L | Piece | √ | √
| Bubble CPAP system starter kits : CPAP Mask, harness, gas supply line two flow meters (BC161-10), nasal prong(different sizes), infant bonnet(different sizes) and nasal tubing. | Set | √ | √
| Central peripherial line (PICC), neonate | Set | √ | √
| Central peripherial line (PICC), pediatric 22G | Set | √ | √
| Coban elastic self-adherent wrap tape | Piece | √ | √
| Disposable Paediatric ECG chest electrode(FOAM): Application : Paediatric; Shape: Teardrop ,Backing : Foam Gel : Solid Gel or Wet Gel; Sensor : Low impedance Ag/Agcl sensor: Size: 27mm | Piece | √ | √
| Disposable ECG leads for neonate compatible for Philips CP monitor MP60/MX600 latex free product code:13953D | Piece | √ | √
| EZ intraosseous needle size 15 mm needle | Set | √ | √
| EZ intraosseous needle size 25 mm needle | Set | √ | √
| EZ intraosseous needle stablizer | Packet | √ | √
| EZ intraosseous Connector | Packet | √ | √
| Extra-Spike® Plus: • Withdrawal cannula for multiple withdrawal from stock bottles and for supplementary injections, safe and handy; • Slim, short spike for easy piercing of stronger stoppers; • Easy opening and closing with one hand by grip plate centrally arranged to spike with integrated, tight snap with colour code; • Easy withdrawal and supplementary injection, reliable protection against microbial contamination by integrated airvent filter; • Luer lock-connector. | Piece | √ | √
| ECG paper compatible with Philips CP Monitor MP60 & Mx600 | Piece | √ | √
| Filter paper for ventilator humidifier compatible with Mindray | Piece | √ | √
| Filter paper for ventilator humidifier compatible with Drager | Piece | √ | √
| Heparinized insulin 1 ml syringe for ABG without needle | Piece | √ | √
| IV extension lines internal diameter 1mm, length 11 cm | Piece | √ | √
| IV extension lines internal diameter 1mm, length 30 cm | Piece | √ | √
| IV extension lines internal diameter 1 mm, length 50 cm | Piece | √ | √
| IV extension lines internal diameter 1mm, length 100 cm | Piece | √ | √
| IV Set (60 drops per ml) | Piece | √ | √ | √ | √
| Incentive spirometer (for Lung exercise) pediatric and adult use 3 chamber of pf 600, 900 and 1200 cc/sec inhaling volume, three coloured balls for easier indentification of inhaled capacity | Set | √ | √
| Intravenous line filter (bacterial) 0.2 micro supor, filters bacteria and particles, luer lock connector, latex free port for supplement medication, effective air ventilation regardless of position | Set | √ | √
| Intraosseous infusion needle, 14 gauge | Piece | √ | √ | √ | √
| Intraosseous infusion needle, 16 gauge | Piece | √ | √ | √ | √
| Intraosseous infusion needle, 18 gauge | Piece | √ | √ | √ | √
| Microdrip Set with burrete | Piece | √ | √ | √ | √
| Meconium aspirator (Pengiun Type) | Piece | √ | √ | √ | √
| Octopus is a polyurethane multi-lumen (two) extension tube for I.V. catheters, transparent and with low priming volume. Priming volume (0.3 to 0.5 ml) with non-return valve. Length 10 cm, internal diameter 1 mm and external diameter 2 mm. | Packet | √ | √
| Paediatric double-lumen central venous catheter: Radiopaque central venous polyurethane catheter for injecting or infusing different solutions and/or medications simultaneously; Tray presentation containing; 1 double-lumen catheter (3Fr) with centimetre markings from 4 cm from distal tip; 1 short I.V. cannula 24G - Length 19 mm; 1 introducer needle 22G - Length 40 mm; 1 Nitinol straight guidewire (particularly kink-resistant) with guide advancer - Length 50 cm; 2 dilators 3Fr (Length 3 cm and 5 cm); 1 secondary fixation wing; 2 injection caps; 1 scalpel; 1 5 ml syringe | Set | √ | √
| Paediatric double-lumen central venous catheter: Radiopaque central venous polyurethane catheter for injecting or infusing different solutions and/or medications simultaneously; centimetre markings from 5 to 12 cm; Tray presentation containing: 1 double-lumen catheter; 1 introducer needle (Ext. Ø 0.8 mm – 21G -Length 40 mm); 1 short I.V. cannula 22G -Length 25 mm; 1 “J” or straight guidewire (Length 40 cm -Ø 0.46 mm) 4,5Fr; 2 dilators (length: 38 mm and 62 mm); 1 secondary fixation wing; 2 injection caps; 1 scalpel; 1 5 ml syringe. | Set | √ | √
| Spare nebuliser chamber with tubings and pediatric mask | Set | √ | √ | √ | √
| Self adhesive urine sample bag - Pediatric | Piece | √ | √ | √
| Sterile IV water bottle (one litre) | Piece/Packet | √ | √
| Sterile Skin cleansing alcohol preparation (webcol) saturated with 70% isopropyal alcohol | Packet | √ | √
| Steri-strip antimicrobial skin closures | Piece | √ | √
| Single channel disposable sensor kit (20) part no M1567A for MP60 Philips CP monitor | Set | √ | √
| Treacheal suction Set (specimen collection for sampling of secretions from tracheobronchial tree, transparent container) | Piece | √ | √
| Umbilical catheters 3.5 F | Piece | √ | √ | √
| Umbilical catheters 5 F | Piece | √ | √ | √
| Vaccu suck. Crown tip-Yankaur suction set. Tube length 2.5 meter. | Packet | √ | √
Consumables of ABG Machine
| Cartridges, Test Cards (OPTI CCA TS) | 25 test/ Packet | √ | √
| Thermal Printing Paper (OPTI CCA TS) | Roll | √ | √
| Calcium Gas, Bottle (OPTI CCA TS) | Bottle | √ | √
| Mindray VS 800 BP cuff size 18 x 26 cms, arm |  | √ | √
| circumference bladder. | Set
| Mindray VS 800 BP cuff size 9.8 x 13.3 cms, arm |  | √ | √
| circumference bladder. | Set
| Welch Allyne, NIBP monitor, BP cuff, |  | √ | √
| size 12 x 16cms | Packet
| Welch Allyn, NIBP monitor, BP cuff, size 9.8 x 13.3 cms. |  | √ | √
Additional for Neonatal consumables
| Coban 3M 1inch | box | √
| Tegaderm 1 1/2 inch | box | √
| Extrathin CGF dressing 10cm | box | √
| Sterilised prepacked umbilical tape for umbilical catheterisation | box | √
| X-ray gonad shields for preterm and neonate | box | √
| Phototherapy eye shields for preterm and neonate | box | √
| Sterilised endotracheal tube adaptor with lure lock side port for use with surfactant for ET size 2.0-4.0 | box | √
| Double lumen umbilical catheter 4Fr and 5fr | box | √
| neonatal extension tubes with filters liquid infusion lines | box | √
| Neonatal liquid filters | box | √
| double lumen octopus with auto flush and non-return valves | box | √
| Triple lumen octopus with auto flush and non-return valves | box | √
| Anti UV lure lock infusion lines (Lectroguard) | box | √
| Cap/bonnet for neonate | box | √
| Cap/bonnet for preterm | box | √
Pharmacy Department
Equipments
| Pharmacy Refrigerator: Temperature range display: +2° C to +8° C LED display should be easy to read for precise control. Capacity, Shelves & Doors: 350-500 ltrs, 5 shelves and 2-4 doors with adjustable wire shelves and sliding wire baskets, doors should be transparent. Dimension and design: As per capacity and design with chip and rust resistant paint finish and should be easy to clean and 100% stainless steel. Alarm & safety device: Both high and low temperature alarm, refrigerator or heating system should shut off if temperature drops below 2° C or rise above 20° C with keyed door lock. Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. Suitable voltage stabilizer. . Certification: CE certification with minimum of 1 year warranty. |  | √ | √
| Weighing balance for extemporaneous compounding: Capacity: 120-300 grams, Sensitivity: 100 mg, Linearity: ± 0.1 g, Operation temperature range: -10° C to +40° C, Display: LCD, Calibration: Internal, Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. stainless steel weighing pan. | Set | √ | √
| Digital weighing scale – 500 mg to 5 kg for compounding | Piece | √ | √
| Pill counting tray: Size of the tray: 20 cm x 14.5 cm x 3.8 cm, material of the set: Plastic (ABS + PS), Material of spatula: Wooden handle + stainless steel. | Set | √ | √
| Spatula, ointment, metal 12 inch SS | Piece | √ | √ | √ | √
| Ceramic Mortar and pestle - size 10 inch: Material: Ceramic, Colour: white | Set | √ | √ | √
| Ceramic Mortar & pestle - size 4 inch: Material Ceramic, colour: white | Set | √ | √ | √
| Tablet counting machine: Portable table top automatic tablet counting machine; Counts coated and uncoated tablets, hard and soft gelatin capsules and dragees; Digitally controlled count adjustment; Accurate counting and dispensing (100%) with excellent repeatability; Tool free removal and change of product contact parts. Weight of the machine: 3-10 kgs. Dimensions of the machine: 30-40 cm height x 30-40 cm Length x 15-20 cm width. Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. Throughput: 15-18 tablets/seconds. Product size: Should enable to count product of dimension 3-20 mm diameter and 2.5-20 mm thickness. Build standard: Should comply with cGMP or equivalent standards. CE marked in accordance with machinery. Operation: Quiet operation (should not have a noise level of more than 80 dBA); Simple operations with fast product change over. Cleaning procedure: The machine should offer easy access for cleaning. Count capacity: Should enable to count up to 9,999. Product contact parts: All product contact parts should be made of SS 316 or better. Partial pill indicator: Alerts when an object <25% the size of the tablet/capsule is detected. Spares and accessories: Should include all the required parts for product change over, spares & consumables and accessories such as air compressor, hoses, etc. | Set | √
| Wash Room Trolley SS | Piece | √ | √ | √
| Pharmacy shelve Dimension-2080x1000x330mm 5shelves with adjustable feet Unit should be assembled with the minimum requirement of tools Shelving config are easily adjusted to suit stocking requirements Ticket railing to allow clear product identification label | set | √ | √ | √ | √
| Ointment Blender | Piece | √ | √ | √
| Laminor work flow bench | Set | √
| Label printer | Pcs | √
| Hot air oven | Set | √
| Autoclave | Set | √
| Digital Hygro-thermometer: displays temperature and relative humidity; wall mountable; Temperature range: -10 to 60°C; Humidity range: 10% to 99% | Pcs | √ | √ | √ | √
Consumables
| Bottle for dispensing extemporaneous preparations: Amber colour PET bottles with child resistant screw cap and 100 ml capacity. | Piece | √ | √
| Personal protective equipment kit: Gown with cuffs and back closure (should be water impermeable fabric). Pair of gloves (specially designed for protection). Pair of goggles. Pair of shoe covers (disposable). Mask (disposable). Absorbable towel | Set | √
| Cytotoxic Drugs spill kit: Gown with cuffs and back closure (should be water impermeable fabric). Pair of gloves (specially designed for protection). Pair of goggles. Pair of shoe covers (disposable). Mask (disposable). Mask (N-95), Absorbable towel, disposable sponges, puncture-proof, leak proof container, container of desiccant powder or granules to absorb wet contents, disposable dust pan, plastic scraper. | Set | √
| label paper for Label Printer | Pcs | √
| Envelope, plastic dispeensing 10cm X 12cm | Pkt of 1000 | √ | √ | √ | √
Physiotherapy Department
Equipments
| Traction Machine for Cervical & Lumbar spine with Versatility, Tension Main Timer, Hold Timer, Patient control switch, Traction pilot light Safety Fuse with table and cervical and Lumbar harness, Thoracic Harness, Spreader bar, Pillow and triangular flexion stool. Power cord with round 3 pin top. Appropriate voltage stabilizer should be provided with the unit. a. Cervical & Lumbar Traction Machine Cables (chains). b. Cervical & Lumbar Traction Machine: Cervical. Environmental operating condition: 2Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. Suitable voltage stabilizer. RH >80%. | Set | √ | √ | √
| Electric muscle stimulator: Diagnostic cum therapeutic unit, suitable for taking S.D curve, producing different types of stimulating currents i.e Galvanic and faradic current . The stimulating current outputs are: Galvanic - continuous DC upto 135 volts for a maximum of 0-75mA load. Interupted Galvanic - Pulse durationof 0.01, 0.03, 0.1, 0.3, 1, 3, 10, 30, 100, 300m sec. Faradic - 80 pulse per second duration with 0.6 ms width. Accessories - Inactive and active plate electrodes, ball tip electrode with handle, 2 pair electrode wires. Environmental operating condition: Relative humidity >80%. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. Suitable voltage stabilizer (delete). | Set | √ | √ | √
| Hot Pack Machine: Stainless steel with on/off switch and thermostatic control. Preset normal operating temperature control, Preset normal operating temperature control. a. Hot Pack Contour cervical Pack with cover. b. Hot Pack Half size Pack with cover. Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. RH >80%. | Set | √ | √ | √
| Infrared Lamp with watt range from 150 - 500 watts,Automatic timer, thread type bulb, . Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. RH >80%. | Set | √ | √ | √
| Interferential Machine: Two oscillators of medium frequency are generated and cross- applied to the body by means of four electrode pads/self adhesive electrodes. Oscillator frequency fixed to 4000 Hz and second can be ranged to 4000 – 4150 Hz. a. Interferential Machine cable to the pads. b. Interferential Machine pad-large with cover/self adhesive electrodes. c. Inferential machine pad -medium with cover/self adhesive electrodes. d. Inferential machine pad small with cover/self adhesive electrodes. e. interferential machine switch cable. Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. >80%. | Set | √ | √ | √
| Short Wave Diathermy SWD: High power short wave Diathermy with patient’s switch & internal cooling system along with additional voltage stabiliser 450 watts. Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. RH >80%. | Set | √ | √ | √
| Therapeutic LASER: Mid laser used by physiotherapist. Helium-neon laser producing red light at 630nm combined with infrared laser at 904nm, or other physiotherapy laser. Environmental operating condition: RH >80%. Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √
| Therapeutic Ultra Sound Machine: Ultra sonic Deep Therapy Unit with complete set including Transducer.1-3 MHz. Environmental operating condition: RH >80%. Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √ | √
| Pocket Transcutaneous Electrical Nerve Stimulator (TENS): Small home stimulation pack (dual channel) with wire attachments, electrodes, 9 V batteries. a. Transcutaneous Electrical Nerve Stimulator (TENS): 9V battery. b. Transcutaneous Electrical Nerve Stimulator. 1. Dual channel: individually isolated circuits. 2. Amplitude: 0 mA - 80 mA; indication only; actual mA will tend to be less than indicated due to electrode impedance. 3. Type: Constant Current. 4. Waveform: Asymmetrical, rectangular bi-phasic with zero DC current. 5. Selectable pulse width: 50 µS - 300 µS [2% accuracy]. 6. Pulse Rate selection: in the continuous mode 2 Hz – 200 Hz [2% accuracy]. 7. Mode: Continuous, Burst or Modulated. 8. Burst mode: Bursts of 9 pulses [200 µS] at 150 Hz, repeating twice every second. 9. Modulation mode: 6-second cycle of concurrent width modulation and pulse repetition rate modulation. Width starting at 200 µS and decreasing exponentially to 100 µS in three seconds and then returning back to 200 µS in the next three seconds. Rate starting at 100 Hz, decreasing exponentially to 65 Hz and then returning to 100 Hz. 10. Time duration of the treatment selectable: 1 minute to 12 hours. 11. Low Battery Indicator: If the battery goes below 6.9 volts +/- 0.2 volts the battery symbol will flash on/off once every second. 12. Environmental conditions for storage & transport: -10 to +40 degrees Centigrade>80% Humidity. 13. Self adhesive electrodes.Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √ | √
| Adaptive sitting chair: A sturdy and comfortable seating system, the High Low Chair (for children up to 45" in height) is a combination high chair and mobile floor sitter for children. Comfortable, Secure Positioning. Easily converts to a mobile floor sitter by detaching the upper frame from the pedestal base. Adjusts in height with the use of a single lever that activates a pneumatic cylinder. The mobile floor sitter frame is constructed of 1" powder coated tubular steel with locking casters. The cushioned seat is covered with a water proof, easy to clean polymer coated fabric. The durable melamine tray remains securely in place with a quick release locking lever. Include H-strap harness, washable cotton quilted trunk support, removable abductor and tray. An accessory head support and lateral supports are available. The high low chair, seat depth: 7" - 10", seat back height: 22", seat width: 12", High chair seat height: 19", Floor sitter seat height: 24", Tilt in space angle: 15°, 20°, 25°, Weight of seat and frame: 46 lbs, Weight capacity of Hi Low chair:75 lbs, Wheel diameter: 2", Foot rest dimensions: 6" x 12", Tray dimensions: 12" x 18". Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √
| Assorted square & round pegboard: wooden board with 20 multi colour wooden pegs (10 round & 10 square) in four rows. Peg size - 76 x 16 mm | Set | √ | √
| Corner Chairs: Provide support for symmetrical posture and mid-line control of head and trunk and improve breathing for children who have low trunk tone. Prevents scapular and shoulder retraction, a pattern frequently seen in children who cannot stabilize against gravity. Age/Size: 8 months - 4 years, 4 years - 8 years. Height of chair with head supports, adjusts from 51 cm - 69 cm, adjust from 74 cm - 104 cm. Removable tray: 38 cm x 43 cm, 46 cm x 51 cm. Seat depth with insert: 20 cm, 26 cm. Seat depth without insert: 30 cm, 36 cm. Removable abductor: 8 cm wide, 10 cm wide. Chair with 4" swivel, locking casters. | Set | √ | √
| Delux revolving hole punch, length 22 cm, throat depth 3.8 cm with replacement tubes of 2.4 mm tube, 2.8 mm, 3.2 mm, 4 mm, 4.4 mm, & 4.8 mm | Piece | √ | √
| Dumbell of Stainless steel with weight of 1 kg, 1.5 kg, 2 kg, 2.5 kg, 3 kg, 3.5 kg, 4 kg, 4.5 kg, 5 kg (one pair each) in a set with stand frame. | Set | √ | √ | √
| Treadmill: 0.1mph to 11 mph adjustable speed (0.1 mph increments). Incline range from 0 -15 % in 1% increments. Support bars on both sides. Surface area - 51 x 140 cm, safely stop button, hand rails, weight capacity 400Ibs. Environmental operating condition: RH >80%. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. | Set | √ | √ | √
| EMG bio-feedback, pressure feedback, electrotherapy and electro diagnostics. The myofeedback unit is provided with separate part for myofeedback, pressure feedback and combination with electrotherapy. Several surface and cavity electrodes can be connected. This makes the unit extremely suitable for all kinds of exercise and relaxation therapies including Pelvic floor re-education. Environmental operating condition: RH >80%. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. | Set | √
| Heat gun: A free standing, light weight heat gun for use with all types of thermoplast with 2 hot air settings 300°C - 500°C. Weight 550 gm. Environmental operating condition: RH >80%, Voltage: AC 210-240V, 50Hz . | Set | √ | √
| Magnetic Field Therapy Unit, MG-Wave is a portable unit for therapy with low-frequency pulsating electromagnetic field. Specification: Field Pilot Signal: rectangular. Pulse Duration: Upto 5ms, Frequency: 1 + 750Hz, Max B Induction: 150 gauss (peak) (± 15%) on Piece coil, Timer: 0+240min, end-of-treatment acoustic signal and automatic interruption of operation. Power Supply: 100-200 Vac, 50-60Hz wall-mounted power supply. Input: 75VA, Safety Class: II- BF type (CEI EN 60601-1). Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. RH >80%. | Set | √
| Treatment Couches 6-Section (Manumed Optimal) : Hydraulic high-low/Electric high-low, head section, flexion position, separate leg section, table top size 280 x 67cm, height adjustment 48-98cm. | Set | √ | √
| Treatment Couches 3 - Section (Manumed Optimal): Hydraulic high-low/Electric high-low, table top size 280 x 67cm, height adjustment 48-98cm. | Set | √ | √
| Mirror (with wheels) for therapeutic purpose - size (27 inch W x72 inch H) | Piece | √ | √ | √
| Parallel Bars adjustable height and width by crank arm 3 m long. with base included - Paediatric | Set | √ | √
| Parallel Bars adjustable height and width by crank arm 4.5 m long. with base included - Adult | Set | √ | √ | √
| Paediatric positioner: Half roll with two lateral supports. Wedge, Log roll, Tray with mirror for creative positioning, Carrying bag, Modules allow for a variety of creative positioning, Three positioning belts help to secure child, Soft covering is easy to clean and is nontoxic. Size: 34"x 19" (86 cm x 48 cm) base | Set | √ | √
| Recumbent Cycle Trainer: Adjustable seat height and distance with back rest and adjustable resistance. Electric or battery operated display to show RPM, time, resistance level, distance, speed. | Set | √ | √ | √
| Shape sorting cube with 10 multi colour wood blocks: size of box - 5.5 inch L x 5.5 inch W x 5.5 inch D. | Set | √ | √
| Compact Shoulder Wheel | Piece | √ | √ | √
| Stacking cones with sequencing & colour identity, size - 4.5 inch H x 2.5 inch diameter | Set | √ | √ | √
| Standing frame - Adult | Set | √ | √ | √
| Standing frame - Paediatric | Set | √ | √
| Tilt Table: Adult size, manual/electric, dimension 78" lengthX 28" widthx 33" height, weight capacity 230 lbs with angles from 0-90 degrees | Set | √ | √ | √
| Wire cutter, heavy duty, length 22 cm | Piece | √ | √
| Crank adjustable Hi-Lo Mat platform. 183cm x 244 cm | Set | √ | √ | √
| Therapeutic Exercise ball , red: Diameter: 20-100cm Should come with inflation valve plug remover. | Piece | √ | √
| Therapeutic Exercise ball, blue: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
| Therapeutic Exercise ball, green: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
| Therapeutic Exercise ball, orange: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
| Therapeutic Exercise ball, yellow: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
| Peanut Therapeutic Exercise ball, Diameter:40-80 cm | Piece | √ | √
| Two-way Air pump: Manually operated pump inflates balls, toys air mattresses etc on up and down strokes cutting effort in half. | Piece | √ | √
| Trampoline: With an easy grip handle, Suitable for both indoor and outdoor use, with tough weatherproof mat and tubular steel frame with a durable coated finish, padded cover around the mat. The handle simply screws off and the legs fold down for ease of storage. Weight: 8 Kg, from handle to floor: 32.25", dimensions: 33.75" x 33.75", maximum weight: supports up to 68Kgs. | Set | √ | √
| Cushion: wedge shaped with firm, wipe down cover, two wrap straps, one abductor module for pediatric, size: Height 4-12 inch, Length:22-28 inch, width:20-24 inch. | Set | √ | √
| Cushion: Bolster shaped with firm, wipe down cover, cylindrical shape, diameter: 8-18 inch, Length: 36-48 inch | Set | √ | √
Consumables
| Self adhesive electrodes for TENS and Interferential Machine | Packet | √ | √ | √
| Cold pack 23 inch neck contour with cover(It should be compatible for Cold pack machine) | Piece | √ | √ | √
| Cold pack Oversize with cover (It should be compatible for Cold pack machine). | Piece | √ | √ | √
| Cold pack Standard size with cover (It should be compatible for Cold pack machine). | Piece | √ | √ | √
| Electric muscle stimulator Electrodes - medium (It should be compatible for Electric Muscle Stimulator machine) | Piece | √ | √ | √
| Electric muscle stimulator Electrodes - small (It should be compatible for Electric Muscle Stimulator machine) | Piece | √ | √ | √
| Finger Hooks, 6.4 mm, steel used in splinting (Packet of 100) | Pkt | √ | √
| Hot Pack Contour cervical Pack with cover (It should be compatible for Hot Pack Machine) | Piece | √ | √ | √
| Hot Pack Oversize Pack with cover (It should be compatible for Hot Pack Machine) | Piece | √ | √ | √
| Hot Pack Standard size Pack with cover(It should be compatible for Hot Pack Machine) | Piece | √ | √ | √
| Interferential Machine Electrodes - medium (It should be compatible for Interferential machine). | Piece | √ | √ | √
| Interferential Machine Electrodes - small (It should be compatible for Interferential machine). | Piece | √ | √ | √
| Interferential Machine Electrodes cover – medium (It should be compatible for Interferential machine). | Piece | √ | √ | √
| Interferential Machine Electrodes cover- small (It should be compatible for Interferential machine). | Piece | √ | √ | √
| Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Yellow: Light) | Piece | √ | √ | √
| Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Red: Medium) | Piece | √ | √ | √
| Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Green: Heavy) | Piece | √ | √ | √
| Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Blue: Extra Heavy) | Piece | √ | √ | √
| Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Black:2X Heavy) | Piece | √ | √ | √
| Air cushion: Providing additional shock and vibration absorption provides a more pain relieving ability and a more stable seating environment. Inflates by mouth or with a small portable hand air pump that comes with the cushion. Standard size: 22" x 22"/56 cm x 56cm, Weight: 1 Kg, Cell Height: 2"/5cm fully inflated (due to the required immersion of the user in the cushion, it will raise the user only (1/2")/(1 cm) above the seat), Cushion material: Neoprene rubber, Cover material: Polyester blend with a non-skid polyester and vinyl base, Water Resistance: Cushion and cover are water resistant. | Set | √ | √
| Gutter Splint/Gaiter: Large | Piece | √ | √ | √
| Gutter Splint/Gaiter: Medium | Piece | √ | √ | √
| Gutter Splint/Gaiter: Small | Piece | √ | √ | √
| Gutter Splint/Gaiter: Pediatric | Piece | √ | √ | √
Consumables - (Mobility Aids)
| Adjustable Axillary Crutch (Aluminum) - Large | Pair | √ | √ | √ | √
| Adjustable Axillary Crutch (Aluminum) - Medium | Pair | √ | √ | √ | √
| Adjustable Axillary Crutch (Aluminum) - Small | Pair | √ | √ | √ | √
| Adjustable Elbow Crutch (Alumunium) - adult | Pair | √ | √ | √ | √
| Adjustable Elbow Crutch (Alumunium) - Small | Pair | √ | √ | √ | √
| Elbow Crutch Rubber Tips | Pair | √ | √ | √
| Walker Rubber Tips | Pair | √ | √ | √
| Adjustable Walker with wheels (Alumunium) - adult | Piece | √ | √ | √
| Adjustable Walker with wheels (Alumunium)- Small | Piece | √ | √ | √
| Adjustable Walker without wheels (Alumunium) - adult | Piece | √ | √ | √
| Adjustable Walker without wheels (Alumunium) - Small | Piece | √ | √ | √
| Adjustable Platform Crutch (Aluminum) - Large | Pair | √ | √
| Adjustable Platform Crutch (Aluminum) - Medium | Pair | √ | √
| Adjustable Platform Crutch (Aluminum) - Small | Pair | √ | √
| Adjustable Platform Walker with wheels (Alumunium) - Large | Piece | √ | √
| Adjustable Platform Walker with wheels (Alumunium) - Medium | Piece | √ | √
| Adjustable Platform Walker with wheels (Alumunium) - Small | Piece | √ | √
| Adjustable Platform Walker without wheels (Alumunium) - Large | Piece | √ | √
| Adjustable Platform Walker without wheels (Alumunium) - Medium | Piece | √ | √
| Adjustable Platform Walker without wheels (Alumunium) - Small | Piece | √ | √
| Adjustable Posterior Walker with wheels (Alumunium) - Adult/ Two fixed direction 5 inch front wheels & two fixed direction 3 inch. One way ratchet rear wheels. | Piece | √ | √
| Adjustable Posterior Walker with wheels (Alumunium) - Pediatric/ Two fixed direction 5 inch front wheels & two fixed direction 3 inch. One way ratchet rear wheels. | Piece | √ | √
| Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 12.5" (inches); Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 14"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set
| Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 15.5"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 17"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 18.5"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 12 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims on the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims on the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 20 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 12 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 20 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 12 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims; Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 12 inches (for children) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches (small adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches (medium adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches (standard adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 20 inches (large adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 8” Seat 200 mm wide x 175-250 mm adjustable depth Seat and frame configuration: If the chair is modular the seat system (combined seat and backrest with optional accessories such as headrest) will be separate from the base frame. The chair or the base frame to be rigid or folding. Infinite adjustable seat system tilt between 0 and minimum 30 degrees by means of a quick release mechanical or hydraulic option. The design should allow for COG adjustment either as a specific wheel/seat adjustment or by means of a mechanism which will always result in COG being safely within base of support. The design of the COG adjustment should allow for appropriate access to the rear wheel. Seat module: The seat to back angle to be adjustable between 90° and 140° with a minimum of 5 settings.The design should always allow the foot rests to be in line with the front of the seat after adjustment of seat depth. Seat base must provide rigid support. The seat will have a length adjustable pelvic strap at 60°- 90° angle over hips. Width of pelvic strap for 12” and smaller = 30mm. Width of pelvic strap for 14” and bigger = 50mm. 12” short and smaller sizes should allow for age appropriate hip abduction. Back post / Push handle requirements: Top of Push handles must be a minimum 93.5cm from ground to top of push handles with 10cm height adjustability. Push handle may in no way restrict self-propelling, affect posture or set-up of back support system. Armrests: Armrests supplied as standard for chairs larger than 12” wide but can be removed as an option. Must be height adjustable with a minimum range of 100mm with 15mm increments (or infinitely adjustable) with the lowest height being 150mm from the top of the seat. Must be adjustable forward/back and must allow for support directly below the elbows. Footrest: Height adjustable hanger with minimum 150 mm range (100mm on 12” chair and smaller) to allow for variations in user profile and seat length adjustments. Must flip-up / fold up or must be detachable (not required on chairs smaller than 12”). Adequate rigidity and must have a non-slip surface. Tray-table: Lap tray to be supplied as standard for chairs smaller than 12” wide. Lap tray Optional device for chairs over 12”. Fitting for tray table to be provided on frame as standard. See specifications below under options. Head support: Standard accessory. Adjustable forward/back and up/down with a range of at least 100mm and 200mm respectively. Front castors and forks: Front castor wheel diameter range: 100 mm and 200 mm (+2mm). Front castor wheel width range: 25 mm and 80 mm. Stem to be fastened to the fork with thread and nut system to allow for re-fastening and adjusting. Rear wheels and axles: Size 12” standard chair and larger: minimum 24” x 1 3/8 diameter with one part semi-solid tyre. Size 12” short chair and smaller: minimum 20” x 1 3/8 diameter with one part semi-solid tyre or foam insert. Optional pneumatic tyre. To be fitted with quick-release axle. Anti-tip wheels: Must be removable. Should allow for adjustability to match the level of independence and balance of the user. Range of adjustability to allow for stability in static position and to adjust for the required degree of stability in propulsion and if applicable, when front castors are raised for obstacle clearance. Weight: Not to exceed 18 kg (complete, but without back system, tray table or head extension). 12” and under not to exceed 14 kg (complete, but without back system, tray table, or head support/extension). Repair / tool kit: All tools required to adjust frame and back rest systems, as well as foot rest system will be supplied with the chair. Back system: Tension adjustable backrest with rigid, height adjustable full length laterals to fit both basic and adjustable folding frame wheelchairs. Standard back height 42cm for chair sizes above 12" and 30cm for 12" & smaller chairs. Tall back height 50cm for chair sizes above 12" and 42cm for 12" & smaller chairs. Lateral depth: 20cm for 18" & 20" chairs, 15cm for 14" & 16" chairs, and 10cm for 12" chairs. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 10” Seat 250 mm wide x 250-350 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 12” short Seat 300 mm wide x 250-320 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 12” standard Seat 300 mm wide x 300-400 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 14” short Seat 350 mm wide x 250-350 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 14” standard Seat 350 mm wide x 350-450 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 16” standard Seat 400 mm wide x 350-450 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
| Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 18” standard Seat 450 mm wide x 350-450 mm adjustable depth. Specifications described in No. 25 | Set | √ | √
Consumables - (Supportive)
| Strapping and Taping Tape: Breathable porous athletic tape protects joints from sprains and injuries. Provides support to help speed recovery from injuries, high strength, versatile tape for a wide range of uses. Easy to tear, no scissors necessary. Size: 3.8 cm x 13.7 - 32 per case | Roll | √ | √
| Strapping and Taping Tape: Breathable porous athletic tape protects joints from sprains and injuries. Provides support to help speed recovery from injuries, high strength, versatile tape for a wide range of uses. Easy to tear, no scissors necessary. Size: 5 cm x 13.7 - 24 per case | Roll | √ | √
| Ankle Support/brace: Extra strong highly porous orthopedic webbing, containing heat resistant rubber with high modulus of elasticity. Hook loop closure - Large | Piece | √ | √ | √
| Ankle Support/brace: Extra strong highly porous orthopedic webbing, containing heat resistant rubber with high modulus of elasticity. Hook loop closure - Medium | Piece | √ | √ | √
| Ankle Support/brace: Extra strong highly porous orthopedic webbing, containing heat resistant rubber with high modulus of elasticity. Hook loop closure - Small | Piece | √ | √ | √
| Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - below 26 cm (child size) | Piece | √ | √ | √
| Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - 26-30 cm (small) | Piece | √ | √ | √
| Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - 30-34 cm (medium) | Piece | √ | √ | √
| Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - 34-38 cm (Large) | Piece | √ | √ | √
| Back supports: 80 cm | Piece | √ | √ | √
| Back supports: 85 cm | Piece | √ | √ | √
| Back supports: 90 cm | Piece | √ | √ | √
| Back supports: 95 cm | Piece | √ | √ | √
| Back supports: 100 cm | Piece | √ | √ | √
| Back supports: 105 cm | Piece | √ | √ | √
| Back supports: 110 cm | Piece | √ | √ | √
| Soft Cervical collar, high density polyurethane foam cover with cotton stockknitte, plastic bushes to provide enough ventilation and Velcro closures for easy application, removal and adjustment, polythylene sheet reinforcement for better immobilization: size neck circumference in cm 38-42 (large) | Piece | √ | √ | √
| Soft Cervical collar, high density polyurethane foam cover with cotton stockknitte, plastic bushes to provide enough ventilation and Velcro closures for easy application, removal and adjustment, polythylene sheet reinforcement for better immobilization: size neck circumference in cm 34-38 (medium) | Piece | √ | √ | √
| Soft Cervical collar, high density polyurethane foam cover with cotton stockknitte, plastic bushes to provide enough ventilation and Velcro closures for easy application, removal and adjustment, polythylene sheet reinforcement for better immobilization: size neck circumference in cm 30-34 (small) | Piece | √ | √ | √
| Crepe Bandage - 5 cm x 450 cm (4.5 m stretch length) | Roll | √ | √ | √
| Crepe Bandage - 10 cm x 450 cm (4.5 m stretch length) | Roll | √ | √ | √
| Crepe Bandage - 15 cm x 450 cm (4.5 m stretch length) | Roll | √ | √ | √
| Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (Size: small - 12" - 14") | Piece | √ | √ | √
| Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (Size: Medium - 14" – 15" | Piece | √ | √ | √
| Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (Size: Large - 15"-17") | Piece | √ | √ | √
| Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (SizeX-Large: 17" - 20") | Piece | √ | √ | √
| Silicon Polymer Heel Pad/cup/Shoe Lift : shoe size 9 and 10 (Large) | Pair | √ | √
| Silicon Polymer Heel Pad/cup/Shoe Lift : shoe size 7 and 8 (Medium) | Pair | √ | √
| Silicon Polymer Heel Pad/cup/Shoe Lift : shoe size 5 and 6 (Small) | Pair | √ | √
| Top Layer Foam: High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m3. Size range: 46 cm wide; 43 cm long; 5cm high. 41 cm wide; 43 cm long; 5cm high. 38 cm wide; 43 cm long; 5cm high. 36 cm wide; 43 cm long; 5cm high. 30 cm wide; 43 cm long; 5cm high. Variance allowed on overall measurement for cushions: +1cm. The entire size range will remain one series. | Series | √ | √ | √
| Top Layer Foam: High-density foam 40/24; Hardness factor 21.6-25.2; Density range 36.1-40kg/m3. Size range: 61 cm wide; 43 cm long; 5cm high.  56 cm wide; 43 cm long; 5cm high.  51 cm wide; 43 cm long; 5cm high.. Variance allowed on overall measurement for cushions: +1cm. The entire size range will remain one series. | Series | √ | √ | √
| Base layer Foam: High density compressed (reconstituted flexible PU) foam 1300; Hardness factor –52-63; Density 130+10% kg/m3. Size range:  61 cm wide; 43 cm long; 5cm high  56 cm wide; 43 cm long; 5cm high  51 cm wide; 43 cm long; 5cm high  46 cm wide; 43 cm long; 5cm high  41 cm wide; 43 cm long; 5cm high  38 cm wide; 43 cm long; 5cm high  36 cm wide; 43 cm long; 5cm high  30 cm wide; 43 cm long; 5cm high Variance allowed on overall measurement for cushions: +1cm The entire size range will remain one series. | Series | √ | √ | √
| Incentive Spirometer: 500 - 4000 ml capacity Latex + DNR free. | Piece | √ | √
| Moleskin Padding with self adhesive backing, roll of 0.8 mm x 300 mm x 4.6 m | Roll | √ | √
| Splinting material (Thermoplast/ Aquaplast/ EZEform): Perforated, thickness 2.4 mm, size - 61 x 91 cm | Sheet | √ | √
| Splinting material (Thermoplast/ Aquaplast/ EZEform): Perforated, thickness 3.2 mm, size - 61 x 91 cm | Sheet | √ | √
| Splinting material (Thermoplast/ Aquaplast/ EZEform): Solid, thickness 2.4 mm, size - 61 x 91 cm | Sheet | √ | √
| Splinting material (Thermoplast/ Aquaplast/ EZEform): Solid, thickness 3.2 mm, size - 61 x 91 cm | Sheet | √
| Tennis Elbow Strap - Large (30-35 cm) | Piece | √
| Tennis Elbow Strap - Medium (25-30 cm) | Piece | √
| Tennis Elbow Strap - Small (18-25 cm) | Piece | √
| Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Extra Large (55-61 cm) | Pair | √ | √
| Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Large (49-55 cm) | Pair | √ | √
| Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Medium (43-49 cm) | Pair | √ | √
| Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Small (37-43 cm) | Pair | √ | √
| Velcro loop, size - 25 mm, reel 25 meter | Roll | √ | √
| Velcro self adhesive hook, size - 25 mm, reel 25 meter | Roll | √ | √
| Wrist brace - strong porous, orthopaedic webbing. Double hook loop locking system enhances the immobilization and provides a firm grip to the wrist. Size 20.0-23.8 cm (Large) | Piece | √ | √
| Wrist brace - strong porous, orthopaedic webbing. Double hook loop locking system enhances the immobilization and provides a firm grip to the wrist. Size 16.3-20.0 cm (Medium) | Piece | √ | √
| Wrist brace - strong porous, orthopaedic webbing. Double hook loop locking system enhances the immobilization and provides a firm grip to the wrist. Size 12.5-16.3 cm (Small) | Piece | √ | √
| Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.25 x 75mm; Packet of 100 needles. | Packet | √
| Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.30 x 50mm; Packet of 100 needles. | Packet | √
| Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.20 x 30mm; Packet of 100 needles. | Packet | √
| Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.25 x 25mm; Packet of 100 needles. | Packet | √
Radiology Department
X-Ray Unit
| X-Ray Machine 300 mA (specification) | set | √ | √ | √ | √
| X-Ray Machine 600mA (specification) | set | √ | √ | √ | √
| X-Ray Machine 60 mA, portable (Moltimobile 2.5 )(specification) | set | √ | √ | √ | √
| Film Auto Processor: Automatic films processors for processing of all types of X-Ray and Imaging Films It should have the following technical features: Should accept film size from 10x10 inches to 14x17 inches, It should process minimum of 60 films per hour both in rapid cycle and standard cycle, The processing time should be variable from 1.5min to 5 min, It should have a developing tank of minimum 10 litres capacity. The container assembly should be made of material, which is non-corrosive, and of latest technology, The main drive should stop automatically during isle time to prevent wear and tear and remain ready in standby, Microprocessor controlled switches and thermostatic control should be there for maintenance of temperature of developer between 30 deg. C and 35 deg, C. The replenishment containers should have capacity of at least 25 liters and the replenishment should be automatic, Wall mounting facility with film output through wall should be possible with the unit, the unit should be quoted with appropriate voltage 220 stabilizer and all essential & optional accessories. Specify the power requirement. | Set | √ | √ | √ |  | where there is X-ray services in DH
| X- Ray film dryer : 25 films drying capacity, 306/316 SS make for anti corrosion and rust free, power supply- 220V-230V, 50Hz,power cord of round 3 pin plug. MCB controlled and on / off switch. Extra heater to be provided. | set | √ | √ | √ | √
| Computed Radiography (CR): GENERAL: Drop-and-go cassette buffer: 5 cassettes of mixed sizes input buffer and 5 cassettes of mixed sizes output buffer. Throughput:35 x 43 cm (14 x 17 inch) = approx. 83 plates/hour. Display for status and error indication: • LCD touch screen.• LED status indicator. Greyscale resolution: • Output to processor: 16 bits/pixel square root compressed. Dimensions and weight: • Covered floor space: (W x D x H): 66 x 51 x 123 cm (26 x 20 x 48.4 inch). • Output buffer included: (W x D x H): 115 x 51 x 123 cm (5.3 x 20 x 48.4 inch). • Weight: approx.: 180 kg (397 lbs). Configuration requirements: • NX, • ID tablet, • CR HD5.0 Detectors and Cassettes, • CR MD4.0R Plates and Cassettes. Power: • 220 - 240V/50-60Hz Standby 87W, peak. 590W, fuse 16A. • 120V/60Hz (USA) Standby 92W, peak. 621W, fuse 15A. • 100V/60Hz (Japan) Standby 92W, peak. 621W, fuse 15A. Environmental conditions DX-G digitizer: • Temperature: 15 - 30 ˚C (59 - 86°F). •Humidity: 15 - 75% RH. • EMC compliant with IEC 60601-1-2. • Rate of change of temperature: 0.5°C/minute (0,9 °F). Environmental effects: • Noise level: max. 65 dB (A). •Heat dissipation: standby 92W, continuous operation 242W. SAFETY Approvals: • ETL classified CUS, CE. Transport details: • Temperature: -25 to +55°C (-4 to 131°F), -25°C for max. 72 hours, +55°C for max. 96 hours. • Humidity: 5 - 95% RH. | Set | √ | √ | √
| Digital Radiography (DR) Machine: 1. Microprocessor controlled high frequency generator (80kW) with Automatic Exposure Control – 1. 2. High speed 600 KHU Rotating Anode X-ray Tube – 1 No. 3. Multi-Leaf collimator – 1 No. 4. Vertical stand for the Digital Detector – 1 No. 5. 3D Ceiling Mounted Tube (CMT) with Servo – 1 No. 6. IDC X3C (3K X 3K) Digital Detector with flat Cesium lodide Scintillator Panel fully integrated with the system – 1 No. 7. Patient table – mobile carbon fiber flat table top with locking wheels – 1 No. 8. Grid – 1 No. A) 180 cm carbon fiber grid (70 Line/cm and ratio of 13:1) – 1 No. B) 100 cm carbon fiber grid (70 Line/cm and a ratio of 13:1) – 1 No. 9. Technologist work station consisting of i) Pentium 4 Duo-processor 3.2 Ghz. ii) 2 GB RAM. iii)160 GB SATA Raud 1 hard Disk drive. iv) 0 Colour flat panel display monitor 1600 x 1200. v) Ethernet 100/1000 Mbps for networking with PACS/RIS/HIS. vi) DV cum CD writer for image exporting. vii) Magellan 3.0 work List software. viii) Built-in Dicom 3.0 enabled functionally. ix)Auto stitching software for special X-rays such as scoliosis and long limbs. Accessories: 10.Konica Dry Laser Camera with 3 online film sizes – 1 No. 11. Voltage stabilizer for the X-ray system – 1 No. 12. Online UPS for the computer and workstation with 30 min. Backup – 1 No. 13.Titanium Apron – 2 Nos. 14. Lead Glass 100 cm x 120 cm 1 No. 15. Book Clarks 100 cm x 120 cm – Each. 16. 1 Box of films for the standard sizes chosen – 1 box for each size. 17. Writable CD/R -1000 Nos. 18. Footrest – 1 No. 19. Hand Grip – 1 No. 20. Abdominal compression binder – 1 No. 21. Footstep – 1 No. 22. Arm rest – 1 No | Set | √ | √
| Cassette 12 inch x 10 inch | Piece | √ | √ | √
| Cassette 12 inch x 15 inch | Piece | √ | √ | √
| Cassette 14 inch x 17 inch | Piece | √ | √ | √
| Lead goggle | Piece | √ | √
| Thyroid shield | Piece | √ | √ | √
| Digital X-Ray film, 14x17 (125 sheets in one packet), compatible with dry view camera) | Packet of 125 sheets | √ | √
| Digital X-Ray film, 11x14 (125 sheets in one packet), compatible with dry view camera) | Packet of 125 sheets | √ | √
| Digital X-Ray film, 8x10 (125 sheets in one packet), compatible with dry view camera) | Packet | √ | √
| Cassette pass box - 2 doors ( when one door opens othecassette to be fit in. 14x17 r shuold be automatically locked)size: | piece | √ | √ | √ | √
| Chest stand ss | set | √ | √ | √ | √
| Safety panel ( Three panels) | set | √ | √ | √ | √
| View box, double films | set | √ | √ | √ | √
| view box, 3 films | set | √ | √ | √ | √
| washing tank with lid, ss 13. 5 Ltr | piece | √ | √ | √ | √
| washing tank with lid, ss 22 Ltr | piece | √ | √ | √ | √
| washing tang with lid, 45 Ltr | piece | √ | √ | √ | √
| Film 12 inch x 10 inch (packet of 50 sheets) | packet | √ | √ | √ | √
| Fiml 12 inch X 15 inch (packet of 50 sheets) | packet | √ | √ | √ | √
| Film 14 inch X 17 inch (packet of 50 sheets) | packet | √ | √ | √ | √
| Fixer 13 . 5 Ltr | packet | √ | √ | √ | √
| Developer 13.5 Ltr | packet | √ | √ | √ | √
| Replenisher 4.5 Ltr | packet | √ | √ | √ | √
| Intensifying screen (speed 400)10inch X 12inch | piece | √ | √ | √ | √
| Intensifying screen (speed 400)12inch X 15inch | piece | √ | √ | √ | √
| Intensifying screen (speed 400)14inch X 117inch | piece | √ | √ | √ | √
| Clip Hanger 10 inch X 12 Inch SS | piece | √ | √ | √ | √
| Clip Hanger 12inch X 15 Inch SS | piece | √ | √ | √ | √
| Clip Hanger 14 inch X 17 Inch SS | piece | √ | √ | √ | √
| Sationary grid 6:1 10inch X 12 inch | piece | √ | √ | √ | √
| Sationary grid 6:1 12inch X 15 inch | piece | √ | √ | √ | √
| Sationary grid 6:1 14inch X 17 inch | piece | √ | √ | √ | √
| Half film blocker 12 inch X 15 inch | piece | √ | √ | √ | √
| Half film blocker 14 inch X 17 inch | piece | √ | √ | √ | √
| Abdominal shield | piece | √ | √ | √ | √
| lead apron | piece | √ | √ | √ | √
| Lead figures, A-Z | set | √ | √ | √ | √
| Lead figures, R-L | set | √ | √ | √ | √
| Lead numbers, 0-9 | set | √ | √ | √ | √
| Glass marking pencil, white |  | √ | √ | √ | √
| Safe light holder : with filter and bulb | set | √ | √ | √ | √
Computed Tomography
Consumables
| Film - dry view laser imaging film: Size 14 x 17 inch (compatible with the machine) | Packet of 125 sheets | √
| Extension line CL Line 155 cm | piece | √
Equipment
| CT Machine 16/64 slice Specification: Brilliance CT 16 slice: Infant cradle, Extended Brilliance Workspace, Bone mineral analysis package PC: HP dx 2000PC, Color laser printer HP 3800n, Kodak Dry view laser camera 8900, Stabilizer, UPS with 30 minutes backup – 125KVA, Pressure injector: MEDRAD Stellant, Defibrillator: Heart start XL, Separate PC for conferencing and teaching aids, Anesthetic machine with pediatric and adult ventilator. Cardiac monitor with stand and Pulse oximeter compatible to the machine. Scan tools Pro, RRCV tools kit, Functional CT/MR image fusion, CV pro analysis package, Virtual colonoscopy, Dental planning package, Lung disease quantification package, Bone mineral analysis, (Installed in Computer). | Set | √
| Operator’s Manuals/ Documentations: Bone mineral analysis program (Q-MAP II) CT operator’s manual, Brilliance CT – Instruction for use – English Volume -1, Brilliance CT – Review – Volume – 2, Brilliance CT – Basic analysis – English – Volume – 3, Brilliance CT – Advanced analysis – English – Volume – 4, Extended Brilliance Workspace – Instruction for use – Volume – 1, Extended Brilliance Workspace – Viewing modes – Volume – 2, Extended Brilliance Workspace – Analysis – Volume – 3, Extended Brilliance Workspace – View forum – Volume – 4, Extended Brilliance Workspace - Supplement version 2.1 Instruction for use, Extended Brilliance Workspace – Cardiac application guide, Brilliance CT – 16 power configuration: Cardiac applications, CT patient log, Kodak Dry view 8900 laser imager – Release 3: User guide, Stellant CT injection system: operation manual, M3/M4 Monitors (M3046A) Measurement server: Instruction for use, Heart start XL - Instruction for use, Heart start XL – User training workbook, Oxywave pulse oximeter: Operation manual, Drager Anesthesia Ventilator, Model Fabius-CE: Operation manual, Drager Anesthesia Ventilator, Model Fabius-CE: Medibus for Fabius CE. | Set | √
| CDs/MODs: Brilliance_ Host_ OS_2.0.0.10, Brilliance CT- Host 2.0.2.13005, Brilliance CT- incremental 2.0.2.13006, Brilliance CT- versions 2.0.2. 13011, CIRS – 4.1.19.5, CIRS- Service- 4.1.19.5, Brilliance CT- Firmware-2.0.2.13001, A plane collimator calibration files: S/N 1758, Systems Diagnostic utilities: Revision:014, Dell Precision Workstation resource CD. | Set | √
| Extended Brilliance Workspace V2.1.1 –Software CD set: P/N 4550 122 02131: Unattended V1.25 – EBW-V2.01: Disk: 1of 6, Environmental creator –EBW-V2.0.1: Disk: 2 of 6, COTS (3rd party) V2.0.1- EBW-V2.0.1: Disk: 3 of 6, Application software – EBW-V2.0.1: Disk: 4 of 6, View forum – EBW- V2.0.1: Disk: 5 of 6, Mandatory Add-on – EBW – V2.0.1: Disk: 6 of 6, Update version –V2.0.11- EBW, EBW-V2.1.0, EBW-V2.1.1, Brilliance Hot fix, NVIDIA FX540 Drivers CD For Dell 670-EBW. | Set | √
| Drivers & Utilities – Dell Precision Workstation Resource CD, Q-MAP II Version 3.15, Q-MAP II Instructional CD: Rev. A, Merge DSP 1.0.2 Dicom storage provider, Microsoft Windows XP Professional Service Pack2: Operating system CD – For HP PC, Microsoft Windows XP Professional Service Pack2: Version 2002 – Compaq PC, HP/Compaq DX 2000 MT drivers, HP restore Plus: For HP dx 2000 MT PC, HP Business Desktop documentation library CD: dx2000 model, HP color LaserJet 3800, Norton Antivirus, LCD monitor – Electronics user’s manual: Version – H6195, Kodak Dry view 8900 laser imager – Version 1.0, Default grid voltage file: Tube No. 105702 MRC 600, Heart start XL: User training CD ROM, Sony magneto optical Disk – 9.1GB: Model EDM 9100C. | Set | √
| Cables and Hardware accessories/spares: Spike buster/Extension box with six outlets, Computer power cables, RJ 45 straight network – long, RJ 45 straight network – short, RJ 45 cross straight network cables – short, DVI Video cable for Philips LCD monitor, VGA cable for Philips LCD monitor, USB cable, Audio cable, Examination table lifting service bracket, Kodak dry view laser camera connector, C7 PCB for consul servo voltage stabilizer, C7 trip PCB for consul servo voltage stabilizer, Transformer CC10 for consul servo voltage stabilizer, Transformer CC11 for consul servo voltage stabilizer, 12 teeth pinion for consul servo voltage stabilizer, Carbon brush for servo voltage stabilizer, Packing sims-1mm, Packing sims-2mm, CT tube repairing kit, Laser aperture tool, Cable for QMAP PC, Parallel cable, Earth cables for ECG Monitor, Binder, Dongle for QMAP PC, VGA adaptor HD15/DB15. | Set | √
| Dry view camera compatable to CT (laser imaging) | piece | √
| Patient Positioning straps (Belts) | Set | √
| UPS with 30 minutes backup-125 KVA | Set | √
| UPS Battery: US-12-125: 12V-125AH. Valve regulated: Constant voltage recharge. Lead acid: Boost-equalize – 14.5-15 OV. Rechargeable battery: Float – 13.5-13 BV. Max. – 37.50 A | Set | √
| Automatic injector | Set | √
CT Accessories
| Extended Phantom performance kit | Kit | √
| Eye Phantom | Kit | √
| Phantom holder | Set | √
| CT slicker cushion | Set | √
| Foot rest extension | Set | √
| Coronal head rest assembly | Set | √
| Standard head rest assembly | Set | √
| Kit couch cushion | Set | √
| Patient restraint upgrade kit | Kit | √
CT and MRI
| Lazer camera compatible to CT-MRI | Set | √
| Stellant syringes: content, 2-200ml syringes, 1-60 inch T tube connector, 1 quick filled tube | set | √
Interventional Radiology
For Angiography (Tace, Bronchial, Uterine and Renal Artery Embolization)
| Puncture Needle 18, 19, 20G sulflo IV Catheter with radio-opaque ETFE. (Length- 2 ¾ inch.), double wall puncture. | Set | √
| Guide wire (puncture) with size 0.035. | Set | √
| Vascular sheath 5, 7 Fr. | Set | √
| Single wall puncture needle 18-21G. | Piece | √
| Cobra catheter 4 & 5 Fr (length 65, 100 cm) | Piece | √
| Simon (Simon II) catheter 4. 5 Fr. (length 65, 100 cm) | Piece | √
| Hydrophilic guide wire size0.032, 0.035 inches (length 80, 100 & 125 cm) | Piece | √
| Angiography suite with Digital Subtraction facility | Set | √
| Simon (I) catheter (4Fr, 5 Fr) length 65, 100 cm. | Piece | √
| Yashiro catheter (4, 5Fr) length 70, 100 cm | Piece | √
| Head hunter (4, 5Fr) length 70, 100 cm. | Piece | √
| Pigtail (intravascular) 6- 8 Fr. | Piece | √
For Ablation of liver tumour
| RFA generator | Set | √
| RFA electrodes | Set | √
Percutaneouse Drainage
| Chiba Needle 18 Gauge (15 cm) | Piece | √
| Chiba Needle 18 Gauge (20 cm) | Piece | √
| Chiba Needle 20 Gauge (15cm) | Piece | √
| Chiba Needle 20 Gauge (20cm) | Piece | √
Guide wires
| Bentson J TIP sizes (0.035, 0038 inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
| Bentson straight tip (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
| Bentson supper stiff J TIP (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
| Amplatz extra stiff J tip (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
| Amplatz extra stiff straight tip (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
| Facial dilators (plastic) standard length: 6 Fr, 7 Fr, 8 Fr, 9 Fr, 10 Fr, 12 Fr. | Piece | √
| Self retaining loop catheter (Pigtail) standard length: 8 Fr, 10 Fr, 12 Fr, 14 Fr. | Piece | √
| Micro punctures (billiary drainage set) consist of: acoustic needle, J-tip guide wire, coaxial introducer, Needle with stylet, extra stiff G W with a fluffy tip | Set | √
| Percutaneous B: spring activated core biopsy needle with coaxial system, 18 gauge (15 cm), 18 gauge (20 cm), 20 gauge (15 cm), 20 gauge (20 cm), spinal needle (15, 20 cm) 18, 20, 22, 23, & 25 G | Set | √
| Mammography Machine, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
Magnetic Resonance Imaging
Equipments
| MRI friendly Suction machine 1/2 HP, mobile fibre casting pressure tube with pressure regulator tube. Twin suction bottle of 2-3 Ltrs capacity. Display of vacuum from 100 to 700mm HG. (MRI compatible) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √
| CO2 cylinder for MRI 'A' type (Re-filling) | Piece | √
| MRI compatible wheel chair | Piece | √
| MRI compatible Patient trolley | Piece | √
| MRI compatible IV Stand | Piece | √
| MRI compatible Anaesthesia machine Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √
| MRI compatible Pulse Oximeter, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √
| Dry view camera, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Piece | √
| MRI UPS | Set | √
| MRI Machine with automatic injector | Set | √
| MRI machine 1.5 Tesla specifications: ADW CPU, ADW monitor, mouse and mouse pad, metal detectors, Kodak dry view 8900 camera, report card CPU, Monitor, key board, CCTV, newmatic slimline head sets, microphone, Onkyo music system with speakers, AM/FM antenna, GOC computer, Monitor, PAC monitor, spectris solaris pressure injector 2 batteries and charger, SCSI tower, keypad, mouse and pad, gradient chiller, 1.5T sumitomo compressor, Cryo chiller, Quad knee coil, 8 channel brain coil, 1.5T split head coil, 8 channel body array, 8 channel Neuro Vascular array, 8 channel CTL coil, 8 channel anterior bridge, general purpose flex coil, phased array shoulder coil, 8 channel breast coil, QD plug coil adaptor extremity, coil ID connector, 1.5T peripheral Vascular coil, 1.5T 5 inch round coil, 1.5T 3 inch round coil, Dual array adaptor, MR compatible adaptor, MR compatible anaesthesia m/c, phantom cart with LV shim phantom, DQA phantom with loader, TLT head sphere, ball phantom, upper anular phantom, CTL phantom, SNR phantom head SNR phantom, ball phantom with base, 8 channel NV array phantom, loading phantom, service ladder, newmatic slimline head sets, invivo pulse oxymeter, ERU, Oxygem monitor with sensor, Ethemet awitch, 1.5T CX K4150 magnet for Twinspeed gradient module, GEDE 100KVA UPS, PDU/HFD cabinet, TAC cabinet, RFS cabinet, CO2 cylinder for door locking, magnet monitor, system main disconnect panel, UPS batteries, MOD media 2.3 GB, patient squeeze bulb audio alarm, signa excite 1.5T phased array cables for Twinspeed, FR cage, viewing box, MRI 3 TR ductable split, TR, 1.5 TR, CTT Linux 4.3.13, excite HD M4 Applications, Gamma correction software, 11.0 M4 patch 00-PA-CD, HD MR Driver Module HVPS, Report Card M3A Installation 2.0, 11.0 M4 patch 00-PA-CD, Functool demo exams, Volume Viewer Demo exam, Functool Applications, Reporting tool, Volume Viewer Applications, MR pating 1.1, AW fast load DVD, AW 4.2P software, 11.0 M4 Prose/Functool Patch CD, 10.0 8 Channel NV Array QA tool, AW demo exam on CDROM, 1.5T Signa Excite 11.0 LRG, Split Head Coil LRG, AW software Config Floppy, Laryngoscope ADULT, Non-magnetic Oxygen Cylinder, Emergency light, HP Color Laser Jet 3800n Printer, Non-Magnetic IV Stand, Stethoscope, Sphygnomanometer, Diagnostic Pro Advantage, MR Machine Datex Ohmeda. | Set | √
| UPS 80KVA | Set | √
| MR compatible Laryngoscope for adult and children | Set | √
| MR compatible anesthesia machine for adult and paediatrics | Set | √
Consumables
| Film - dry view laser imaging film: Size 14 x 17 inch (compatible with the machine) | Packet of 125 sheets | √
| Ear Plugs (spongy foam) | Piece | √
| MRI compatible ECG leads | Piece | √
| Patient Positioning straps (Belts) | Set | √
Accessories
| MRI UPS Battery: Highlit, 12xLB420, 12v 416w/cell/15 min to 1.67v/cell at 25 degree C (77 degree F), Float voltage 13.62v at 20 degree C (68 degree F), 13.56v at 25 degree C (77 degree F), connection torque 8-9Nm/70-80 mlb. | Set | √
| Hand Held Meta Detector: Hand held super scan metal detector weighing 10 oz. | Set | √
| 1.5T 8 Channel Brain Coil: This eight element head coil for the use with 8 channel EXCITE 1.5T MR system provides unprecedented high SNR. Critically important for high throughput, the coil is optimized for ASSET parallel imaging with a speedup factor of two. With a 24 cm diameter and S/I coverage of 30 cm, the coil features an open, patient friendly design with a lookout mirror to let patients view the world outside the system. This coil will change your expectations for clinical brain imaging intracranial MRA and paediatric imaging (8-ch option is mandatory) | Set | √
| 1.5T Channel CTL Array: This 12 element, multi-station array for use with 8-channel EXCITE 1.5 T MR system delivers extraordinarily high SNR, spatial resolution and coverage over the full range of spine, soft tissue neck and carotid studies. It’s packaged to conform to the spin’s normal curvature and includes a patients comfort pad and restraint. Patients can be positioned head first or feet first and the coil’s extensive coverage 75 cm in the S/I direction-accommodates imaging of the entire spine, making it a truly high productivity solution for your site. This coil’s unique split top design extends its clinical flexibility. The removable top facilities routine neck imaging, including the anterior cervical spine and neck and eliminates the need to switch to a dedicated neck coil. Coil markers make accurate positioning at imaging isocenter surprisingly quick and easy. The result improved productivity and patient throughput. This package includes applications manuals, operating instructions, SNR testing phantoms and formfitting patients comfort pads. | Set | √
| 1.5 T Open Neurovascular Array: With a four-channel, quadrature phased array design and extended, 48 cm S/I coverage, this receive only coil is specially suited to high-SNR imaging of the neck, head, brain and vascular system. It features a two piece removable top with many openings and an adjustable mirror system to reduce claustrophobia. | Set | √
| 1.5T 8 Channel body Array: This is a 12 element coil designed for use in creating extraordinary images of the chest, abdomen and pelvis on 8 channel EXCITE 1.5T MR system. This coil offers such extensive coverage that repositioning of the coil for combined abdomen-pelvis exams typically is not necessary. It provides higher sensitivity than anybody coil offered previously by GE, to enable the highest possible signal-to-noise ratio. And particularly important for body imaging, the coil was optimized for the use of ASSET98 ch option is mandatory | Set | √
| USAI 1.5T Phased Array shoulder coil: The USAI 1.5 T shoulder array coil offers the increased signal of the phased array coil along with a unique sleeve design that gives the user exceptional joint imaging capabilities. Added benefits of teh coil include. This shoulder coil is able to image the shoulder region providing images with clear definition of the shoulder joint, specifically the head of the humerus, clavical, acromion bone, supraspinatus muscle and ligaments. This shoulder coil is able to image the shoulder region providing images with clear definition of the shoulder joint, specifically the head of the humerus, clavical, acromion bone, supraspinatus muscle and ligaments. This shoulder coil is able to image the shoulder region providing images with clear definition of the shoulder joint, specifically the head of the humerus, clavical, acromion bone, supraspinatus muscle and ligaments. The shoulder coil is able to image both shoulders (not at the same time). The shoulder coil provides patients comfort for patient imaging of up to 40 minutes procedures. The shoulder coil is ASSET compatible and delivers optimum FATSAT imaging. | Set | √
| Peripheral Vascular phased Array Coil: Coil is a specially designed 12 elements phased array coil with 1.26 meter extended coverage for peripheral vascular run-offs from renal to dorsal digital arteries with smartstep2000. The other applications include visualization of lower abdomen and lower extremities. High signal uniformity. 1.26 meter coverage in S/I direction (3 x 44 cm steps with 3 cm overlap). Physical dimension 1.37 m long and 74 cm wide. Flexible and 2 part coil structure for close proximity to patient. Received only phased array design. Wide positioning strips with Velcro captures is provided with the coil to hold the coil closely to the peripheral leg vessels into a common plane. Must ensure that M1087CY pre amp kit also ordered with the system. Includes signal to noise test scanning phantom for image quality assurance procedures. | Set | √
| MAI QUAD EXTREME COIL W ID | Set | √
| Open breast phased Array coil: Coil is specially designed for imaging and visualization of the female breast and auxiliary tissue. It is possible to perform either unilateral imaging of the either left or right breast or bilateral imaging of the both breast. | Set | √
| 3 Inch circular general purpose coil: This is a 3” circular general purpose coil for applications in orbits, IAC, ankle and wrist. Can be combined with another 3 inch GP coil for simultaneous high resolution dual imaging of TMJ, Orbits and IAC (Dual array adapter is mandatory for such applications). | Set | √
| 1.5T GP flex coil: Coil is designed to optimize imaging of irregular anatomy such as the branchial plexus, hip, thigh, elbow, large knees, ankle and foot. | Set | √
| 5 inch circular general purpose coil: This coil is useful for a wide range of small parts applications-including the orbits, IAC, ankles and wrist. And with the help of a dual array adaptor, two of the coils can be used concurrently for simultaneous high-resolution of TMJ, orbits or IAC. | Set | √
| Dual Array Adapter: Allows for combining two receive only non phased array coils. This helps in high SNR imaging of symmetrical anatomy such as TMJ | Set | √
Radiology Ultrasound
Equipments
| USG Machine - Type II: 'USG Machine - Type II: Specification: Full backlit Alphanumeric keyboard, telescoping, tilt, swivel, high adjustment. The system includes patient entry preset menus archiving, edge enhancement, frame averaging, storage and recall. Features: Truscan Architecture - software that gives computational power, image manipulation capability workflow flexibility and product upgradeability. Truaccess - a raw data processing technology for image enhancement. Codescan - Improve image qua. Transducer : 3.5 - 5 MHz Convex probe for ABD, OB GYN, Linear probe 7 - 10 MHz Additional probe Port. Accessories: Footswitch, cable set for B/W printer, Fixture for B/W printer, 2KVA UPS. Measurement capabilities: Distance, Circumference, area and volume. Details: Multipurpose imaging System design for abdominal, vascular, obstetrics, gynecology, cardiology, neonatology, urology, transcranial and small parts application. Servo controlled Voltage stabilizer. | Set |  |  | √
| USG Machine - Type III:'USG Machine - Type III: Transducer : 3.5 - 5 MHz Convex probe for ABD, OB GYN, Linear probe 7 - 10 MHz Additional probe Port. Details: 17" Flat screen colour doppler with comprehensve archiving facilities. Recording facilities: VCR, CDROM/FFD/DVD/UBS thum drive. Features: State of the art system with DVD and standby mode, coded contrast imaging, cross beam, report writer, stress echo package and PDF. Scanning modes include B-mode, M-mode, colour flow, pulsed wave, power doppler and real time triplex mode. Accessories : Footswitch, cable set for B/W printer, cable set for VCR, Fixture for B/W printer, VCR-TV Monitor and CD writer built in, 2KVA UPS and Fixture for VCR/Colour printer. Endovaginal probe, endorectal probe with 7 to 10 MHz. TRUS guided biopsy probe.Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top and suitable UPS | Set | √ | √ | √ |  | where there is X-ray services in DH
| Ultrasound Machines (Portable with laptop type), with color Doppler facility and atleats 2 probes (convex and linear), with complete set of printer. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. | set | √ | √ | √
Consumables
| Thermal printing paper type I (Sony), ultra sound (110mmx20m) | Roll | √ | √ | √
| Ultrasound gel | 250 ml | √ | √ | √
| Probe: curvi linear for trans abdomin ultrasound (compatable with existing machine) |  | √ | √ | √
Probe: linear (compatable with existing machine, mention the Specs)
| Probe: transvaginal (compatable with existing machine) |  | √ | √ | √
| Printer for ultrasound machine |  | √ | √ | √
Surgery Department
Endoscopy
Equipments
| Argon Plasma Coagulation (APC) unit: For open Endoscopic and Laparoscopic. Power supply 220-240v, 50Hz. All power cord should be of round 3 pin plug top. | Piece | √ | √
| Argon Gas Cylinders: Compatible to the APC Generator | Piece | √ | √
| Argon Plasma Coagulation (APC) hand piece: compatible to the energy generator | Piece | √ | √
| Video endoscope Medical Monitor: Flat screen, 14” to 16” | Piece | √ | √
| Video endoscope Processor, 3 chips camera high definition | Piece | √ | √
| Gastro-Scope compatible to above processor | Piece | √ | √
| Colono scope compatible to the above processor | Piece | √ | √
| UPS for the entire system | Piece | √ | √
Accessories
| Biopsy forcep for upper GI ENDOSCOPY | Piece | √ | √
| Biopsy forcep for Colonoscope | Piece | √ | √
| Hemoclip for GI bleeding | Set | √ | √
| Multiband Ligator for Esophageal verices | Piece | √ | √
| PEG (Endoscopic Gastrostomy Set) | Set | √ | √
| Esopageal dilator set | Set | √ | √
| Esopageal stents | Piece | √ | √
| Esopageal varices sclerotherapy injector (needle) | Piece | √ | √
Laproscopy
Equipments
| Laproscopic camera 3 chips high definition: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
| Laproscopic camera unit: flat screen, 18 inches | Piece | √ | √
| Laproscopic Medical Monitor: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
| Laproscopic CO2 insufflator: Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √
| Laproscopic Light source: compatible to the above light source | Piece | √ | √
| Laproscopic Light cable: 0 degree,10 mm | Piece | √ | √
| Laproscope lens: 30 degree,10 mm | Piece | √ | √
| Laproscope lens | Piece | √ | √
Consumables
| Cleaning brush: Size: 16 mm | Piece | √ | √
| Disposable 12 mm abdominal trocar | Piece | √ | √
| Disposable 10 mm abdominal trocar | Piece | √ | √
| Reducer: (from 10 mm to 5 mm ) compatible to Trocar | Piece | √ | √
| Reducer: (from 12 mm to 5 mm ) compatible to Trocar | Piece | √ | √
| Disposable 5mm abdominal trocar | Piece | √ | √
| Laparoscopic Bipolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
| Laparoscopic Monopolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
| Liga clip: LT 300 | Piece | √ | √
| Liga clip: LT 400 | Piece | √ | √
| Oil for the instruments maintenance | ml | √ | √
| Single clip applicator: Titanium clip mediun for LT 300 (Ethicon) | Piece | √ | √
| Single clip applicator: LT 400 ( large Titanium clip) | Piece | √ | √
| Claw extractor with monopolar with Ratiched handle 33.2 cm, size 5mm | Piece | √ | √
| Curved dissector with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
| Grasper with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
| Laparoscopic Aspiration needle: size bigger than 20 Fr. | Piece | √ | √
| Curved dissector with Bipolar cautery 33.2 cm: size 5mm | Piece | √ | √
| Laproscopic Babcock (10 mm), size 10mm | Piece | √ | √
| Curved scissor with monopolar cautry | Piece | √ | √
| Gall bladder holding forceps fenestrated (grasping forceps). 5mm, 33.2 cm | Piece | √ | √
| Irrigation and suction cannula | Piece | √ | √
| Laproscopic Hook dissector : 5mm diameter shaft length 34 cm | Piece | √ | √
| Merry lamp forcep: 32.2 cm, 5 mm | Piece | √ | √
| Metal trocar: size 5mm | Piece | √ | √
| Metal trocar: size 10mm | Piece | √ | √
Accessories
| CO2 regulator for D-type cylinder | Piece | √ | √ |  |  | for D-type cylinder
| CO2 regulator for B-type cylinder | Piece | √ | √ |  |  | for B-type cylinder
| CO2 cylinder D type | Piece | √ | √ |  |  | D-type
| CO2 cylinder B type | Piece | √ | √ |  |  | B-type
| High pressure Hose (CO2), connection length 10m | Piece | √ | √ |  |  | connection length10 m.
Neurology
Equipment
| Intra Cranial Pressure Monitoring system with hydrophobic filter with two convient sites, sliding chamber, locking bracket, with drainage bag of capacity 700cc and system panel and patient connecting wire with barium impregnated silicon catheter tubing and ventricular silicon catheter of internal diameter 1.5 mm, outer diameter 2.8 mm, length 35 cm with CSF flow control valve of medium pressure 12 cm size. | Set | √
| Endovenous LASER System: Laser Type Diode laser, CW, class IV, Centre Wavelength 1470 nm, Maximum Power 12 W to the laser output port, Aiming Beam Red Class IIIa diode laser, ≤5 mW at port, wavelength 635–655 nm, Pulse Interval 100 ms–1000 ms, Power Supply 100–240 V AC, 50–60 Hz, 500 VA max, Dimensions (H x W x D) 150 mm x 370 mm x 325 mm (± 5 mm), Weight 12 kg max. DIODE LASER:VenaCure 1470 nm Laser MMC Card-64 MB Interlock Connector Foot Switch Travel Case Spare Key-VenaCure 1470nm Laser. ACCESSORIES: USB Memory Card Reader/Writer Warning Signs Premium Laser Safety Glasses Standard Laser Safety Glasses Emergency Reset Device User Manual Spare Fuses. Procedure Kit set, Micro Introducer Kit set and catheters | set | √
| May Field Cranial stabilization skull clamps, Codman 360 degree rotation of rocker arc with clamp and stainless pins. | Piece | √
| Micro Chair (Hydraulically-operated): Drive System: DC 24V, Elevation/lowering: 44cm-86cm, Back/forth & right left sliding: free movement. Rotation counter and clockwise reclining up to 6 degree from vertical with arm rest. Power consumption 400VA. Power source 100-110V, 50/60Hz. Insulation resistance of 400M ohm, lekage current 50 micro amperes. weighs approximately 130kgs (Should be compatible with microscope leica M651) | Set | √
| Multi-Purpose Head Frame: Basal frame/E.Semicircular frame, Hand rest with slider /instrument support, Spring scalp HoOK retractors (6nos), Cotton plate, Head holder with 4 standard pins, Table attachment, Head pin standard, Self retaining retractors (Large, Small), Soft brain spatulas (Large tip width 6mm, medium tip width 4mm ,small tip width 2mm), Universal attachment 1no, Slide adjuster 1no, -shaped head rest (large and small) 1no Piece, Self retain retractor attachment 1no. | Set | √
| Neuro Surgery Operating Table: Power Source voltage: AC 100V, 220V, Driving voltage: DC 24V, Battery: DC 12V x 6, Power ConsumptionVA (MST- 7000N), Table top: 450mmx1895mm, Elevation Stroke: 480-980mm (High/low speeds), Trendelenburg (high/low speeds): head –down 40 degree Head –up 20 degree. a. Operating Table Horseshoe Head pad - adult (20.5 OD x 7.5 ID x 4.5cm) b. Operating Table Horseshoe Head pad - pediatric (14 OD x 5.7 ID x 3.2cm) c. Operating Table Supine Head pad - 23x25x5.7cm d. Operating Table Lateral Head pad with centre dish - 23x25x2.5cm e. Operating table prone head rest - Medium 22.2x19.7x12cm f. Operating table prone head rest - Medium 22.2x19.7x12cm g. Operating table prone head rest - Large 27x22.9x13.4cm h. Positioning Chest rolls - Adult 50.8x10cm i. Positioning Chest rolls - Paediatric 30.5x5cm. Operation theatre stirrup pads - boot/ stirrup pad set 68.5x36T/12.5B/6cm. | Set | √
| Neurosurgery operating microscope (portable) | Set | √
| Nitrogen Gas cylinder - D type | Piece | √
| Ultrasonic aspirator system : 1. universal handpiece 2. disposable tips for soft tissue -superlong straight tip, spetzler barracuda large diameter tip, straight tip, straight large diameter tip, straight micro diameter tip, spetzler barracuda tip, ring curette tip; 3. accessories - step torque wrench offset 7mm, cleaning wire, sterilization tray, console 110V(including foot pedal and IV pole), plastic suction canister reusable, disposable suction canister liner, disposable tubing set & extender filter tubing, cart. | Set | √
| Neuro-endoscopy set consisting of adult and child:Neuroendoscopy Instrument Set includes Injection Needle 1.7mm; Deflecting Mechanism for LASER probe and ring-grip handle, diameter 2.9mm,length 38cm; Grasping Forceps-diameter 2.7mm, working lenght 30cm; Biopsy Forceps with working length 30cm; Scissors- pointed with single action jaws(diameter 2.7mm) and doble action jaws(1.7mm) with working length 30cm; Forceps for ventriculostomy,flexible with diameter 1mm; Coagulating Electrodes-both unipolar and bipolar with diameters 1.7mm; Autoclavable Holding Systems consisting of socket, straight articulated stand, and clamping jaw; Hopkins Wide Angle and Straight Forward Telescopes 6degree,with angled eyepiece, with instrument channel diameter 3mm, lenght 15cm, autoclavable, fiber optic light tranismission incorporated, color code:green; Hopkins Wide Angle forward-oblique telescope 30degrees diameter 4mm, length 18cm color code:red; and Straight Forward TelescopesOperating Sheaths- outer diameter 6.5mm, working length 13cm; Obturators; Telescope Bridges; Irrigation Tube with Luer-lock; suction cather disposable; and Adapters; containers for sterilization for teklescopes,angled telescopes and instruments. | Set | √
| Surgical Clipper with fixed head new, unique cordless/ coded operation | Set | √
| Electric Craniotome: Motor speed should be atleast 75,000 rpm at 8 bar pressure, motor should be light(preferable 85g or less), main motor unit should be detachable from cable, straight and angled attachment of various lengths should be available for cranial and spinal surgery, change of hand piece with mounted tool should be possible, motor should be converted to an angulated position with or without an adaptor, sound level should be very low, no intra-operative oiling of motor should be necessary, duraguard should be detachable from craniotome hand piece, so that same hand piece can be used for paediatrics, adults and spine surgery, quick coupling attachment should be available, single use and reusable burrs should be available, sterilization through flash or regular steam autoclave, perforator driver with cutter should be available , should be able to use saw hand piece with same system, in-built or optional irrigation pump should be available, five years warranty followed by AMC quote for next five years. Power supply 210-240V. The following accessories should be included: (Should be compatible with Leica microscope 651) | Set | √
| Following accessories should be included with the machine (Electric Craniotome): A. Hand Piece: Hand Piece, straight, 120 mm. Hand Piece, straight, 90 mm. Hand Piece, straight, 160 mm. B. Craniotomy attachment: Craniotome hand piece. Fixed duraguard, adult. Fixed duraguard, paediatric. C. Craniotome cutter: Craniotome cutter, paediatric. Craniotome cutter, adult. D. Perforator: Perforator driver. Cranial perforator, 9 x 12 mm, Hudson type. Cranial perforator, 6/9 mm, Hudson type. Hudson chuck. Spare cutter for perforator, 9 x 12 mm. Spare cutter for perforator, 6 x 9 mm. E. Burrs: Rosen Burr D 3.1 mm for 120 mm hand piece. Diamond burr D 3.2 mm for 120 mm hand piece. Barrel burr D 4.0 mm for 160 mm hand piece. Barrel burr D 4.0 mm for 120 mm hand piece. Neuro cutter D 2.3 for 120 mm hand piece. Neuro cutter D 3.1 for 120 mm hand piece. Neuro cutter D 3.1 for 160 mm hand piece. Acorn burr D 6.0 mm for 90 mm hand piece. Pin point cutter D 1.5 for 120 mm hand piece. Twist drill D 1.5 mm for 90 mm hand piece. F. Micro sagital saw attachment: Micro sagital saw pencil shape. Saw blade for micro sagital saw 9/13/0.3/0.3 mm. G. Storage and maintenance: Oily spray for high speed motor and hand pieces. Oily spray for perforator. Perforator basket with covering lid with holders for motors, all hand pieces, hose tools and all other accessories. | Piece | √
| Operating binocular Loupes with 4X times magnification | piece | √
| Skull bone fixation system (cranial fixation system) Instruments consisting of applier,holding forceps, pin cutter, removal forceps, storage tray; instrument oil (aerosol);Sterilit instrument oil(bottle) ; Implants with clamp, 11mm (box of 12), clamp, 16mm (box of 6), clamp, 20mm (box of 6) Sterilit instrument oil (bottle) Container for storage tray | set | √
| Negative Pressure Wound Therapy system or Vacuum assisted closure system:Unit Weight with empty canister3.35kg (7.4 lbs with a 500cc empty canister and no solution bag) Dimensions with empty canister217mm x 260mm x 191mm (8.55in x 10.25in X 7.5in) Power SupplyClass I Equipment Ordinary Equipment Battery TypeCustom lithium battery Battery LifeApproximately 6 hours Battery Recharge TimeApproximately 4 hours Battery IndicatorThree-segment; low-battery alarm when <2 hours Electrical Data (Power Supply)Voltage 100 - 240 VAC Frequency 50Hz - 60Hz Power 60W. Other accessories includes Foam dressings , Pad and Tubings. | Set | √
Consumables
| Aneurysm Clips 1/1(mm): Bayonet (s) opening width 11.0(10mm)bayonet (l) opening width 13.0(12mm) | Piece | √
| Aneurysm Clips 1/1(mm): Bayonet opening width 9.0(7mm) | Piece | √
| Aneurysm Clips 1/1(mm): Bent opening width 8.0 (8mm) | Piece | √
| Aneurysm Clips 1/1(mm): Bent, opening width 11.0(12mm) | Piece | √
| Aneurysm Clips 1/1(mm): Side ward bent, opening width 11.0(12mm) | Piece | √
| Aneurysm Clips 1/1(mm): Slightly curved, opening width 10.0(12mm) | Piece | √
| Aneurysm Clips 1/1(mm): Straight (large) opening width 9.0(10mm) | Piece | √
| Aneurysm Clips 1/1(mm): Straight (small) opening width 8.0 (7mm) | Piece | √
| Aneurysm Clips 1/1(mm): Straight, opening width 11.0(12mm) | Piece | √
| Fenestrated large type 1/1 (mm) – Straight (L) opening width 12.0 (12 mm) | Piece | √
| Fenestrated large type 1/1 (mm) – Straight (M) opening width 11 (9 mm) | Piece | √
| Fenestrated large type 1/1 (mm) – Straight (S) opening width 10 (6 mm) | Piece | √
| Cottonoids with X-ray detectable markers & string attached, small 1.5 x 5 cm | Pkt | √
| Cottonoids with X-ray detectable markers & string attached, large 2 x 5 cm | Pkt | √
| Surgical patties for tissue protection (x-ray detectable), size1/2 x 3 | Piece | √
| Operating microscope draping sheet plastic, disposable, compatible with microscope Leica M651 | Piece | √
| Artificial dural graft, expanded polytetrafluoroethylene 4 x 4 cm | Piece | √
| Artificial dural graft, expanded polytetrafluoroethylene 5 x 5 cm | Piece | √
| Low adherent absorbant dressing – 10 x 10 cm | Piece | √
| Low adherent absorbant dressing – 10 x 20 cm | Piece | √
| Hemostatic glue solution made of gelatin or fibrin 2-5ml | Piece | √
| Pure cotton bandage with cohesive effect 15 cm x 3.5 m (stretched) | Piece | √
| Pure cotton bandage with cohesive effect 10 cm x 3.5 m (stretched) | Piece | √
| Neurosurgical patties-sterile packages 5 x 5 cm | Piece | √
| Neurosurgical patties-sterile packages 10 x 10 cm | Piece | √
| Neurosurgical patties-sterile packages 20 x 20 cm | Piece | √
| Cranioplastic kit, sterile (methylmethacrylate medical grade in powder form with liquid polymerizing agent) | Set | √
| Peripheral vascular patch prepared from bovine pericardium, sterile packed, 1 cm x 6 cm | Piece | √
| Peripheral vascular patch prepared from bovine pericardium, sterile packed, 1 cm x 10 cm | Piece | √
| Carotid vascular patch made of collagen coated knitted polyster 2 x 8 cm | Piece | √
| Carotid vascular patch made of collagen coated knitted polyster 12 x 25 cm | Piece | √
| Bard inflation device with 30 atmosphere and 30 cc syringe | Piece | √
| Carotid endarterectomy shunt- T-shaped silicone with 15 cm length and shunt outer diameter of 3 mm distal, 4 mm proximal and bulb proximal and distal outer diameter 4.3 mm | Piece | √
| Fogarty Balloon catheters 2F, with 60 cm length | Piece | √
| Fogarty Balloon catheters 3F, with 60 cm and 80 cm lengths | Piece | √
| Fogarty Balloon catheters 4F, with 40 cm and 80 cm length | Piece | √
| Fogarty Balloon catheters 5F, with 80 cm length | Piece | √
| Fogarty Balloon catheters 6F, with 80 cm length | Piece | √
| Fogarty Balloon catheters 7F, with 80 cm length | Piece | √
| Dacron or PTFE prosthesis grafts 30 cm long and 6 mm diameter | Piece | √
| Venticulo Peritonial (VP) Shunt System Catheters (Silicone rubber) - length 62cm, Internal Diameter 1.3mm, Outer Diameter 2.5mm, Medium Pressure | Piece | √
| Venticulo Peritonial (VP) Shunt System Catheters (Silicone rubber) - length 90cm, Internal Diameter 1.3mm, Outer Diameter 2.5mm, Medium Pressure | Piece | √
| External cerbrospinal fluid drainage system: consisting of one ventricular catheter, one collecting system consisting of three way stop cock, connecting tube, drip chamber and collection bag consisting of two chambers and a folding scale. | set | √
| Transparent film dressing frame style 6x7cm | piece | √
| Transparent dressing with label 10x12cm | piece | √
| Sacrum shaped adhesive hydrocellular foam dressing 17x17cm | piece | √
| Adhesive hydrocellular foam dressing 12.5x12.5cm | piece | √
| Vascular graft for arterio-venous fistula (6mm x 40cm) | piece | √
| Vascular graft for arterio-venous fistula (5mm x 40cm) | piece | √
| Bone wax 2.5gm | piece | √
| Ommaya reservoir set: burr hole style with dome size 2.5cm and with one ventricular catheter | piece | √
| Hydrocellular foam heel dressing | piece | √
Instrunments
| Adson Hypophyseal Forceps, stainless: 5 mm diameter 20.5 cm | Piece | √
| Adson Hypophyseal Forceps, stainless: 8 mm diameter 20.5 cm | Piece | √
| Hypophyseal forceps – 3 mm diameter 17 cm | Piece | √
| Specula for transphenoidal hypophysectomy diameter 15 mm x 70 mm | Piece | √
| Pituitary curette- horizontal 45 °, 18.5 cm | Piece | √
| Pituitary curette- horizontal 90 °, 18.5 cm | Piece | √
| Pituitary curette- vertical 45 °, 18.5 cm | Piece | √
| Pituitary enucleator blunt 24.5 cm | Piece | √
| Bipolar Forceps (non-stick), Bayonet - tip 0.8 mm, 20 cm | Piece | √
| Bipolar Forceps (non-stick), Bayonet right curved - 45 degree tip 0.8, 21cm | Piece | √
| Bipolar Forceps (non-stick), Curved upward - tip 1.8 mm, 20 cm | Piece | √
| Brain spatula, Malleable Large 6 mm with 20 cm | Piece | √
| Brain Spatula, Malleable Medium 4mm with 20cm | Piece | √
| Brain Spatula, Malleable Small 2mm with 20cm | Piece | √
| Cloward Kerrison Rongeur- 40° angled 3.5 mm bite, 17 cm | Piece | √
| Complete clip set with the sterilizable metal case for all the clip sizes | Piece | √
| Cranial Elevator Stainless – Large 21.5 cm, 2 cm, 8.5 mm | Piece | √
| Cranial Elevator Stainless – Small 21.5 cm, 1.7 cm, 7 mm | Piece | √
| Dissecting tumor and tissue grasoing forcep, size-30.48 cm | Piece | √
| Double action Rongeur- Angular, bite 5 mm, length 22 cm | Piece | √
| Dura forceps- 12 cm | Piece | √
| Dura Hook – 13.5 cm | Piece | √
| Dura Hook – 2 hooks 13.5 cm | Piece | √
| Dura Raspatorium, Elevator stainless – A length 20 cm | Piece | √
| Dura Raspatorium, Elevator stainless – B length 18 cm | Piece | √
| Dura Raspatorium, Elevator stainless – C length 21 cm | Piece | √
| Roux retractors double-ended stainless steel, with curved blades: 3 sizes, size 1 with handle length 14 cm with one blade 22 x 20 mm and other blade 28 x 25 mm; size 2 handle length 15 cm with one blade 26 x 23 mm and other blade 38 x 29 mm; size 3 with handle length 16.5 cm and one side blade 29 x 29 mm and other side blade 44 x 34 mm. | Set | √
| Baby-mixter dissecting and ligature forceps | Piece | √
| Tangential occlusion clamps 25.5 cm | Piece | √
| Multipurpose vascular clamps with toothing 15 cm long | Piece | √
| Tunnelling instruments for bypass grafting 10 mm diameter tunnelling tube and 680 mm length with handle | Piece | √
| Intimal spatulas and dilator: Robb 20.5 cm long with tip diameter 2 mm and 4 mm flat side | Piece | √
| Vascular dilators: Debakay 19 cm long, malleable diameters 0.5 mm, 1 mm, 1.5mm, 2mm, 2.5 mm, 3mm, 3.5mm, 4mm, 4.5mm, 5mm, 7mm, 9mm. | Piece | √
| Gelpi laminectomy retractors 18.5 cm | Piece | √
| Vascular scissors-dissecting with 1 probe-pointed blade 12 cm long | Piece | √
| Hudson Brace with Cranial Burs and drills, Brace, with snap-luck chuck, length 25.5 cm stainless with following parts: | Set | √
| Bur with 12 blades, 8 cm, 15 cm width | Piece | √
| Bur with 8 blades, 8 cm, 15 cm width | Piece | √
| Bur with 8 blades, 11 cm, 15 cm width | Piece | √
| Drill 7 cm, 2 mm width | Piece | √
| Drill 8 cm, 3 mm width | Piece | √
| Perforator drill – 11 cm,15 mm width | Piece | √
| Perforator drill - 8cm,15mm width | Piece | √
| Micro Forceps - 11cm | Piece | √
| Micro needle holders, SS - Length 13 cm with catches | Piece | √
| Micro needle holders, SS - Length 13 cm without catches | Piece | √
| Micro Scissors Slightly curved, SS – 13 cm | Piece | √
| Penfield dura dissector, SS 19.5 cm length | Piece | √
| Penfield dura dissector, SS 21.5 cm length | Piece | √
| Penfield dura dissector, SS 24 cm length | Piece | √
| Penfield dura dissector, SS 29 cm length | Piece | √
| Raney Carbon steel clips | Piece | √
| Raney Disposable plastic clips | Piece | √
| Raney Scalp clip applier, SS | Piece | √
| Raney scalp clip Frame - length 19 cm, stainless | Piece | √
| Raspatorium double ended, SS - Large 21cm, 16 mm, 7 mm | Piece | √
| Raspatorium double ended, SS - Small 14 cm, 10.5 cm, 4.5 mm, | Piece | √
| Retractors, Beckman – Adson, with Hinged arms, SS, 5 prong, blunt blades length 29 cm 3 cmW - 1 pair, 5 prong, blunt blades length 24.5 cm 3 cmW, 2.8 cmH – 1 pair | Piece | √
| Rongeur Small type, Curved, SS - Length 15 cm | Piece | √
| Self retaining retractors (titanium) Double – 37 cm | Piece | √
| Self retaining retractors (titanium) Double – 42 cm | Piece | √
| Self retaining retractors (titanium) Single – 37 cm | Piece | √
| Self retaining retractors (titanium) Single – 42 cm | Piece | √
| Suction Tube, Frazlersuction Tube, Malleable – 1.5 mm diameter, 23 cm | Piece | √
| Suction Tube, Frazlersuction Tube, Malleable – 2.5 mm diameter, 23 cm | Piece | √
| Suction Tube, Frazlersuction Tube, Malleable – 3.5 mm diameter, 23 cm | Piece | √
| Suction Tube, Frazlersuction Tube, Malleable – 5 mm diameter, 23 cm | Piece | √
| Sugita dissectors silver tip, malleable L-18 cm | Piece | √
| Sugita dissectors silver tip, malleable L-23 cm | Piece | √
| Sugita dissectors silver top, slim type malleable L-18 cm | Piece | √
| Sugita dissectors silver top, slim type malleable L-23 cm | Piece | √
| Tumor forceps - Large 10 mm width 18 cm | Piece | √
| Tumor forceps - Small 7 mm width 18 cm | Piece | √
| Venticulo Peritonial (VP) Shunt System; Malleable catheter passer length 55 cm - Child | Piece | √
| Venticulo Peritonial (VP) Shunt System; Malleable catheter passer length 90 cm to 120 cm - Adult | Piece | √
| Ventricular Needle - Length 12.5 cm, 1.5 mm diameter | Piece | √
| Ventricular Needle (silver) - Length 13.5 cm | Piece | √
| Ventricular Puncture Needle - Large 3 mm length 12.5 cm | Piece | √
| Ventricular Puncture Needle - Small 2 mm length 12.5 cm | Piece | √
| Wire Saw Guide Single, SS - 5mm width 30cm | Piece | √
| Wire Saw Guide Single, SS - 6mm width 30cm | Piece | √
| Wire Saw, Olive-krona, SS - length 35cm | Piece | √
| Gigli saw handle, SS | Piece | √
| Aneurysm Clip applying forceps with locking device-Angled 30 °. | Piece | √
| Aneurysm Clip applying forceps with locking device-Angled 15 °. | Piece | √
| Aneurysm Clip removal-Angled 30° | Piece | √
| Aneurysm Clip removal straight | Piece | √
| Hydrocellular cavity dressing | Piece | √
| Hydroactive gel for slough removal 30gm Tube | piece | √
| Hydrocellular sacral wound dressing | Piece | √
| Saline solution spoits-14 cm long | Piece | √
Urology
Equipments
| Karl stroz light source 485B stand cold light, Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top and suitable Stabilizer | UE | set | √
| Telescope 30 deg. 27005 B | UE | Piece | √
| Telescope zero deg. 27005 A | UE | piece | √
| Hamou Endoment for Laproscopy and Hysteroscopy, Power supply: 220V-230V,50Hz, Note: All Power cord should be of round 3 pin plug top. suitable stabilizer with all its accessories should be provided | LE | set | √
Consumables
| 27651B cleaning brush, flexible, long for6 Fr channel | UE | Piece | √
| Ball electrode, 4Fr, length 73 cm | UE | Piece | √
| Ball electrodes,length53cm :27160AA - 3fr, 27160A -4fr, 27160B- 5fr, 27160D - 7fr ,27160E - 8fr, 27160F-10fr | UE | Piece | √
| Bard U.S.C poly urethome ureteral catheter 4fr | UE | Piece | √
| Bard U.S.C poly urethome ureteral catheter 3fr | UE | Piece | √
| Bard U.S.C poly urethome ureteral catheter 5fr | UE | Piece | √
| Cidex tray with cover (10 litres capacity) | UE | Piece | √
| Glycine solution (Iso-osmotic) 1.5% | UE | litres | √
| Hand piece for Valley Lap (Ligasure) vessel sealing system R3M859702 LS 2070 | UE | Piece | √
| Integral ureter stent (Double J stent), size 26 mm, Ch.5 | UE | Piece | √
| Integral ureter stent (Double J stent), size 26 mm, Ch.6 | UE | Piece | √
| Integral ureter stent (Double J stent), size 26 mm, Ch.8 | UE | Piece | √
| Pig tail percutanous nephrostomy catheter with needle (romsons)size:5,6,7 | UE | Piece | √
| Silicone foleys catheter, size - 14 fr | OT G | Piece | √
| Silicone foleys catheter, size - 16 fr | OT G | Piece | √
| Silicone foleys catheter, size - 18 fr | OT G | Piece | √
| Suprapubic catheterisation set disposable, catheter 10M1Cr 12Ch | UE | Piece | √
| Toomey syringe glass | UE | Piece | √
| Urostomy pouch with Colostomy WAFFER - Size 45mm |  | Piece | √
| Urostomy pouch with Colostomy WAFFER - Size 57mm |  | Piece | √
Instrunments
| 27023FE grasping forceps for small fragments single action jaw, flexible, 5Fr. Length 73 cm | UE | Piece | √
| 27023VK stone basket,5Fr. Length 60 cm consisting of: 1x27023 VT3-ring handle, 3 x 27023VR basket ,3x27023VS coil | UE | set | √
| 27023ZE biopsy forceps,single action jaw,flexible,5Fr. Length 73 cm | UE | piece | √
| 27040BK resectoscope sheat with LUER lock stopcock,including connecting tube for irrigation,24fr,short beak with obturator 270400D colour code yellow. | UE | set | √
| 27040SI resectoscope sheath including connecting tubes for in and outflow, 2 LUER-lock adaptors 26fr ,oblique back,inner tube with ceramic insulation,for use with working element 27040 L/E and 27050 C/D/E/EN colour code yellow | UE | set | √
| 27074B stone crushing forceps, single action jaw. 27074SZ adapters for use with resectoscope sheath, 2704SL/SD/SE/SM outer sheath only. | UE | set | √
| 27078AJ cystoscopic rongeur 24Fr,single action jaw with obturator 270778AO,including 2 LUER lock adapter for continous suction and irrigation | UE | set | √
| Biopsy forceps, double action jaws, flexible 5fr,length 40cm (271562H) | UE | Piece | √
| Biopsy forceps, through cutting, double action jaw for unipolar coagulation (27072DD) | UE | Piece | √
| Catheter deflecting mechanism with 1instrument channel with rachet optical(27026F) | UE | piece | √
| Cysto-urethescope sheath 24fr | UE | piece | √
| Half cut trocar for urlolgy, 5mm | UE | piece | √
| Half cut trocar for urlolgy, 6mm | UE | piece | √
| Half cut trocar for urlolgy, 8mm | UE | piece | √
| Instrument channel 6Fr for adult.Direction of view 0degree,angle of view110 degree,working length 35cmand OD;15.5Fr | UE | piece | √
| Lithotrite 27076A incorporating both handle and turning screw action 24fr including 2 LUER lock adaptors | UE | Piece | √
| Optical jawbiopsy forceps,double action jaws (27072BL) | UE | Piece | √
| SACHSE MEATOTOME with scale to set cutting sizes 10fr-36fr | UE | Piece | √
| Stone punch:27077A punch working element,27077B punch sheat with central valve,including connecting tubes for suction band irrigation25Fr, straight beak obturator270077DO | UE | set | √
| Stonebasket: basket with 6 wires,5fr (27186A), store basket 5fr,length 60cm (27023VK), ZEISS urethral stone dislodger 5fr(27190) | UE | set | √
| Urethroscope sheath size 22 fr with obturator(026B) | UE | Piece | √
| Urethroscope sheath size17fr with obturator(27026U) | UE | Piece | √
| Urethroscope sheath size19 fr with obturator(27026D) | UE | Piece | √
| Working Elements(27050EH),sets consisting of:27050E 1 working element also to use with optical urethrotome, 2 cuttuing loops (27050), 2 coagulating electrodes(27050), 2 high frequency cords (277)and 1 protection tubes(280) | UE | set | √
Accessories
| 11025E pressure compensation cap for ventilation during gas sterlization | UE | Piece | √
| 27170 connecting cord for electrodes | UE | Piece | √
| 27224 ELLIK evacuator | UE | Piece | √
| 27770E coagulating electrode 8Fr,unipolar,length53cm | UE | Piece | √
| 495NL fibre optic light cable size3.5mm, length180cm | UE | Piece | √
| Adapter for use with resectoscope sheaths (27094) | UE | Piece | √
| Argon gas | UE | litres | √
| Coagulating electrode 8fr(27770E) | UE | Piece | √
| Coagulating electrode,pointed 24fr.colour code yellow (27050L) | UE | Piece | √
| Cold knife,straight (27069K) | UE | Piece | √
| cutting loop,angled 24fr.colour code yellow(27050 G) | UE | Piece | √
| Luovik cold knife,straight waveform(27969M) | UE | Piece | √
| Telescopes bridge 27025G with one instrument channel | UE | Piece | √
Heamodialysis
| 5 micron filter compatible with R/O machine | Piece | √ | √
| AV Fistula needle 15G | Piece | √ | √
| AV Fistula needle (16G X 1) | Piece | √ | √
| Blood line with bubble catcherpart No.N-A00/V003 with filter 8mm | Set | √ | √
| Haemodialysis machine, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS | Set | √ | √
| Dialyzer Polysulfone F5, F6, F7, compatible to Dialysis Machine | Piece | √ | √
| Diasafe filter plus, compatible with Dialysis machine | Piece | √ | √
| Double Lumen Catheter (DLC) Jugular 13.5 or 14.5 cm.11.5F, with 18G introducer needle 1no, injection cap 1no | Set | √ | √
| DLC Jugular 13.5 or 14.5 cm.11F (3.7 mm) x 15 cm, with 18G introducer needle 1no, injection cap 1no | Set | √ | √
| DLC Subclavian set with wings 16cm, 11.5F (flexible guide wire 1no, Catheter with wings 1 no, 12F dilator 1no, 18G Introducer needle 1no, Injection cap 1no.) | Set | √ | √
| DLC Femoral 19.5 cm, 11.5F (flexible guide wire 1no, Catheter with wings 1 no, 12f dilator 1no, 18G Introducer needle 1no, Injection cap 1no.) | Set | √ | √
| Frekaderm spray 250 ml | Piece | √ | √
| Frekaderm spray 500ml | Piece | √ | √
| Infusion Line set (i/v set) | Set | √ | √
| Leukopore 2.5cm X 9.14m | Roll | √ | √
| Leukopore 5cm X 9.14 m | Roll | √ | √
| Micro shield hand rub 250 ml | Piece | √ | √
| Micropore 1" (2.5 cm x 9.14 m) | Roll | √ | √
| Tranpore 0.5" | Roll | √ | √
| Port cap (Dialyzer cap -rubber) | Piece | √ | √
| Reverse Osmosis (RO) Water Hardness testing kit | Piece | √ | √
| Syringe 20 ml for heparin pump (disposible) | Piece | √ | √
| Syringe 50 ml (disposable) | Piece | √ | √
| Tegaderm (6 cm X 7cm), Sterile (Transparent film dressing, frame style) | Piece | √ | √
| Tegederm 15cmx20cm, Sterile (Transparent film dressing, frame style) | Piece | √ | √
| Tran spore 1.25cm X 9.1m | Roll | √ | √
| Transducer A/V pressure monitor compatible with Dialysis machine | Piece | √ | √
| Peritoneal Dialysis transfer set, compatible with Peritoneal Dialysis machine | Set | √ | √
| Peritoneal Dialysis Catheter set (Adult), compatible with Peritoneal Dialysis machine | Set | √ | √
| Stainless Steel dressing tray with cover, size: small | Piece | √ | √
| Stainless Steel dressing tray with cover, size: medium | Piece | √ | √
| Twist lock ISO fitting [MPC-45] dialyzer blood port cap | Piece | √ | √
| Electronic weighing scale,capacity 100 kg [DIGITAL] | Piece | √ | √
| AMSD- 650 Dialyzer [Gamma sterlized] | Piece | √ | √
| Micron filter compatible | Piece | √ | √
General Consumables
General Consumables
| Catheter, Foley 2 ways (Siliconized) - size 6 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized)- size 8 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized)- size 10 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized) - size 12 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized)- size 14 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized)- size 16 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized) - size 18 | Piece | √ | √ | √ | √
| Catheter, Foley 2 ways (Siliconized)- size 20 | Piece | √ | √ | √
| Catheter, Foley 2 ways (Siliconized) - size 22 | Piece | √ | √ | √
| Catheter, Foley 3 ways (siliconized) - size 14 | Piece | √ | √
| Catheter, Foley 3 ways (Siliconized) - size 16 | Piece | √ | √
| Catheter, Foley 3 ways (Siliconized) - size 18 | Piece | √ | √
| Catheter, Foley 3 ways (Siliconozed) - size 20 | Piece | √ | √
| Catheter, Foley 3 ways (Siliconized) - size 22 | Piece | √ | √
| Catheter, Plain (Siliconized) - size 5 | Piece | √ | √ | √ | √
| Catheter, Plain (Siliconized) - size 6 | Piece | √ | √ | √ | √
| Catheter, Plain (Siliconized) - size 8 | Piece | √ | √ | √ | √
| Catheter, Plain (Siliconized) - size 10 | Piece | √ | √ | √ | √
| Catheter, Plain (Siliconized) - size 12 | Piece | √ | √ | √ | √
| Catheter, Plain (Siliconized) - size 14 | Piece | √ | √ | √ | √
| Catheter, Plain (Siliconized) - size 16 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line - size 6 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line - size 8 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line- size 10 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line - size 12 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line- size 14 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line- size 16 | Piece | √ | √ | √ | √
| Catheter, suction (PVC), with X-Ray opaque line- size 18 | Piece | √ | √ | √ | √
| Chest drainage tube - size 10 | Piece | √ | √ | √
| Chest drainage tube - size 12 | Piece | √ | √ | √
| Chest drainage tube - size 14 | Piece | √ | √ | √
| Chest drainage tube - size 16 | Piece | √ | √ | √
| Chest drainage tube - size 18 | Piece | √ | √ | √
| Chest drainage tube - size 20 | Piece | √ | √ | √
| Chest drainage tube - size 22 | Piece | √ | √ | √
| Chest drainage tube - size 24 | Piece | √ | √ | √
| Chest drainage tube - size 26 | Piece | √ | √ | √
| Chest drainage tube - size 28 | Piece | √ | √ | √
| Chest drainage tube - size 30 | Piece | √ | √ | √
| Chest drainage tube - size 32 | Piece | √ | √ | √
| Chest drainage bag water seal, 1000 ml | Piece | √ | √ | √
| Colostomy bag (transparent) with WAFFER (Flange) tail closer (clip), perforated marking size of 25mm - 64mm | Set | √ | √ | √
| Condom catheter (male external device) - Small | Piece | √ | √ | √ | √
| Condom catheter (male external device)- Medium | Piece | √ | √ | √ | √
| Condom catheter (male external device) - Large | Piece | √ | √ | √ | √
| Disposable Wooden Tongue Depressor with paper cover | Piece | √ | √ | √ | √
| Disposable Sharp Container | Piece | √ | √ | √ | √
| Disposable Face mask, three ply with back-tie | Piece | √ | √ | √ | √
| ECG paper 50mmx20 metre for 6108T machine | Roll | √ | √ | √
| ECG Paper Schiler AT/AT-2/Plus | Packet | √ | √
| Gel, high conductivity for ECG | Piece | √ | √ | √
| AQUA Gel for USG, CTG, fetal Doppler | Piece | √ | √ | √ | √
| Adhesive Cardiac Electrode pad (disposable) | Piece | √ | √ | √
| Gauze Than 36"x20 yds | Packet | √ | √ | √ | √
| Disposable Gloves non-sterile - small | Pair | √ | √ | √ | √
| Disposable Gloves non-sterile - medium | Pair | √ | √ | √ | √
| Disposable Gloves non-sterile - large | Pair | √ | √ | √ | √
| Gloves sterile, latex with plastic wrap - size 6.5 | Pair | √ | √ | √ | √
| Gloves sterile, latex with plastic wrap - size 7 | Pair | √ | √ | √ | √
| Gloves sterile, latex with plastic wrap - size 7.5 | Pair | √ | √ | √ | √
| Gloves sterile, latex with plastic wrap - size 8 | Pair | √ | √ | √ | √
| Gloves sterile, latex, elbow-length, size: 6.5 | Pair | √ | √
| Gloves sterile, latex, elbow-length, size: 7 | Pair | √ | √
| Gloves sterile, latex, elbow-length, size: 7.5 | Pair | √ | √
| Gloves sterile, latex, elbow-length, size: 8 | Pair | √ | √
| Gluco strip compatible with Glucometer | 50strips/ pkt | √ | √ | √ | √
| IV infusion dripSet, Specification: 3mm non-toxic PVC tubing with big PVC moulded chamber with 15mm spike cover filter with airvent spike | Piece | √ | √ | √ | √
| Stop Cock, 3 ways connector for IV drips. | Piece | √ | √ | √
| IV Cannula 2 ways, size 14 with injecting port | Piece | √ | √ | √ | √
| IV Cannula 2 ways, size 16 with injecting port | Piece | √ | √ | √ | √
| IV Cannula 2 ways, size 18 with injecting port | Piece | √ | √ | √ | √
| IV Cannula 2 ways, size 20 with injecting port | Piece | √ | √ | √ | √
| IV Cannula 2 ways, size 22 with injecting port | Piece | √ | √ | √ | √
| IV Cannula 2 ways, size 24 with injecting port | Piece | √ | √ | √ | √
| IV Cannula 2 ways, size 26 with injecting port | Piece | √ | √ | √
| IV Cannula 2 ways, size 24 with injecting port without wings | Piece | √ | √
| IV Cannula 2 ways, size 26 with injecting port without wings | Piece | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned without side opening at the tip, Length 52cm - size 4 FG | Piece | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip, distal end coned without side opening at the tip, Length 52cm - size FG 5 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned without side opening at the tip, Length 52cm - size FG 6 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes, Length 52cm - size FG 8 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes, Length 52cm - size FG 10 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes - size FG 12 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes- size FG 14 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes - size 16 | Piece | √ | √ | √ | √
| Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes- size FG 18 | Piece | √ | √ | √ | √
| Rectal Tube (Flatus), size 6 | Piece | √ | √ | √ | √
| Rectal Tube (Flatus), size 8 | Piece | √ | √ | √ | √
| Rectal Tube (Flatus), size 10 | Piece | √ | √ | √ | √
| Rectal Tube (Flatus), size 12 | Piece | √ | √ | √
| Rectal Tube (Flatus), size 14 | Piece | √ | √ | √
| Rectal Tube (Flatus), size 16 | Piece | √ | √ | √
| Rectal Tube (Flatus), size 22 | Piece | √ | √ | √
| Rubber (Mackintosh) roll of 25 meter | 25mtr/ roll | √ | √ | √ | √
| Condom | Gross | √ | √ | √ | √
| Urine Drainage Bag with T-type, bottom outlet and sampling port, capacity:1000/2000ml | Piece | √ | √ | √ | √
GC- Dressing
| Adhesive plaster 7.5cm x 5mts | Roll | √ | √ | √ | √
| Bandage rolled 5cm x 3mtr Length | Pkt of 12 | √ | √ | √ | √
| Bandage rolled 10cm x 3mtr Length | Pkt of 12 | √ | √ | √ | √
| Bandage rolled 15cm x 3mtr Length | Pkt of 12 | √ | √ | √ | √
| Cotton (zig-zag) 400g | Packet | √ | √ | √ | √
| Elastic Adhesive Bandage 10cm x 6mts (stretch length), flesh colour with fast edges | Packet | √ | √ | √
| Adhesive surgical tape (paper) 25mm x 9.1 mts | Roll | √ | √ | √
| Adhesive surgical tape (paper) 50mm x 9.1 mts | Roll | √ | √
| Adhesive surgical tape (paper) 75mm x 9.1 mts | Roll | √ | √
| Transparent dressing (Tegaderm) :15 cm x 20 cm/6 inch x 8 inch | Piece | √ | √
| Transparent dressing (Tegaderm) :6 cm x 7cm/ 2 3/8inch x 2 3/4inch. | Piece | √ | √
| Transparent dressing (Tegaderm) :4.3 cm x 4.3cm/ 2 3/8inch x 2 3/4inch. | Piece | √ | √
GC- Linen
| Hand towel 12 inch x 12 inch - white | Piece | √ | √ | √ | √
| Patient Trolley matttress, foam with water proof synthetic cover 72 inch x 21 inch x 2 inch with strips at both ends | Piece | √ | √ | √
| Mattress, coir with water proof synthetic cover : size length 78" x Width 37" x Breadth/thickness 4" | Piece | √ | √ | √ | √
| water proof synthetic mattress cover : size length 78" x Width 37" x Breadth/thickness 4" | Piece | √ | √ | √ | √
| Pillow cover (cotton): size 25 inch x 17 inch - Sky Blue | Piece | √ | √ | √ | √
| Synthetic foam pillow, washable: size 24 inch x 16 inch, 1.5 kg | Piece | √ | √ | √ | √
| Bed sheet with elastic band at four corners Size: length 78" x Width 37" x Breadth/thickness 4" | Piece | √ | √ | √ | √
| Wollen Blanket: Green colour. Weight-2kg Size: 60 inches × 90 inches.composition 80% wool and 20% other fibre | Piece | √ | √ | √ | √
| Draw sheet: size 72 inch x 36 inch - Green Colour | Piece | √ | √ | √
GC-Utility
| Apron: plastic with cloth piping, back and neck tie, standard size | Piece | √ | √ | √ | √
| Apron plastic with cloth piping with back and neck tie. Full sleeves with elastic band at wrist. Size small | Piece | √ | √
| Apron plastic with cloth piping with back and neck tie. Full sleeves with elastic band at the wrist. Size Medium | Piece | √ | √
| Apron plastic with cloth piping with back and neck tie. Full sleeves with elastic band at wrist. Size Large | Piece | √ | √
| Polypropylene Bio-hazard bag, Autoclavable and Bio-degradable with biohazard symbol, size: Length 87cm x Breadth 75cmltr (Red colour). Thickness not less than 9micron. | Roll | √ | √ | √ | √
| Polypropylene, Bio-degradable: size Length 87cm x breadth 75cm (Green colour). Thickness not less than 9micron. | Roll | √ | √ | √ | √
| Baby liquid medicine dispencer. Maximum capacity 1 tsp/1ml. Simple | Piece | √ | √
| Digital forehead (Sensor) medical thermometer: Forehead Thermometer with sensor; High accuracy; 6-8 seconds Measurement; Memory function; LCD Display; Fever alarm; Batteries included | Piece | √ | √ | √ | √
| Digital Thermometer with battery |  | √ | √ | √ | √
| Safety goggles: With high impact resistance, panoramic and crystal clear view. Glare reducer. Anti scratch properties. | Piece | √ | √ | √
| Face mask N95: Comfortable, Lightweight. Particulate respirator with 95% filtration efficiency. Contoured, aluminium nose Piece with foam | Piece | √ | √ | √
| Bone marrow aspiration Needle, sterile - size 16G | Set | √ | √
| Bone marrow aspiration Needle, sterile - size 18G | Set | √ | √
| Bone marrow aspiration Needle, sterile - size 20 G | Set | √ | √
| Paraffin gauze dressing, 10 x 10 cm, sterile | Pkt of 10 dressings | √ | √ | √
| Paraffin gauze dressing, 15 x 15 cm, sterile | Pkt of 10 dressings | √ | √ | √
| Paraffin gauze dressing, 20 x 20 cm, sterile | Pkt of 10 dressings | √ | √ | √
| Paraffin gauze dressing, 30 x 30 cm, sterile | Pkt of 10 dressings | √ | √ | √
| Tourniquet | Piece | √ | √ | √
| Wrist band for patient- Paediatric and neonate: Precision ID Bracelets, Insert Style. | Piece | √ | √
| Wrist band for patient- Adult: Precision ID Bracelets, Insert Style. | Piece | √ | √
| Needle 16G (Disposable) | Piece | √ | √ | √ | √
| Needle 18G (Disposable) | Piece | √ | √ | √ | √
| Needle 20G (Disposable) | Piece | √ | √ | √ | √
| Needle 22G (Disposable) | Piece | √ | √ | √ | √
| Needle 24G (Disposable) | Piece | √ | √ | √ | √
| Syringe with needle 1ml (Disposable) | Piece | √ | √ | √ | √
| Syringe with needle 2ml (Disposable) | Piece | √ | √ | √ | √
| Syringe with needle 5ml (Disposable) | Piece | √ | √ | √ | √
| Syringe 10ml with 21G needle (Disposable) | Piece | √ | √ | √ | √
| Syringe without needle 20ml (Disposable) | Piece | √ | √ | √ | √
| Syringe without needle 50ml (Disposable) | Piece | √ | √ | √ | √
| Cover all: Large | Packet | √ | √
| Cover all: Medium | Packet | √ | √
| Cover all: Small | Packet | √ | √
| Face shield | Piece | √ | √
| Piles binding rubber bands | Piece | √ | √
General Equipment
| Air mattress: Electric, 220V, 50Hz with pump, PVC 3ply, thickness 0.3-0.4mm. | Set | √ | √ | √
| Ampoule cutter: Anodised Aluminum.Silicone (O-ring). Stainless steel spring, Compatible ampoule sizes* 1-15ml . Slim light weight design, pocket size | Piece | √ | √ | √ | √
| Bed pan plastic: Type plastic be perfection with lid. Length 35.0cm, width 28.5cm, front height 6.0cm, rear height 8.5cm | Piece | √ | √ | √ | √
| Bed Fowler's with wheel -Semi/half: Size, 206cm (L) X90cm (W) x 60 cm (H) function: one function movements and controls. Back lift manual, Top, two sections. Tubular head and foot bows of unequal height. Frame: mild steel, finish with epoxy power coated, stainless steel panels with provisions for IV rod in two sections & wheels | Piece | √ | √ | √ | √
| Bed Fowler's with wheel -Full: Overall size: 78" L x36"W x 24" H ( 206L x 90W x60 H cms ) Back rest and knee rest adjustable maneuvered by 2 separate screw handle from foot end. Four section perforated CRCA MS sheet top. Provisions for stainless steel IV rod with four locations. Removable stainless steel tubular head & foot bows with laminated panels. Bed mounted on 100mm Dia castor two with brakes & wheels. Pre-teated epoxy powder coated. | Piece | √ | √ | √
| Bed Fowler's with wheel- Full : Orthpedic patient bed with balken beam and pulleys with adjustable height | Piece | √ | √
| Bedside Locker: Full Body SS (Stainless Steel) size: 16"x 16" x 32" HT 16" x 16" x 32" HT Body machine pressed from 20 SWG M.S. sheet with three sides closed. Have one box and one drawer with SS.Mounted on 5 cm dia. swinging castors or legs on rubber shoes. Size: 46 x 46 x 81 cm. | Piece | √ | √ | √ | √
| Bedside Screen 4 fold: with spring and cloth size 168 cm H x 244 cm W. Frame work made of ERW tubular pipe in four folds Piece fold 60cms dia and two inner folds fitted with one castor Piece of 5cms dia, Supplied with spring wire and curtains, pretreated and epoxy power coated. | Piece | √ | √ | √ | √
| Burns cradle | Piece | √ | √ | √
| Cardiac Table:Rectangular Frame Work Mounted On 5cms Castors30×16 Laminated Top Adjustable By Side Knob. Pre-treated and epoxy powder coated. | Piece | √ | √ | √
| Cheatle forceps : Size, 10 inch, Stainless Steel | Piece | √ | √ | √ | √
| Cheatle forcep stand: Stainless steel | Piece | √ | √ | √ | √
| Trolley for clean linen: Polymer design, Dimension 724mmDx219mmLx1736mmH, Capacity of 1.36 cubic metres | Piece | √
| Trolley for dirty linen : Box type, Dimension: 1060x610x610mm, Stainless steel frame work with push handle at one end, mounted on 5cms castors with wheels. Body: Superior quality plastic (ABC material) | Piece | √ | √
| Dressing Drum: Size, 12 inch, Stainless Steel: Domed top and bottom with sliding removable band and drop down-hasps (stainless steel). Made of stainless steel grade 316, thickness 0.5mm. | Piece | √ | √ | √ | √
| Dressing Drum: Size, 15 inch, Stainless Steel:Domed top and bottom with sliding removable band and drop down-hasps (stainless steel). Made of stainless steel grade 316, thickness 0.5mm. | Piece | √ | √ | √ | √
| Dressing Drum: Size, 9 inch, Stainless Steel: Domed top and bottom with sliding removable band and drop down-hasps (stainless steel). Made of stainless steel grade 316, thickness 0.5mm. | Piece | √ | √ | √ | √
| Digital weighing Scales (bathroom) 120 kg | Piece | √ | √ | √ | √
| Electronic Weighing scale: LCD Display, with tray, capacity 25 kg. Tolerance +-0.5 gm, Precision minimum 10gm, Power supply: 220V-230V, Power cord of round 3 pin type plug top | Piece | √ | √ | √
| ECG Machine (Portable) : Single Channel ECG Recording, 12 leads, Last memorized ECG Recording and additional copies available, Automatic and Manual Recording Modes. Dual Power supply: Mains and Battery operation. Power supply 220-230V. Power cord with round 3 pin top . Built in recharge | Set | √ | √ | √
| ECG machine( Table Model) : Dimensions: 400x330x100 mm, Weight: 5.0 kg (5.35kg with full paper tray). Mains Supply: 100 to 115/220 to 240 VAC, 50/60 Hz. Battery: Built-in 12 V lead-acid battery (rechargeable). Battery capacity: 4 hours normal use- 300 printouts. Power Consumption: Recording 40 VA max. Leads: Standard/Cabrera. Paper speed: 5/10/25/50 mm/s (direct). Sensitivity: 5/10/20 mm/mV, either automatically adjusted or manually selected. Chart Paper: Thermoreactive – Z-folded, 210 mm wide, perforation 280 mm. Printing Process: High-resolution thermal print head, 8 dots per mm / 200 dots per inch (amplitude axis). 40 dots per mm / 1000 dots per inch (time axis 25 mm/s). Recording Tracks: 6 channels, positioned at optimal width on 200 mm, automatic baseline adjustment. Automatic Lead Programs: Data Record: Printout of all 12 leads, Listing of ECG recording data. Version C: ECG measurement results (intervals, amplitudes, electrical axis), Sokolow Index, average complexes with optional measurement reference markings, and interpretation. | Set | √ | √
| ECG Storage: Circular input memory for 10 s, 12-lead ECG. Memory Option: Memory for c.45 ECG recordings with transmission facilities over an RS-232 interface. Frequency Range of Digital Recorder: 0 to 150 Hz (IEC), 0 to 150 Hz (AHA). ECG Amplifier: Simultaneous, synchronous registration of all 9 active electrode signals (=12 standard leads). Sampling frequency: 5µV. Dynamic range: ±9.5 VAC. Max. electrode potential: ±300 mVDC. Time constant: 3.2 s. Frequency response: 0.05 to 150 Hz (-3 dB). Input impedance: >2.5MOhms at 10 Hz. Myogram Filter (muscle tremor filter): 25 Hz or 35 Hz, programmable (not active on average waveform). The stored ECGs can be printed with or without filter. Distortion-free suppression of superimposed 50 or 60 Hz sinusoidal interferences by means of an adaptive digital filter. Line Frequency Filter: Inpatient Input: Fully floating and isolated, defibrillation protected. | set | √ | √
| Safety Standard: CF according to IEC and complying with the following. RL 93/42/EEC. EN 60601-1:1990. IEC 601-1. IEC 601-2-25:1993. pr EN 1441:1994. CISPR 11: 1985, EN 55011:1992. IEC 801-2:1991, IEC 801-3:1984, IEC 801-4:1988. IEC 801-5. Safety Class: I according to IEC 601-1 (with internal power supply). IIa according to RL 93/42/EEC, CE-0123. This device is not designed for outdoor use (IP 20). Environmental Conditions: Temperature, Operating: 10° to 40°C. Temperature, Storage: -10° to 50°C. Relative humidity: 25 to 95% (non condensing). Atmospheric pressure: 700 to 1060 hpa.Power supply 220-230V. Power cord with round 3 pin top. Control Panel: Rubber keys. | set | √ | √
| Enema can SS with marking: Measuring: 50ml - 1000 ml with 3 extra tubing | Set | √ | √ | √ | √
| Examination table with mattress and Foot step: Stainless Steel with rubber topping | Set | √ | √ | √ | √
| Pulse Oximeter, Finger Tip | Piece | √ | √ | √ | √
| Forceps Dressing: Thumb ( Non-toothed): Size 8 inch | Piece | √ | √ | √ | √
| Forceps Dressing: Gross Maier : Size 8 inch | Piece | √ | √ | √ | √
| Galli pot: 100ml, Stainless Steel | Piece | √ | √ | √ | √
| Gauze cutting machine: features XD-100 can cut from a single ply of fabric, right up to 1" high, making it a very versatile - lightweight machine. Octagonal cutting blade.Built-in sharpening mechanism .Extra-thin base plate.High-power electric motor .Carbide-tipped counter blade | Piece | √ | √
| IV stand with wheels: Stainless steel adjustable with four hooks/hanger | Piece | √ | √ | √ | √
| Jug: Plastic, graduated, 1 Ltr | Piece | √ | √ | √ | √
| Kick bucket: Foot step with wheel, Stainless Steel | Piece | √ | √ | √ | √
| Kidney dish: Size: 10 inch, Stainless Steel | Piece | √ | √ | √ | √
| Kidney dish: Size: 8 inch, Stainless Steel | Piece | √ | √ | √ | √
| Knee Hammer with plastic handle | Piece | √ | √ | √ | √
| Medicine cup - graduated 30 ml, plastic | Piece | √ | √ | √ | √
| Medicine Trolley with four drawers (SS): Overall approx. Sizes : 76Lx46Wx8H cms, Frame work of MS Tube mounted on 10cms castors,Two SS shelves with three side railing on top shelf, Two drawers under the each shelf, Whole body frame SS (Stainless steel) | Piece | √ | √ | √
Emergency crash trolley
| Nebulizer machine: Nebulizer compressor system, Spare nebulizer chambers with mask, Power supply 220-230V. Power cord with round 3 pin top | Piece | √ | √ | √ | √
| Nebulizer chamber along with tubings and mask, Adult | Piece | √ | √ | √ | √
| Non electric weighing scale with tray – capacity upto 25 kgs, tolerance ± 1 gm, Precision minimum 10 gm. | Piece |  |  | √ | √
| Oxygen trolley: B-type | Piece | √ | √ | √ | √
| Oxygen trolley: D-type | Piece | √ | √ | √ | √
| Patient chart holder | Piece | √ | √ | √ | √
| Patient transport stretcher trolley: 210Lx65Wx65Hcms, with safety side railings, IV rods, Cylinder cage, Adjustable height by hydraulic foot operated pump, 15cm dia castors, pretreated & epoxy powder coated | Piece | √ | √ | √
| Proctoscope: For adult, Stainless Steel | Piece | √ | √ | √
| Proctoscope: For Pediatric, Stainless Steel | Piece | √ | √ | √
| Piles banding forcep | Piece | √ | √
| Scissor: For slough cutting, Stainless Steel | Piece | √ | √ | √
| Scissor: Tailor, 9 inch | Piece | √ | √ | √ | √
| Scissor: Bandage, 7 inch SS | Piece | √ | √ | √ | √
| Sphygmomanometer : Dial type Single-hand operation Non-stop pin reliable manometer. Big dial face 65mm & easy to read, Latex free cuff | Set | √ | √ | √ | √
| Spotlight on stand, fibre optic | Piece | √ | √ | √ | √
| Sputum cups with lid: Size: 100ml, Stainless Steel | Piece | √ | √ | √ | √
| Stethoscope bivalve | Piece | √ | √ | √ | √
| Stretcher: Hand carrying, aluminum pole | Piece | √ | √ | √ | √
| Suction machine: Portable, Stainless steel body with wheels, high pressure tubes, pressure regulating switch. twin suction bottle of 2-3 lits. Capacity, display of vacuum from 100 to 700 mm Hg. Power supply AC 220-230V. Power cord with round 3 pin top | Piece | √ | √ | √ | √
| Suction machine (foot operated) | Piece | √ | √ | √ | √
| Suture removing scissor: Size: 8 inch | Piece | √ | √ | √ | √
| Swab container with cover: Size, 4 inch, Stainless Steel | Piece | √ | √ | √ | √
| Tray with cover: Size, 10" x 12", Stainless Steel | Piece | √ | √ | √ | √
| Tray with cover: Size, 12" x 18", Stainless Steel | Piece | √ | √ | √ | √
| Tray with cover: Size, 9" x 5", Stainless Steel | Piece | √ | √ | √ | √
| Trolley, dressing: Stainless Steel: tubular pipe framework. Mounted on 100mm diameter castor wheels, two with brakes. Two SS shelves with railings. Fitted with SS basin and pail. Dimensions (LXWXH): 910X460X800mm | Set | √ | √ | √
| Urinal plastic: With marking from 50 - 1000ml | Piece | √ | √ | √ | √
| Vial Opener: Stainless Steel | Piece | √ | √ | √ | √
| Weighing scale stand model: For weight & height, 120 kg (Sliding Lever mechanism) | Piece | √ | √ | √ | √
| Wheel Chair with commode | Piece | √ | √
| Oxygen Concentrator : Small, quiet and light weight, with oxygen sensing device. Oxygen purity 85 - 99% at all flow rates, Flow rate : 0.5 - 5 LPM, Sound level: low, Electrical requirement: Power 240V,50Hz, Accessories: Ozygen mask with tubing for adult, nasal soft-tip cannula with tubing for pediatric and adult patients, oxygen bubble humidifier, warranty 3 - 5 years. Power supply: 220-240V. Power cord with round 3 pin top | set | √ | √
| Needle Cutter (Mini Destructor for needles and syringes): Balcan | piece | √ | √ | √ | √
Sutures
General Sutures
| Catgut Plain 2-0 NW 3/8 circle round bodied 30mm, 76 cm | 12Pcs/ Pkt | √ | √ | √ | √
| Chromic Catgut 1, 1/2 circle round bodied, 40mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Chromic Catgut 2-0, 1/2 circle round bodied 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Chromic Catgut 3-0, 1/2 circle reverse cutting 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Chromic Catgut 4-0, 3/8 circle round bodied 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
| Chromic Catgut 5-0, 3/8 circle round body 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
| Nylon 1-0, reverse cutting 45 mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Nylon 2-0, 3/8 circle reverse cutting 45mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Nylon 3-0, 3/8 circle reverse cutting 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Nylon 4-0, 3/8 circle cutting 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
| Nylon 6-0 reverse cutting 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
| Prolene 1,1/2 circle reverse cutting Heavy 100cm | 12Pcs/ Pkt | √ | √ | √
| Prolene 2-0, 1/2 circle round body 30mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Prolene 3-0, 1/2 circle round body 25mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Prolene 3-0, reverse cutting 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Prolene 4-0, 1/2 circle round body 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Prolene 6-0 round bodied 8mm, 60cm double needled | 12Pcs/ Pkt | √ | √
| Prolene 7-0, 8mm round bodied 60cm double needle | 12Pcs/ Pkt | √ | √
| Silk 2-0, 3/8 circle reverse cutting 45mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Silk 2-0, 1/2 circle round bodied 30mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Silk 3-0, 3/8 circle reverse cutting 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
| Silk 3-0, 1/2 circle round body 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Silk 4-0, 3/8 circle round bodied 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Silk sutupak 2-0 (120 cm) | 12Pcs/ Pkt | √ | √
| VICRYL 1, 1/2 circle round bodied 40mm needle heavy, 120 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Vicry 1-0, 1/2 circle round bodied 40mm needle, 90 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Vicryl 2-0, 1/2 circle round bodied 30mm, 90 cm suture length | 12Pcs/ Pkt | √ | √
| Vicryl 2-0 reverse cutting, 45mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Vicryl 3-0 reverse cutting, 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Vicryl 3-0, 1/2 circle cutting, 22mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
| Vicryl 4-0, 1/2 circle round bodied needle 20mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
| Vicryl 5-0, 1/2 circle round bodied 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
| Vicryl 3-0, 1/2 circle round bodied, 20mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √
Eye Sutures
| Nylon 10-0 | 12Pcs/ Pkt | √ | √
| Silk 5-0, 3/8 circle round bodied 16mm | 12Pcs/ Pkt | √ | √
| Silk 6-0, 1/4 circle black braided 8mm | 12Pcs/ Pkt | √ | √
| Silk 6-0, 3/8 micropoint spatulated 6mm double | 12Pcs/ Pkt | √ | √
| Vicryl 5-0 on ap2 needle | 12Pcs/ Pkt | √
| Vicryl 6-0, micropoint 1/4 circle spatulated | 12Pcs/ Pkt | √ | √
| Vicryl 8-0, micropoint 3/8 circle spatulated | 12Pcs/ Pkt | √ | √
| Silk 8-0, micropoint 3/8 circle spatulated | 12Pcs/ Pkt | √ | √
| Orthopedic Surgery |  | √ | √ | √ | √
| Ethibond 1-0 reverse cutting 45mm | 12Pcs/ Pkt | √ | √ | √
| Ethibond 2 taper cut 45mm | 12Pcs/ Pkt | √ | √ | √
| Ethibond 2-0 tapercut 45mm | 12Pcs/ Pkt | √ | √ | √
| Ethibond 5 tapercut 50mm | 12Pcs/ Pkt | √ | √ | √
| Braided Polyblend suture (fibre wire) size 5, 1/2 circle, 45mm, 96.5 cm | 12Pcs/ Pkt | √ | √
| Braided Polyblend suture (fibre wire) size 2, 1/2 circle, 45mm, 96.5 cm | 12Pcs/ Pkt | √ | √
| Braided Polyblend suture (fibre wire) size 2.0, 1/2 circle, 30mm, 96.5 cm | 12Pcs/ Pkt | √ | √
| Braided Polyblend suture (fibre wire) size 2.0, 1/2 circle, 40mm 75-90 cm | 12Pcs/ Pkt | √ | √
| Braided Polyblend suture (fibre wire) size 4.0, 1/2 circle, 20mm, 70cm | 12Pcs/ Pkt | √ | √
| Nylon 4-0 1/2 circle reverse cutting, 20 mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √
| PDS (PolyDioxanone), size: 1, double strand, 45 mm with length of 90 cm | 12Pcs/ Pkt | √ | √
| Surgical Department |  | √ | √ | √ | √
| PDS (PolyDioxanone), size: 1, double strand, 45 mm with length of 90 cm | 12Pcs/Pkt | √ | √
| PDS (PolyDioxanone), size: 1.0, 1/2 circle round body 25mm, 76 cm suture length | 12Pcs/Pkt | √ | √
| PDS (PolyDioxanone), size: 2/0, 1/2 circle round bodied 30mm, 90 cm suture length | 12Pcs/Pkt | √ | √
| PDS (PolyDioxanone), size: 3/0, 1/2 circle round bodied, 20mm needle, 76 cm suture length | 12Pcs/Pkt | √ | √
| PDS (PolyDioxanone), size: 4/0,1/2 circle round bodied needle 20mm, 76 cm suture length | 12Pcs/Pkt | √ | √
| PDS (PolyDioxanone), size: 5/0 , 1/2 circle round bodied 16mm, 76 cm suture length | 12Pcs/Pkt | √ | √
| Monocryl 2-0, 3/8 circle round bodied 30mm, 76 cm | 12Pcs/Pkt | √ | √
| Pediatric Department |  | √ | √ | √ | √
| PDS (PolyDioxanone), size: 4/0,3/8 circle cutting needle 19mm, 76 cm suture length | 12Pcs/Pkt | √
| Vicryl 4-0, 3/8 circle cutting needle 19mm, 76 cm suture length | 12Pcs/Pkt | √
| Vicryl 5-0, 3/8 circle cutting 13mm, 76 cm suture length | 12Pcs/Pkt | √
| Dental Department |  | √ | √ | √ | √
| Plain catgut 6.0 (Fast absorbing), PC-1 13mm Needle, 3/8 circle, 45 cm suture length | 12Pcs/Pkt | √
| Plain catgut 5.0 (Fast absorbing), PC-1 13mm Needle, 3/8 circle, 45 cm suture length | 12Pcs/Pkt | √
| PDS (PolyDioxanone), size: 4/0,3/8 circle reverse cutting needle 19mm, 76 cm suture length | 12Pcs/Pkt | √
| Monocryl 6.0, 3/8 circle, 11mm needle, Reverse cutting 45 cm suture length | 12Pcs/Pkt | √
Royal Centre for Disease Control
National Tuberculosis Reference Laboratory
| Biological Safety Cabinet: Class II, Type B1, NU_427-400E NuAIRE Bio-Safety Cabinet- Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
| Biological Safety Cabinet: Class II, Type B2, NU-427-400E NuAIRE | Piece
| Autoclave: ES-315 TOMY ES-315 TOMY High pressure steam sterilizer ES-315 use for sterilization and decontamination of laboratory equipments and waste. The vertical autoclave with a top-open lid, which can be operated easily with a single hand or foot. Operating temperature range (during sterilizing): 105 - 132°C, Operating pressure range: 0 - 186kPa, Maximum operating pressure 216kPa, Temperature control: Digital, microprocessor controlled, Temperature display / display range: Digital / -15 - 180°C, Pressure display / display range: Analog / 0 - 400kPa, Heat source: 2.0kW electric heater, Timer control: Digital, microprocessor controlled, Timer display: Digital. Safety devices: • Inside the chamber overheat prevention, • Outer wall of the chamber overheat prevention, • Overpressure prevention, • Temperature sensor disconnection prevention, • Empty heating prevention, • Leakage breaker, • Safety valve. Malfunction prevention devices: • Low water level detection, • Exhaust valve knob open/close detection, • Insufficient sterilization detection, • Chamber lid open/close detection. Chamber material: SUS304. Capacity / material of exhasut bottle: 3ℓ / polyethylene. Type / material of lid gasket: Accommodating internal pressure / silicone rubber. Dimensions of main unit (mm): 490W x 560D x 1090H (Height from floor to control panel:875). Net weight: 80kg. Rated Voltage AC 120 / 220 / 230 / 240V. Power input: 20A (120V), 10A (220V), 9A (230 / 240V), Power supply requirements: Single-phase 220 / 230 / 240V AC (50 / 60Hz) 15A or above. Power Comsumption (calorific value): 2kW (1720kcal/h). Accessories: • Stainless steel basket with slatted bottom plate 1, • Water level sensor 1, • Inspection sheet 1, • Warranty card 1, • Customer card 1, • Operator's manual 1,• Clear plastic accessory case 1, • Attachment screw for clear plastic accessory case 1. Stainless basket (φ300 x 182mm) | Piece
| Autoclave: Ospery/ J8153 | Piece
| Refrigerated Centrifuge: Zentrifuge-Hittich, Rotana 460 R | Piece
| Refrigerated Centrifuge: Zentrifuge-Hittich, Rotana 460 R | Piece
| Micro-Centrifuge: MIKRO-200 Hettich, Germany | Piece
| TwinCubator: Hain Lifescience, Germany | Piece
| Thermal Cycler: GTQ Cycler 96,Hain Lifescience, Germany | Piece
| Water Bath: GFL, Type-1002, No. 11264413J | Piece
| Water & Oil Bath, MEMMERT | Piece
| Weighing Balance, CA-224 contech Instruments ltd CA-224 | Piece
| Weighing Balance, KERN PLS-310-3F, Germany | Piece
| Weighing Balance, UMT2 Mettler UMT2 | Piece
| Ph Meter, lab 850 INTERTEK lab 850 | Piece
| Light Microscope, CH20I OLYMPUS CH20I | Piece
| Liquid culture system, MGIT 960 Becton Dickinson MGIT 960 | Piece
| Isspissator (Open System), VFIN GRANT INSTRUMENT LTED VFIN | Piece
| Isspissator (Open System): TB-GD100IN GRANT INSTRUMENT LTED TB-GD100IN | Piece
| Isspissator (Closed System): J-IWL JESICO J-IWL | Piece
| Refrigerator: 2-8°C, Samsung | Piece
| Refrigerator, 2-8°C, Vision Scientific, VS 1302L | Piece
| Ultra Low Temperature Refrigerator: MDF-433V-PE PANASONIC MDF-433V-PE | Piece
| Freezer: Electrolux, TCW 1151 | Piece
| Universal Centrifuge: Z323K Hermele Labortechnik Z323K | Piece
| Shaker: KS501D10.010381 IKA-WERKE KS501D10.010381 | Piece
| Hot plate: HT1 BIBBY STERILIN LTD HT1 | Piece
| Incubator: MEMMERT Model 600 D6062 | Piece
| Incubator: MEMMERT INE600 | Piece
| Incubator: MEMMERT INE-800 | Piece
| Incubator: MEMMERT INE-800 | Piece
| Incubator: MEMMERT | Piece
| Hot Air Oven: UFB400 MEMMERT UFB400 | Piece
| Slide drying bench/Slide Warmer: with black anodised working plate, temp. Up to 100°C, regulation in 1°C steps, digital display, overheating protection, size WxDxH 160x290x30mm: with black anodised working plate, temp. Up to 100°C, regulation in 1°C steps, digital display, overheating protection, size WxDxH 160x290x30mm | Piece
| Safety Gas Burner & Safety gas tubing: 143000 INTEGRA BIOSCIENCES AG 143000: Safety gas burner, with foot switch, IR sensor, with rest heat display and splash guard. can be operated with touch free IR-Sensor, button function or foot pedal. With the DualKnob for adjusting gas and air flow, the flame can be precisely regulated. can be operated with stationary natural gas and propane / butane gas supplies, cartridge gas or gas cylinders. Safety Gas Tubing: safety gas tube 750mm long, with clamps. Inner diameter 9mm | Piece
| Glass brushing machine: Electrically operated, with a quick release chuck taking a wide variety of interchangeable nylon brushes for cleaning syringes, test tube, McCartney bottles, and round bottles up to 250ml. Mounted on a cast aluminium base with a perspex anti-splash shield, a white enamelled tray and incorporating a 1/16hp motor for 250V/50Hz single phase supplies. Dimensions: (approx.) 62.5x27cmsx25cms.; Approx.wt.:50-60 kgs., Accessories: Brushes, nylon, for round bottles capacity 30,120,250ml; brushes, nylon, for McCartney bottles- diameter 25mm. | Piece
| Vortex mixture: Vortexer is designed with a solid-state electronic control that provides adjustable speed from 40 to 3200rpm. Units can be operated in touch or contimnuous mode using the three-way switch. Supplied with 3-inch Platform and Pop-Off Cup. The 3-inch Platform accepts single or multiple test tubes, beakers and various flasks. The removable cover, made of thermoplastic rubber, is highly impervious to chemicals. The Pop-off Cup can accommodate tubes of virtually any size or shape from micro tubes to those up to 40mm in diameter and 200mm in length. Power required: 230 V, Hz: 50, AMPS: 0.5. Weight: 4 Kg (8.8 lb) Base Dimensions: (DxWxH) 165x122x165mm (6.5x4.8x6.5in). | Piece
| High Precision Weighing Balance for DST: Electronic analytical balance: ISO/GLP compliant print out, in conjunction with printer or PC, Data interface port RS232C, Display Backlit LCD, Calibration/ Adjustment: internal, isoCAL, Pan size Ø 80 mm, Repeatability ≤± 0.02 / 0.05 / 0.1 mg, Linearity ≤± 0.03 / 0.1 / 0.2 mg | Piece
| LED Microscope: Zeiss’ Primo Star iLED microscope, with following product specifications: high-grade optics; reflected rather than transmitted blue light for fluorescent applications; battery power on AC failure; easy switching between fluorescent and brightfield viewing; utility outside a darkroom; >10,000 hour bulb life; LED and filter sets that would allow use of Auramine O and acridine orange for TB and parasitologic examinations. Binocular microscope for use with electric light via power line or alternatively via battery pack. Battery pack for uninterruptible power supply, and is usually included as an accessory  Observation tube: binocular, 30 deg inclination (viewing angle) and 360 deg rotation  Stage: rectangular, built in mechanical stage with vernier scale (minimum: 14 mm x 135mm). No polymer belts, metal cables, timing belt systems or non-metallic components are acceptable in the drive mechanism. Coaxial controls must be low mounted for ease of use. Stage finger assembly is to be slide friendly so that it does not damage or break slides | Piece
| Slide Cabinet: Slide Cabinet: To store atleast 20,000 slides of 75x25mm slide - Horizontal with anodized aluminum slide tray of 350x210x35mm, uniformly slotted rows are numbered 1 to 50 to identify the slides.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Piece
| Trolley, stainless steel: deep pressed, rounded cornered, stainless steel shelves. Sound deading pads below each tray. Strong, welded rigid construction. Technical specification: No. of shelves/trays/tiers or bins: 2; Dimensions [w x d x h], mm: 1050x558x948. | Piece
| Deep freezer: Deep freezer [[-20ºC]: Upright Model: CFC free high efficiency double refrigeration system for cooling and freezing filled in the bottom. Temperature: -36oC to 0oC + - 0.5oC Temperature Control: (i) Digital temperature controller (ii) Microprocessor Control/Microcontroller-for temperature setting (iii) Alarms for-Voltage, Over heat, Over cool Size: 500 litres with minimum 3 Stainless Steel, rust free shelves. Fixed in casters for easy maneuverability. Electricity: 230 volts AC, 50Hz single phase. Refrigeration system: Heavy-duty maintenance free refrigeration system with hermetically sealed refrigeration compressors and reliable cascaded refrigeration to minimize noise and vibration. Air-cooled with security lock to prevent unintentional switch off. Short cooling time of 4 to 5 hours at 400C ambient temperature. The equipment should be of continuous duty and frostfree. Access port for CO2 back up. Audiovisual Electronic Alarm System independent of power supply Electrically heated doors for quick opening of frozen doors. | Piece
| PCR hood: PCR hood: Features: UVTect™ microprocessor controller; Class 100 (ISO 5) vertical laminar flow air; Polycarbonate and polypropylene design to reflect UV energy; Full access folding sash; Unibody plastic design - no joints or gaps in construction; Digital UV light timer 0-59 minutes; Lab event timer; 24″, 32″, and 48″ widths; nominal depth 24″; 220V AC versions. UVTect™ Features: Variable speed blower control; 254 nm ultraviolet light with user set digital timer; Constant HEPA filter monitoring; Constantly monitors and alerts effectiveness of the shortwave UV bulb life; Lab event timer. External Dimension (in inches): width x height x depth (32"x39"x24"); Internal Dimension (in inches): width x height x depth (31"x27"x23") | Piece
Infectious Disease Serology Laboratory
| CENTRIFUGE: 5000 RPM (Gx2441-G000) | Piece
| INCUBATOR: SANYO (10202410) | Piece
| WATER BATH: Memmert (DIN 40050-IP-20) | Piece
| VORTEX: Genie-2/Abdos (2.41159/2.126764, VB3B008722) | Piece
| Rotatest Shaker: 4159B | Piece
| REFRIGERATOR (+2 -8): Vision Scientific, VS-1302L | Piece
| ELISA Plate Washer (12 manifold): Biotek, 207837/207836 | Piece
| ELISA Plate Washer (8 manifold): Biotek, 191866/183278 | Piece
| ELISA Plate reader: Sunrise 1002000585 | Piece
| ELISA Plate reader: ELx808, 207370 | Piece
| Freezer (-20): Whirlpool, CF41T300348002360 | Piece
| Balance: Mettler Toledo, B540527801 | Piece
ENTERIC DISEASE LABORATORY
| Autoclave: HV-50 HICLAVE, HIRAYAMA | Piece
| Autoclave: ES-215, TOMY | Piece
| Automatic pipetter (E fill) | Piece
| Bunsen burner (Footswitch): Fireboy plus, INTEGRA FIREBOY, Switzerland | Piece
| BACTEC: BACTEC 9050, BD | Piece
| Compound Light microscope: CH20i BIMF, Olympus | Piece
| Distillation Plant: GFL, Germany | Piece
| Dark field microscope: Lab.A1, Zeiss | Piece
| Densimat Turbidometer: 99234, Densimat, Biomerieux | Piece
| Electronic weighing balance: PL202-S, Mettler Toledo, Switzerland | Piece
| Fluroscence Microscope: BX51TRF, Olympus | Piece
| Gel Electrphoresis chamber: PowerPAC basic, BIORAD | Piece
| Gel Doc imaging system: XR+ GelDOC, BIORAD | Piece
| Incubator: B-30, Memmert, Germany | Piece
| Incubator: Memmert, Germany | Piece
| Incubator: INE-800, Memmert, Germany | Piece
| Incubator: 100-180, Memmert, Germany | Piece
| Incubator: Memmert, Germany | Piece
| Loop sterilizer and holder: OTO6, Dinkelberg Analytics | Piece
| Microcentrifuge: MIKRO-200, Hettich | Piece
| Micro-oven: SANYO | Piece
| Microcentrifuge(spinner): PRISM MINI, C1801, PRISM | Piece
| Micropipette 2-10µl | Piece
| Micropipette 5-40µl | Piece
| Micropipette 20-200µl | Piece
| Micropipette 20-1000µl | Piece
| PH meter: Lab 850, Schott Instrument | Piece
| PCR Thermal cycler: PIKO 24, Thermoscientific | Piece
| PCR Thermal cycler: T100 Thermal cycler, BIORAD | Piece
| PCR Cabinet: UVP PCR cabinet | Piece
| Refrigerator: YPM-110P, SANDEN, INTERCOOL | Piece
| Refrigerator: 54-7-12105, SNOWLAND | Piece
| Upright freezer: 900 series -86C, Thermoscientific | Piece
| Vortex mixer: MX-S, swirlex, Abdos | Piece
| Vortex mixer: 21915, Whirlimixer, NICKEL ELECTRO LTD | Piece
| Vortex mixer(3000rpm): VB821200180150 | Piece
| Waterbath: W 14, Grant Instrument LTD, UK | Piece
| Bio-Safety Cabinet: Bio-Safety Cabinet- Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
| Autoclave: Autoclave: Chamber Capacity: 53 Litters, Temperature setting range: 105°C to 132 °C,Temperature control: Digital, Microprocessor controlled, Pressure display range: 0 to 400 kPa, Maximum operating pressure: 216kPa, Operating pressure range: 0-186kPa, Sterilization heat source: 2.0kW electric heater, Dimensions of main unit (mm): 490 W x 560D x 1090H (Height from floor to control panel:875) , Chamber diameter: 325 x 740mm, Net Weight: 80 kg, Power input: 20A (120V), 10A (220V), 9A (230/240V) | Piece
| Distillation Plant: WATER STILL 4 LIT/HOUR WITHOUT STORAGE TANK 230V /1 PH 50/60HZ AND HOSE SET 2940 GFL, Germany | Piece
| Refrigerator/Freezer: Refrigerator/Freezer: side to side refrigerator/freezer, Poduct details: GWB207FVQA, Capacity: - Refrigerator 349L / Freezer 178L, Energy Consumption: - 540kWh / Energy Saving Rating, Colour: Silver/Stainless steel, Dimensions: 894(w) x 1753(h) x 725(d) mm, Freezer Capacity 178L, Fridge Capacity 349L, Frost Level No Frost, Interior Light (Freezer)- Yes, Interior Light (Fridge) -Yes, Weight 100kg, External Features: - Silver/stainless steel / LED Digital Display Interior Features: - 5x Digital Sensors / Express Freezing / No Frost Cooling / Child Lock / Door Open Alarm / 12kg Freeze Capacity / N Climate Class Refrigerator Compartment: - Multi-Flow / 3x Tempered Glass Shelves / Door Baskets / Veg Box - Moist Balance Crisper / Egg Bank Freezer Compartment: - Multi Air Flow / Ice Maker / Door Baskets / 3x Tempered Glass Shelves / 1x Drawer | Piece
| Micro oven: 30 litre, 900 watt silver free standing microwave oven with an electronic timer, safety door locks, turnable table, hub and bearing washer, control panel, auto defrost, child lock | Piece
| Aspirator bottle: Polythene heavy duty aspirator bottle for storing distilled water with stop cock, carrying handle, 50 litres capacity | Piece
| Electronic weighing balance: Readability: 0.01g, Maximum load:210g, Adjustmment: Cal Ext, Repeatability: 0.008g, Linearity: 0.02g, Balance size: 195x225x67mm | Piece
| Electronic weighing balance: Readability: 0.001g, Maximum load: 81g, Adjustmment: Cal Ext, Repeatability: 0.0008g, Linearity: 0.002g, Balance size: 195x225x137mm | Piece
| Safety Gas Burner & Safety gas tubing: Safety gas burner, with foot switch, IR sensor, with rest heat display and splash guard. can be operated with touch free IR-Sensor, button function or foot pedal. With the DualKnob for adjusting gas and air flow, the flame can be precisely regulated. can be operated with stationary natural gas and propane / butane gas supplies, cartridge gas or gas cylinders. Safety Gas Tubing: safety gas tube 750mm long, with clamps. Inner diameter 9mm | Piece
INFLUENZA/MOLECULAR LABORATORY
| Class II, Type B1-Biological Safety Cabinet: Model No-NU-427400E ( Brand-LABGARD) | Piece
| Class II, Microbiological Safety Cabinet: ENVAIR | Piece
| Water Bath | Piece
| Ependrof Centrifuge 5810 R: Sl No-5811BM567216 | Piece
| Inverted Microscope: Nikon | Piece
| Micro-Centrifuge: Hero Lab | Piece
| Centrifuge 5415 D | Piece
| Vortex Mixture | Piece
| 2-8°C Refrigerator | Piece
| ABI Machine ( 7500 Fast Dx Real Time PCR Instrument: Applied Biosystems | Piece
| Thermal Cycler-Rotor Gene: QIAGEN | Piece
| Thermal Cycler-Bio Rad: Bio-RAD | Piece
| Micro-Oven | Piece
| Electrophoresis Chamber | Piece
| Gel Doc ( Molecular Imager Gel Doc): Bio- RAD | Piece
| Freezer -30°C: Sanyo | Piece
| PCR Hood: ESCO | Piece
| Combi-Spin | Piece
| Fluroscent Microscope | Piece
| Egg Incubator | Piece
| Deep Freezer -80°C: SANYO | Piece
| Freezer -20-40°C: INSEGNIA | Piece
| Deep Freezer -80°C: INSEGNIA | Piece
| Dry Ice Machine | Piece
| Humidified Incubator | Piece
| Benchtop centrifuge: Benchtop centrifuge for microtitre plates: Spin speed up to 2000 rpm for microtitre plates ; Timed operation (1 to 90 minutes) with automatic switch-off; digital display; power supply 230V; weight 13-14 Kg; accompanied with rotors (changeable) for accomodating at least 2 standard 96 well-microtitre plates; | Piece
| Personal microcentrifuge: Personal microcentrifuge for 0.2 ml tubes: Holds two 8-strip or 16 individual 0.2 ml tubes, 6000 rpm, 2000 x g; Convenient on/off switch | Piece
| CO2 Cylinder: CO2 Cylinder for Dry Ice maker with siphon ( FILLED WITH CO2 GAS): Cylinder type 50 lb, Grade Industrial, Purity 99.5%,Volume (cu.Ft) 436, Volume (Kg) 22.7, Volume (lb) 50, Pressure (PSIG) 835 | Piece
Water Quality Testing Laboratory
| Ammonia Gas sensing electrode: 95-12 Orion 95-12 | Piece
| Arsenic Kit: Lovibond | Piece
| Conductivity/Tds Meter: 44600 Hach 44600 | Piece
| Del Agua water testing kit: Del Agua Del-501-00 | Piece
| Digital Burette: Bibby | Piece
| Digital Burette: Titrettee | Piece
| Distillation Plant: 2001/4 GFL 2001/4 | Piece
| DO Meter: ELE International | Piece
| Electronic Balance: Metler Toledo BD202: Metler Toledo BD202 | Piece
| Filter Manifold: Millipore | Piece
| Filter Manifold: Gelman Sceiences: Gelman Sciences | Piece
| Hot Air Oven: SANYO Gall OMT150xx2 | Piece
| Hotop: Tarsons MC01-5020 | Piece
| Incubator: WTW | Piece
| Incubator: SANYO MIRI 162 | Piece
| Oven(Thermolyne): Sybron F6030 CM | Piece
| Oximeter: WTW oxi330 | Piece
| Oximeter(BOD Meter): WTW OxiTop | Piece
| Ph Meter: Jenway 3070 | Piece
| Ph Meter: Schott Instruments Lab 850 2001759 | Piece
| pH/ORP meter: Hanna HI 98172 | Piece
| Photometer: Palintest | Piece
| Pipette Washer: ABC LABO | Piece
| Refrigerator: Samsung RT26M | Piece
| Sample Digester: Hach 44336-62 | Piece
| Spectrophotometer(UV-160A): Shimadzu PH26402A2 | Piece
| Spectrophotometer: Hach DR2000(44800-60) | Piece
| Turbidity Meter: Hanna HI93703 | Piece
| Turbidity Meter: Jenway 6035 | Piece
| Turbidity Meter: Lutran TU-2016 | Piece
| Vaccum Pump: KnF VDE0530 | Piece
| GARMIN(GPS): OREGON-650 | Piece
| Arsenic meter: Lovibond | Piece
| Bio-Safety Cabinet: Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
| Laboratory filtration pump: The lab filtration pump that transfers filtered liquids directly to waste; Ideal for use both on the work bench and in the laminar flow hood where its small footprint; No need to use additional containers to capture the waste liquid; Quiet design:Optimized to reduce the noise down; Maintenance-free: pump technology and the materials of construction provide reliable vacuum performance and make the pump maintenance-free. The diaphragm is made of highly durable chemically resistant polytetrafluoroethylene | Piece
| Manifold filter: Manifold filter 6-place for rubber stopper tulip; Overview: Vacuum manifolds support simultaneous filtration of six test samples. Use Millipore glass, plastic or stainless steel filter holders fitted with standard No. 8 perforated stoppers over holder outlets. Manifold is available with the new Microfil head for use with presterilized, ready-to-use funnels and membranes; Application: General microbiological analysis, Water Monitoring; Materials:Stainless Steel. | Piece
| Acid Digestion Fume Hood with base cabinet: Acid Digestion Fume Hood with base cabinet: should be designed to perform digestion applications; should built with unplasticised polyvinylchloride (u-PVC) internal surfaces (including the work surface) and a polycarbonate sash window. UL classified; Full uPVC or PP interior; Excellent protection against acids; Excellent aesthetics; 5 degree sloped front; Superior containment at 0.3m/s; High energy savings; Superstructure: Tri-wall construction for maximum robustness; U-PVC internal chamber constructed of unplasticised polyvinylchloride; Chain and Sprocket Sash support system; Hood lighting should be pre-wired. Electronically ballasted lighting, energy efficient, instant start. Typical light intensity on work surface should be >1076 lux in zero ambient conditions; U-PVC exhaust collar ensures superior chemical resistance. Safety Features: Shaping vanes increase airflow "sweep" at the critical area at the sidewall to improve containment, especially when laboratory personnel walk fast in front of the hood; Sash made from polycarbonate to prevent hydrofluoric acid fume etching; Sash stop limits sash movement beyond 457mm (18.0"), 'encouraging' user to work at safe positions; When sash is raised above 457mm (18.0"), it should automatically and gently fall back to the safe level unless held in place. Allowing the sash to be raised temporarily to the full open position for set up of equipment and apparatus in hood, while enforcing regular operation of the hood with lowered sash; Aerodynamic foil entry provides maximum airflow to "sweep" the critical boundary layer near the work surface level to reduce turbulence and eliminate backflow; Serviceability: U-PVC internal work zone allowing for easy cleaning of all surfaces; Removable baffle system allowing for thorough cleaning inside the hood; Chain and sprocket system should require minimal service; Internal access panels on side walls should facilitate access to plumbing connections for service especially when hoods are installed next to each other; Removable front panel facilitates easy access to lighting and other electrical components mounted above the hood chamber. Local ducting and exust fan should be included. | Piece
| Glassware Dryers: External – (h) x (w) x (d) -1626 x 1168 x 610(mm); Internal – (h) x (w) x (d)-1430 x 1165 x 590; Capacity (Liters)-1000; Air circulation-by hot air fan; Voltage- 230v | Piece
| Autoclave: High pressure steam sterilizer ES-315 use for sterilization and decontamination of laboratory equipments and waste. The vertical autoclave with a top-open lid, which can be operated easily with a single hand or foot. Operating temperature range (during sterilizing): 105 - 132°C, Operating pressure range: 0 - 186kPa, Maximum operating pressure 216kPa, Temperature control: Digital, microprocessor controlled, Temperature display / display range: Digital / -15 - 180°C, Pressure display / display range: Analog / 0 - 400kPa, Heat source: 2.0kW electric heater, Timer control: Digital, microprocessor controlled, Timer display: Digital. Safety devices: • Inside the chamber overheat prevention, • Outer wall of the chamber overheat prevention, • Overpressure prevention, • Temperature sensor disconnection prevention, • Empty heating prevention, • Leakage breaker, • Safety valve. Malfunction prevention devices: • Low water level detection, • Exhaust valve knob open/close detection, • Insufficient sterilization detection, • Chamber lid open/close detection. Chamber material: SUS304. Capacity / material of exhasut bottle: 3ℓ / polyethylene. Type / material of lid gasket: Accommodating internal pressure / silicone rubber. Dimensions of main unit (mm): 490W x 560D x 1090H (Height from floor to control panel:875). Net weight: 80kg. Rated Voltage AC 120 / 220 / 230 / 240V. Power input: 20A (120V), 10A (220V), 9A (230 / 240V), Power supply requirements: Single-phase 220 / 230 / 240V AC (50 / 60Hz) 15A or above. Power Comsumption (calorific value): 2kW (1720kcal/h). Accessories: • Stainless steel basket with slatted bottom plate 1, • Water level sensor 1, • Inspection sheet 1, • Warranty card 1, • Customer card 1, • Operator's manual 1,• Clear plastic accessory case 1, • Attachment screw for clear plastic accessory case 1. Stainless basket (φ300 x 182mm) | Piece
| Vortex Mixer: 3-inch Platform, Pop-off Cup, Speed (RPM), 600-3200, (2700 for 50Hz models), Dimensions (base), (D x W x H), 165 x 122 x 165mm, (6.5 x 4.8 x 6.5in) | Piece
| Incubator: Temperature range: Ambient + 5°C to 80°C. Temperature control: Digital P.I.D. controller. Dual LED display. Capacity: 50 - 100 liters.Power supply: AC 220V- 230V, 50HZ ; shelf provided:-2-3 | Piece
| HPLC: For detection of organic contamination and Pesticides in water. Ultra high pressure pump optimized for fast liquid chromatography, capable of pressures up to 1000 bar (15000 psi) and fl ow rates up to 5 ml/min. The electronically controlled pump drive provides for low pulsation, a highly precise fl ow rate and a stable baseline. Flow rate- 0.01 – 5.0 ml/min or 0.001 ml increments; Flow rate accuray- ± 1 %;n Flow rate precision- < 0.1 %; Features: touchscreen, active pressure and pulsation compensation, RFID automatic pump head detection (GLP), interchangeable pump heads; | Piece
| ICP-OES: Detector Type-High performance solid-state CID86 chip; Wavelength Range- 166 to 847nm; Auxiliary Gas- Mass flow control, 0-2 L/min; Data Acquisition Modes-Precision mode, Speed mode; Depth (Metric)- 750mm; Features- Intelligent introduction and rinse, Speed data acquisition mode, Get ready: automatic performance checks; Full-frame Capability- Yes; Height (Metric)- 590mm; Nebulizer Gas- Mass flow control, 0-2 L/min; Peristaltic Pump- 4-channel, mini peristaltic pump, Speed: 0 to 125rpm; Plasma Gas- Mass flow control, 0-2 L/min; Plasma Viewing-DUO or Radial; RF Source-27.12MHz solid state, 750 to 1500w output power (DUO restricted to 1350w); Spectral Bandpass- 7pm at 200nm; Spectrometer-Simultaneous echelle type, 52.91 grooves/mm ruled grating, 383mm effective focal length, 9.5° UV fused silica cross dispersion prism; Standard Sampling Kit- Concentric glass nebulizer, Glass cyclonic spray chamber, Semi-demountable EMT torch, 1.5mm bore quartz center tube (Radial), 2mm bore quartz center tube (DUO); Width (Metric)-840mm; Function- Detection of trace elements in water | Piece
National Drug Testing Laboratory
| UHPLC: knauer | Piece
| Dissolution Apparatus: 708 DS Agilent Technologies | Piece
| Disintegration Apparatus: Labindia | Piece
| Autotitrator: Metrohm | Piece
| Autotitrator: Fischer Scientific | Piece
| UV-Vis Spectrophotometer: Perkin Elmer | Piece
| Friability Apparatus: Pharma test | Piece
| Analytical Balance: Shimadzu • Perform basic weighing with below balance, counting, percent, animal weighing, and density determination functions • Anti-theft feature secures balance to work station • Functional in 8 languages • Document critical weights in GLP mode via RS-232 communication with PC or printer • Sliding doors and removable shield chamber parts make cleanup easy • Backlit high-contrast display delivers excellent viewability in any lighting condition • Calibration:external• Capacity (g): 220g • Dimensions (" W):11.9• Dimensions (" H):13• Dimensions (" D):9.1 • Display: back light LCD• Readability (g):0.0001• Linearity (g):0.0002 • Repeatability (g):0.0001 • Weighing units: g, mg, kg, ct, oz, troy, lb, pennyweight, grain, Newton, momme, mesghal, tael • Stabilization time:205 secs• Power (VAC):120/220 • RS-232 interface: yes with miniprinter• Output:RS232 miniprinter • Accesorries: miniprinter | Piece
| Microscope-olympus | Piece
| pH meter - schaut instruments | Piece
| Water Deionizer- millipore | Piece
| Water Deionizer/Distillation Unit: millipore | Piece
| Centrifuge: Fischer Scientific | Piece
| Automatic shaker- werke | Piece
| Automatic shaker- bibby scientific | Piece
| Sonicator - rivotek | Piece
| Hot Plate- Fischer Scientific | Piece
| Polarimeter: :sugar scale ( with automatic temperature compensation).Measurement readings= angle of rotation, international sugar scale ( without temp compensation) , international sugar scale ( without automatic temp compensation) , specific optical rotation, concentration and purity. • Measurement range = angle of rotation: -89.999 to +89.999° , international sugar scale : -130 to +130 °z. • Resolution= angle of rotation: 0.001°(Default 0.01°z) , international sugar scale : 0.001°z ( default 0.01°z). • Measurement accuracy = angle of rotation displayed value +_ 0.01° ( -35 to +35°) , relative precision +_ 0.2% (-35.01 to -89.99° , +35.01 to +89.99°).. • International sugar scale displayed value +_ 0.03°z ( -101 to +101°z.. relative precision +_ 0.2% (-130 to - 101.01°z , +101.01 to +130 °z) • Length of observation tube= upto 200mm • Temperature = 10 to 30°c• Temp accuracy = +_ 0.5°c• Temp resolution = 0.5 °c • Display panel = backlit color LCD • Graphic display = indicating the conformity of measurement values of measurement with limit settings. • Output= 1) digital printer ( dot printer matrix printer) 2) RS-232 C input / output ports to a computer system. • Light source = halogen lamp• Accessories = observation tube 100mm n 200mm • Power supply = AC 100v to 240 V• Dimension = 48.5 × 28.5× 14.5 mm , 14.4 kg | Piece
| Particulate Matter Counter: Benifits: Defect Reduction,100% sample volume for greater accuracy,A precise sample volume within 0.1 ml is achieved with the syringe sampler, producing repeatable results,SamplerSight-Pharma software comes with a Validation Information Notebook, providing comprehensive software development documentation and full IQ, OQ, and PQ protocolsSamplerSight-Pharma complies with FDA21 CFR 11User-Friendly, Reduce operator error and ensure process ,accuracy with recipes for repeat sample processing to USP, EP, or JP standards • Automated particle counter calibration wizard for full calibration or routine verification • Alarm levels for pass/fail criteria ensure quality control Comprehensive System: Easy-to-use, pull-down menus,Security password maintains system integrity, Particle measurements are reported per ml or per container values,Small sample volume minimizes waste of expensive product,Various custom report formats are available.with a computer system provided. Features: Sizing capability from 1.5 to 125 µm, Fifteen user-selectable sizing channels,Meets and exceeds all USP, EP and IP requirements • Sample volumes from 0.2 ml to 1 liter • Adaptors for SVI and IV bag testing • Automated sensor calibration Applications: Pharmaceutical parenteral monitoring to USP, EP, IP and FDA standards,Parts/medical device cleanliness testing, Laboratory water sampling,WFI and Purified Water testing,Filter efficiency testing. | Piece
| FTIR: 1. Interferometer:25 mm, permanently aligned, Michelson, 45°, mechanical flexure 2. Spectrometer interface:USB-2 3. Size :16 x 31 x 13 cm (6 x 12 x 5 in) (W x D x H with ATR) 4. Weight :3.8 kg (8 lb) (with ATR) 5. Sample interface:DialPath, TumblIR, Diamond ATR, Diffuse Reflectance, Transmission 6. Software: PC software, Automated IQ/OQ, 21 CFR Part 11 compliant, Resolutions Pro for advanced data analysis. 7. Power:220 – 250 V AC, 60/50 Hz 8. Spectral range: KBr 6300–350 cm-1 and ZnSe 5100–600 cm -1 9. Spectral resolution: < 2 cm-1 10. Wavenumber accuracy:0.05 cm-1 Measured with ASTM 1921 11. Wavenumber reproducibility:0.005 cm-1 Measured with ASTM 1921 12. computer system and Printer attached. 13.Warranty: Five (5) year interferometer, Five (5) year laser, Three (3) year source Location dependent,Hardware support period. Seven (7) years from date of last unit manufacture. After this time,parts and supplies will be provided if available 14.Software support: Software upgrades to add additional functionality .. IQ, OQ and PQ to be conducted at the site of installation. | Piece
| Water Bath: Chamber Capacity: 34.5 liters (9.1 gallons) Temperature Display: Ambient +5 °C to 99.9 °C Dimensions (L x W x H) Work Area: 12 x 27 x 7.5 in (30.5 x 68.6 x 19.0 cm) Overall without cover: 14 x 36.5 x 9.5 in. (35.6 x 92.7 x 24.1 cm) Shelf: 11 11 x 26.5 in. (27.9 x 67.3 cm) Electrical Specs (50/60Hz) 120V : 1550 Watts 230V: 1550 Watts Temperature controller provides +/-0.05 °C uniformity at 37 °C and sensitivity of +/-0.05 °C with stainless steel gable cover • Water flow directed around bath perimeter • Elevated sample platform for more thorough circulation • Easy-to-maintain, corrosion resistant construction • Automatic sensor to protect bath from overheating • Single switch control for heater and circulator pump • Meets ASTM E715 Class IIA standards for uniformity | Piece
| Fumehood: Fume hood- should be designed to perform digestion applications; should built with unplasticised polyvinylchloride (u-PVC) internal surfaces (including the work surface) and a polycarbonate sash window. UL classified; Full uPVC or PP interior; Excellent protection against acids; Excellent aesthetics; 5 degree sloped front; Superior containment at 0.3m/s; High energy savings; Superstructure: Tri-wall construction for maximum robustness; U-PVC internal chamber constructed of unplasticised polyvinylchloride; Chain and Sprocket Sash support system; Hood lighting should be pre-wired. Electronically ballasted lighting, energy efficient, instant start. Typical light intensity on work surface should be >1076 lux in zero ambient conditions; U-PVC exhaust collar ensures superior chemical resistance. Safety Features: Shaping vanes increase airflow "sweep" at the critical area at the sidewall to improve containment, especially when laboratory personnel walk fast in front of the hood; Sash made from polycarbonate to prevent hydrofluoric acid fume etching; Sash stop limits sash movement beyond 457mm (18.0"), 'encouraging' user to work at safe positions; When sash is raised above 457mm (18.0"), it should automatically and gently fall back to the safe level unless held in place. Allowing the sash to be raised temporarily to the full open position for set up of equipment and apparatus in hood, while enforcing regular operation of the hood with lowered sash; Aerodynamic foil entry provides maximum airflow to "sweep" the critical boundary layer near the work surface level to reduce turbulence and eliminate backflow; Serviceability: U-PVC internal work zone allowing for easy cleaning of all surfaces; Removable baffle system allowing for thorough cleaning inside the hood; Chain and sprocket system should require minimal service; Internal access panels on side walls should facilitate access to plumbing connections for service especially when hoods are installed next to each other; Removable front panel facilitates easy access to lighting and other electrical components mounted above the hood chamber. Local ducting and exust fan should included. | Piece
| Hardness Tester:: Tablet hardness tester 1. Tablet size= 2mm to 30mm with accuracy of +0.5% 2. Hardness measurement by load cell up to= 500N 3. Units of measurement= N,kgf, and lbf 4. PC connectable 5. Programmable upper T1 and lower T2 tolerance limit 6. Automated Semi calibration 7. Can display of test value and calculate average 8. LCD display with back light 9. Acrylic glass cover with disposable container 10. Operates with rechargeable batteries and 6 v dc adaptor 11. Accessories =IQ,PQ ,OQ documents, printer,6V dc adaptor | Piece
| Trolley, stainless steel: deep pressed, rounded cornered, stainless steel shelves. Sound deading pads below each tray. Strong, welded rigid construction. Technical specification: No. of shelves/trays/tiers or bins: 2; Dimensions [w x d x h], mm: 1050x558x948 | Piece
| UV-Spectrophotomter: Spectrophotometer with ES Software and PC System and printer • Principle :Double-beam, double monochromator, ratio recording spectrophotometers. • Optical System:All-reflecting optical system (SiO2 coated) with holographic grating monochromators. Littrow mounting, sample thickness compensated detector optics. • Detector: Photomultiplier (UV/Vis) • Wavelength range: 190-1100 nm • Accuracy: ±0.1 nm • Absorbance Range: -3 to 4 A • Bandwidth: 0.5-4 nm, variable • Stray Light: <0.01% at 220 nm • Photometric Accuracy: ±0.001 A • Dimensions: 65W x 56D x 23.3H cm (25 x 22 x 9") Weight: 25 kg (57 lbs.) IQ, OQ and PQ to be conducted at the site of installation. | Piece
| Table Top Digital Timer: • 10-hour timing range tabletop clock with timer • Includes alarm, count-up, time-out function, and four-mode stopwatch • One-inch digital display for easy visibility • Runs on one C battery | Piece
| Digital LCD Alarm Thermometer Hygrometer Temperature Humidity Meter Features: Convenient wall mount and mini design Can check the max / min temperature and humidity Memorize or erase the last minimum and maximum temperature / humidity Ideally placed on the desk Indoor temperature, humidity, displayed at the same time, it has °F / °C to choice Hourly chime function, alarm clock function Ideal household decoration and ensure you the most desirable ambient temperature and humidity Multifunction instrument can display temperature, time and humidity all at the same time Powered by one AAA battery. (not included) Specification: Range (indoor): 0°C~50°C; 32°F~122°F (Temperature); 10%~99% (Humidity) Accuracy (temperature): ±0.1°C Accuracy (humidity): ±5% Resolution: 0.1°C (temperature); 1% (humidity) Indoor temperature, humidity, time is displayed at the same time °C / °F LCD screen dimension: 65 x 35 mm Dimension: Approx. 8 x 7.5 x 2cm (W x H x T) Max key: the highest measurement environment Min key: the lowest temperature / humidity | Piece
| Hot air oven: Interior: Stainless steel, material 1.4301 (ASTM 304),with all-round deep-drawn ribs to integrate the large-area heating with ceramic-metal sheath. Insertions: Stainless steel grid shelves Housing: Textured stainless steel, rear zinc-plated steel,with 4 feet, sizes 450, 750 and 1060 mounted on lockable castors, with mains cable and plug. Fresh air: Admixture of pre-heated fresh air by electronically adjustable air flap. Stainless steel interior Volume approx. l/cu. ft. :74/2.6 Width (A) mm/in: 400/15.8 Height (B) mm/in:560/22.0 Depth (less 39 mm/1.5 in for fan) (C) mm/in :330/13.0 Stainless steel grid shelves: 2 Max. number of grid shelves/shelves number: 6 Max. load per grid shelf/shelf : kg/lbs 20/44.1 Max. load of chamber kg/lbs :120 /2640 Textured stainless steel exterior Width (D) mm/in : 585/23.0 Height (size 450, 750, 1060 with castors) (E) mm/in: 944/37.2 Depth (without door handle),door handle + 56 mm (2.2 in) (F) mm/in : 514/20.2 Working-temperature range °C/°F: at least 5 (9.0) (UN/UNplus) at least 10 (18.0) (UF/UFplus) above ambient temperature to +300 (572) Setting temperature range °C/°F :+20 (68) to +300 (572) Setting accuracy °C/°F: up to 99.9 (211.8): 0.1 (0.2) / from 100 (212): 0.5 (0.9 | Piece
| HPLC: Physical specifications Features Description Pump system Hydraulic system :Dual pistons in series pump with proprietary servo-controlled variable stroke design and smooth motion control for active damping. Pump resolution step size :300 pl step size Settable flow range: 0.001 – 5 mL/min, in 0.001 mL/min increments (- executed in 300 pl/step increments) Flow precision :≤ 0.07 % RSD, or 0.01 min SD, whatever is greater (0.2 - 5.0 mL/min); based on retention time at constant room temperature Flow accuracy :}1% or }10 μL/min, whatever is greater; pumping degassed H2O at 10 MPa (100 bar) Maximum operating pressure:Operating range up to 120 MPa (1200 bar) up to 2 mL/min, ramping down to 80 MPa (800 bar) up to 5 mL/min Pressure pulsation :< 1 % amplitude or < 0.5 MPa (5 bar) whatever is greater at 1 mL/min water Compressibility compensation:Automatic, pre-defined, based on mobile phase selection Gradient formation: Low pressure quaternary mixing Delay volume: Standard configuration: < 350 μL With optional V380 Jet Weaver Mixer: < 500 μL Composition range: Settable range: 0.0 – 100.0 % Recommended range: 1.0 – 99.0 % or 5 μL/min Composition precision :< 0.15 % RSD, or 0.02 min SD whatever is greater (1 mL/min); based on retention time at constant room temperature Composition accuracy :}0.4 % absolute (1 - 99 % B, 0.5 - 2.0 mL/min with water/caffeine tracer, 400 bar) Integrated degassing unit:Number of channels: 4 Internal volume per channel: 1.5 mL Materials in contact with solvent:TFE/PDD copolymer, FEP, PEEK, PPS, stainless steel, polyimide Communications: Controller-area network (CAN), RS232C, APG remote: ready, start, stop and shutdown signals, LAN Safety and maintenance :Extensive diagnostics, error detection and display through company LabAdvisor, leak detection, safe leak handling, leak output signal for shutdown of the pumping system. Low voltage in major maintenance areas. GLP features :Early maintenance feedback (EMF) for continuous tracking of instrument usage in terms of seal wear and volume of pumped mobile phase with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors. | Piece
| HPLC continuation: Autosampler Physical specifications Weight: 15.5 kg (34.2 lbs) Dimensions :200 x 345 x 440 mm (height x width x depth) (8 x 13.5 x 17 inches) Line voltage: 100 – 240 VAC, ± 10 % Wide-ranging capability Line frequency :50 or 60 Hz, ± 5 % Power consumption :300 VA Maximum (apparent power) Power consumption (active power): 200 W Maximum Ambient operating temperature: 4–55 °C (39–131 °F) Ambient non-operating temperature: -40–70 °C (-40–158 °F) Humidity: < 95 %, at 25–40 °C (77–104 °F) Non-condensing Operating altitude: Up to 2000 m (6562 ft) Non-operating altitude :Up to 4600 m (15091 ft) Performance specifications Injection range: 0.1 – 20 μL With standard loop 0.1 – 40 μL With extended loop in 0.1 μL increments Precision Typically :< 0.25 % RSD from Measured with injections 5 – 20 μL, Typically < 0.5 % RSD of benzyl alcohol. from 2 – 5 μL volume, Typically < 0.7 % RSD from 1 – 2 μL volume Accuracy :± 1 % (10 μL, n-10) Pressure range: Up to 1200 bar Cooling/Thermostatting Compatible with G1330B Autosampler Thermostat, 4-40°C Sample viscosity range 0.2 – 5 cp Sample capacity: Capacity 2 x well plates (MTP) + 10 x 2 mL vials, 108 x 2 mL vials in 2 x 54 vial plate plus 10 additional 2 mL vials, 30 x 6 mL vials in 2 x 15 vial plate, plus 10 additional 2 mL vials, 54 Eppendorf tubes (0.5/1.5/2.0 mL) in 2 x 27 Eppendorf tube plate Injection cycle time :Typically < 19 s without needle wash and typically <25 s with 3 s needle wash; conditions: default draw (100 μL/min) and eject speed (400 μL/min). Injection volume :< 5 μL (water) | Piece
| HPLC contiuation: Thermostatted Column Compartment G1316C Physical specifications Weight :11.2 kg (24.7 lbs) Dimensions: 140 x 345 x 435 mm (height x width x depth) (5.5 x 13.5 x 17 inches) Line voltage :100 – 240 VAC, ± 10 % Wide-ranging capability Line frequency :50 or 60 Hz, ± 5 % Power consumption :320 VA / 150W / 512 BTU Maximum Ambient operating temperature: 0–55 °C (32–131 °F) Ambient non-operating temperature: -40–70 °C (-40–158 °F) Humidity :< 95 %, at 25–40 °C (77–104 °F) Non-condensing Operating Altitude: Up to 2000 m (6562 ft) Non-operating altitude :Up to 4600 m (15091 ft | Piece
| HPLC continuation: Performance specifications Temperature range :10 °C below ambient to 80 °C (5 mL/min) 10 °C below ambient to 100 °C (2.5 mL/min) Temperature stability: ± 0.05 °C Temperature accuracy: ± 0.8 °C ± 0.5 °C With calibration Column capacity • 2 columns of 300 mm length Up to three 1290 Infinity with individual solvent heating TCCs can be clustered • 3 columns of 300 mm length in a system. These will in series appear with a single joint • 4 columns of 100 mm length user interface in the with opt. Low dispersion control software.* heat exchanger • 6 columns (max. 300 mm) with triple-TCC clustering and column selection valves • 8 columns (max. 100 mm) with dual-TCC clustering and column selection valves Warm-up/cool-down time 5 min from ambient to 40 °C 10 min from 40 – 20 °C Temperature zones • 2 (single TCC) • 4 (dual-TCC cluster) • 6 (triple-TCC cluster) Individual standby temperatures definable for cluster configurations. Dead volume • 1.6 μL low dispersion heat The use of low dispersion exchanger heat exchangers is rec- • 3 μL left heat exchanger ommended. 3 μL and 6 μL • 6 μL right heat exchanger heat exchanges are available for backwards compatibility. Available valves • 8 position / 9 port, 8x column The optional valve drive selection valve, min. (G1316C #058) needs to be 2 1290 Infinity TCCs required, present. Valve kits have to (G4230A 400 bar, be ordered as separate (G4230B 1200 bar) products. • 2 position / 6 port (G4231B),1200 bar • 2 position /10 port (G4232B), 1200 bar | Piece
| HPLC continuation: Diode Array Detector G4212A Physical specifications Weight: 11.5 kg (26 lbs) Dimensions: 140 x 345 x 435 mm (height x width x depth) (5.5 x 13.5 x 17 inches) Line voltage :100 – 240 VAC, ± 10 % Wide-ranging capability Line frequency: 50 or 60 Hz, ± 5 % Power consumption :160 VA / 130 W / 444 BTU Maximum Ambient operating temperature: 5-40 °C (41–104 °F) Ambient non-operating temperature :-40–70 °C (-40–158 °F) Humidity Operating: < 80 % Non-condensing Operating altitude Up :to 2000 m (6562 ft) Non-operating altitude :Up to 4600 m (15091 ft) For storing the module | Piece
| HPLC continuation: Performance specifications Detection type :1024-element photodiode array Light source UV-lamp Equipped with RFID tagthat holds lamp typical information. Wavelength range :190 – 640 nm Short term noise (ASTM) Single < ± 3 x 10-6 AU at 230/4 nm, and Multi-Wavelength slit width 4 nm, TC 2 s ASTM Drift < 0.5 x 10-3 AU/hr at 230 nm Linear absorbance range > 2.0 AU (5 %) at 265 nm Typically 2.5 AU (5 %) Wavelength accuracy ± 1 nm After recalibration with deuterium lines Wavelength bunching 2 – 400 nm Programmable in steps of 1 nm Number of signals 8 Slit width 1, 2, 4 , 8 nm Programmable slit Diode width ~ 0.5 nm Signal data rate up to 160 Hz Spectra Data rate up to 160 Hz Flow cells Max-Light Cartridge Cell Cartridge type, equipped (10 mm, V(s) 1.0 μL), with RFID tags that holds 60 bar (870 psi) cell information. pressure maximum Max-Light Cartridge Cell (60 mm, V (s) 4.0 μL), 60 bar (870 psi) Max-Light Cartridge Test Cell Accesorries: Computer system with a printer. IQ, OQ and PQ to be conducted at the site of installation. | Piece
| Karl Fischer titrator: Measuring interface A measuring interface with one high-ohm measuring input (Ind.) for pH, redox or ISE electrodes, one measuring input (Ref.) for a separate reference electrode, one measuring input (Temp.) for temperature sensors (Pt1000 or NTC), one measuring input (Pol.) for a polarizable electrode and one connector (iConnect) for an iConnect Titrando: • Type plate:Contains specifications concerning mains voltage, instrument type and serial number. • USB connector (USB 1 and USB 2) USB ports (type A) for connecting printer,keyboard, barcode reader, additional Titrandos, USB Sample Processor, etc • Controller connector Connection of a Touch Control or PC with installed PC software. Mini DIN, 9-pin • MSB connector (MSB 1 bis MSB 4 Serial Bus. For connecting external dosing devices, stirrers or remote boxes. Mini DIN, 9-pin. • Mains connection socket • Measuring interface 1 (Input 1) • Electrode connector (Ind.) For connecting pH, redox or ISE electrodes with integrated or separated reference electrode. Socket F. • Electrode connector (Ref.) For connecting reference electrodes, e.g. Ag/AgCl reference electrode. Socket B, 4 mm. • Temperature sensor connector (Temp.) For connecting temperature sensors, e.g. Pt1000 or NTC sensors. Two B sockets, 2 mm. • Electrode connector (Pol.) For connecting polarizable electrodes, e.g. double Pt electrodes. Socket F. • iConnect connector (iConnect) For connecting electrodes with integrated data chip Titration vessel / 50-150 mL: Material: Clear glass ,Height (mm): 80, Outer diameter (mm): 78 ,Volume (mL): 20 ...90 Titration vessel / 20-90 mL: Material: Clear glass,Height (mm): 80,Outer diameter (mm): 78,Volume (mL): 50 ...150 KF titration vessel lid:For volumetric KF titrations. Material: PTFE,Note, Material: Insert, KF adsorber tube:For KF equipments. Including cover and O-ring.,Height (mm): 113,Outer diameter (mm): 8 Septum / 12 mm / 5 pieces:Set of 5 pieces. Material: Silicone rubber.Height (mm): 2,Outer diameter (mm): 12 Stirring bar / 16 mm: Stirring bar with magnetic core, PTFE covering, 16 mm long.,Material: PTFE,Outer diameter (mm): 4,Length (mm): 16 Stirring bar / 25 mm:Stirring bar with magnetic core, PTFE covering, length 25 mm,Material: PTFE,Outer diameter (mm): 5,Length (mm): 25 Holding clamp for bottles:Holding spring for reagent bottles and exchange units. Glass weighing spoon:Material: Glass,Length (mm): 120 | Piece
| Karl Fischer titrator continuation: Molecular sieve:Molecular sieve. Bottle of 250 g. Pore size: 0.3 nm. Without moisture indicator. For Rancimats and Karl Fischer instruments. Stirrer:Magnetic stirrer with base plate, support rod and electrode holder. With permanently attached cable for MSB (Metrohm Serial Bus). Ti Stand Compact titration stand with magnetic stirrer and built-in pump for adding solvent and aspirating the titration vessel content. Suitable for Karl Fischer titrations with Titrandos. MSB connector. Including extensive accessories, solvent and waste bottle. Dimensions in mm (B/H/T): 106/ 101/ 220.Remark Dimensions: Height with support rod 412 mm Touch Control Operating element for the Titrandos, USB Sample Processors and the 846 Dosing Interface. Touch-sensitive color display. Simple and intuitive user guidance. Two PCMCIA card slots. Default dialog language English, one additional language can be added later. Including CompactFlash card with adapter, demo version of the PC Control software. KF Thermoprep The KF Thermoprep is used for thermal sample preparation in Karl Fischer titration. Many substances are not suitable for a direct Karl Fischer titration because they are not soluble, react with the KF reagent or because they release their water content only very slowly or only at high temperatures. With the aid of tightly sealed sample vials the sample are brought to the oven and can afterwards be determined by every volumetric or coulometric KF titrator. Dosing unit 5 mL Dosing unit with integrated data chip with 5 mL glass cylinder and light protection, mountable to a reagent bottle with ISO/DIN GL45 glass thread. FEP tubing connection, antidiffusion buret tip.Volume (mL): 5 Dosing unit 10 mL Dosing unit with integrated data chip with 10 mL glass cylinder and light protection,mountable to a reagent bottle with ISO/DIN GL45 glass thread. FEP tubing connection, antidiffusion buret tip. Volume (mL): 10 Dosing unit 20 mL Dosing unit with integrated data chip with 20 mL glass cylinder and light protection, mountable to a reagent bottle with ISO/DIN GL45 glass thread. FEP tubing connection, antidiffusion buret tip. Volume (mL): 20 | Piece
| Karl Fischer titrator continuation: KFT:Karl Fischer titration for water determination. As measuring mode can be selected: – Ipol (voltametric measurement with adjustable polarization current) – Upol (amperometric measurement with adjustable polarization Voltage pH electrode Redox electrode ISE electrode Reference electrode Polarizable electrode Temperature sensor iConnect Printer IQ, OQ and PQ to be conducted at the site of installation. | Piece
Bio-Safety Level 3
| Bio-SAfety Cabinet: Bio-Safety Cabinet-Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
| Incubator- TOMY High pressure steam sterilizer ES-315 use for sterilization and decontamination of laboratory equipments and waste. The vertical autoclave with a top-open lid, which can be operated easily with a single hand or foot. Operating temperature range (during sterilizing): 105 - 132°C, Operating pressure range: 0 - 186kPa, Maximum operating pressure 216kPa, Temperature control: Digital, microprocessor controlled, Temperature display / display range: Digital / -15 - 180°C, Pressure display / display range: Analog / 0 - 400kPa, Heat source: 2.0kW electric heater, Timer control: Digital, microprocessor controlled, Timer display: Digital. Safety devices: • Inside the chamber overheat prevention, • Outer wall of the chamber overheat prevention, • Overpressure prevention, • Temperature sensor disconnection prevention, • Empty heating prevention, • Leakage breaker, • Safety valve. Malfunction prevention devices: • Low water level detection, • Exhaust valve knob open/close detection, • Insufficient sterilization detection, • Chamber lid open/close detection. Chamber material: SUS304. Capacity / material of exhasut bottle: 3ℓ / polyethylene. Type / material of lid gasket: Accommodating internal pressure / silicone rubber. Dimensions of main unit (mm): 490W x 560D x 1090H (Height from floor to control panel:875). Net weight: 80kg. Rated Voltage AC 120 / 220 / 230 / 240V. Power input: 20A (120V), 10A (220V), 9A (230 / 240V), Power supply requirements: Single-phase 220 / 230 / 240V AC (50 / 60Hz) 15A or above. Power Comsumption (calorific value): 2kW (1720kcal/h). Accessories: • Stainless steel basket with slatted bottom plate 1, • Water level sensor 1, • Inspection sheet 1, • Warranty card 1, • Customer card 1, • Operator's manual 1,• Clear plastic accessory case 1, • Attachment screw for clear plastic accessory case 1. Stainless basket (φ300 x 182mm) | Piece
| High Spec Centrifuge-Max. Speed 15200rpm Capacity 1.6L (4 x 400) Refrigerated No Max. RCF 25830 x G Hertz 50/60Hz Voltage 230V Control Microprocessor Display Digital Height (English) Exterior 14.2 in. Height (English) Open Cover 34.3 in. Height (Metric) Exterior 36cm Height (Metric) Open Cover 87cm Length (English) Exterior 23.8 lb. Length (Metric) Exterior 60.5cm Memory 99 programs (5 with direct access key) Motor Type Direct, brushless induction, low-profile Net Weight (English) 126.8 lb. Net Weight (Metric) 57.5kg Timer 99 hrs., 59 min. plus HOLD Width (English) Exterior 17.3 in. Width (Metric) Exterior 44cm Certifications/Compliance IVD compliant; IEC 61010-1, IEC 61010-2-020, IEC 61010-2-101; 230V models meet EN 292, EN 61326, EN 55011B. UL, CE, CSA, Certified Biosafety - See more at: http://www.thermoscientific.com/en/product/heraeus-multifuge-x1-centrifuge-series.html#sthash.JsGJZZUE.dpuf | Piece
| Autoclave-Chamber Capacity: 53 Litters, Temperature setting range: 105°C to 132 °C,Temperature control: Digital, Microprocessor controlled, Pressure display range: 0 to 400 kPa, Maximum operating pressure: 216kPa, Operating pressure range: 0-186kPa, Sterilization heat source: 2.0kW electric heater, Dimensions of main unit (mm): 490 W x 560D x 1090H (Height from floor to control panel:875) , Chamber diameter: 325 x 740mm, Net Weight: 80 kg, Power input: 20A (120V), 10A (220V), 9A (230/240V) | Piece
| Freezer (-80 deg); uprightfreezer, thermo Scientific. Model: 900 series- 86° C | Piece
| 2-8 Refrigerator-Laboratory refrigerator of Capacity 1500L, temperature range 2℃-8℃, accuracy 0.1℃, Audio and visual alarm for: Height or low temperature; System & Sensor failure; Door ajar(except BXC-V50M); Power failure alarm , Auto defrost,Rigid polyurethane insulation material, Height efficiency air-cooled condenser, made of copper , Three-ply transparent glass door with lock ,Stealing joint strip with magnetic force ,Power Supply 220V±10%, 50Hz/60Hz with adequate lighting, Stainless steel Material (Internal Cold-rolled steel with anti-bacteria powder coating, External Cold-rolled steel with anti-bacteria powder coating ) | Piece
Food Lab
| Analytical balance: Capacity up to 320g and Readability 0.01 mg / 0.1mg. MonoBloc HighSpeed cell, full auto internal adjustment. Touchscreen, weighing-in guide, test management. Dismentalable weighing chamber, dishwasher proof. | Piece
| Spectophotometer: Measuring wavelength range:220 to 800 nm. Spectrum band width:3 nm. Pathlength: 0.2 mm, 0.7 mm (manual selection). Light source: Xenon flash lamp. Monochrometer: Holographic grating. Detector: Photo diode array. Auto wiping function. Auto sample mount function. Quantitation rate:Pathlength 0.2 mm, 1 to 75 OD (50 to 3,700 ng / µL). | Piece
| Vortex: Speed: 3000rpm Tube Capacity Power:Voltage 230V Platform Material Foam Hertz 50/60Hz Amperage 0.3A | Piece
| Stomacher | Piece
| Double-door refrigerator: Transparent Double door high performance refrigerator with temperature Control Microprocessor. Min. capacity (volume) 1200L. Must be frost free and CFC free. Temperature Range (Metric) 1° to 8°C Power:Voltage 250V/50Hz. | Piece
| Deep Freezer (-86): Temperature Range (Metric) -50° to -86°C. Capacity :28 cu. ft. Min. capacity (volume) 700-800L. Power Voltage: 230V Amperage 12A. | Piece
| Deep Freezer (-40): Temperature Range (Metric) -20° to -40°C. Capacity :17.3 cu. ft. Min. capacity (volume)250-300L with CO2 & Liquid Nitrogen option. Power Voltage: 230V Amperage 12A. | Piece
| Centrifuge with refrigerator: Max. Speed 15200rpm Capacity 1.6L (4 x 400) Refrigerated No Max. RCF 25830 x G Hertz 50/60Hz Voltage 230V Control Microprocessor Display Digital Height (English) Exterior 14.2 in. Height (English) Open Cover 34.3 in. Height (Metric) Exterior 36cm Height (Metric) Open Cover 87cm Length (English) Exterior 23.8 lb. Length (Metric) Exterior 60.5cm Memory 99 programs (5 with direct access key) Motor Type Direct, brushless induction, low-profile Net Weight (English) 126.8 lb. Net Weight (Metric) 57.5kg Timer 99 hrs., 59 min. plus HOLD Width (English) Exterior 17.3 in. Width (Metric) Exterior 44cm Certifications/Compliance IVD compliant; IEC 61010-1, IEC 61010-2-020, IEC 61010-2-101; 230V models meet EN 292, EN 61326, EN 55011B. UL, CE, CSA, Certified Biosafety - See more at: http://www.thermoscientific.com/en/product/heraeus-multifuge-x1-centrifuge-series.html#sthash.JsGJZZUE.dpuf | Piece
| Waterbath: Microprocessor controlled with integrated autodiagnostic system with fault indication, digital timer from 1 min. up to 999 hours for:ON, DELAYED ON, HOLD or HOLD set temperature with HOLDING time facility.Resolution of display and setting accuracy: 0,1 °C below 99,9 °C, 1 °C above 100 °C. Material 1.4301 (ASTM 304), corrosion resistant.Temperature range: Min. 5 °C above ambient up to +200 °C heating capability. Dimension(mm): 600 x 350 x 230 mm approx. or 40-45L. Power: 230-250V/50Hz. | Piece
Note on B2074: dell: add near gelfoam

## Anesthesia
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
1 | Anesthesia Department
1.1 | Equipments
1.1.1 | Anesthesia Machine: Gas delivery: 3 gas version- Nitrous oxide, oxygen and air. Integrated safety functions: The Oxygen-Ratio Controller should guarantee a minimun oxygen in the fresh-gas flow. N20 cut off mechansim should ensure that no N2O is delivered in case of an oxygen failure. Audible oxygen supply failure alarm with oxygen flush. Pin index trolley: 2 yokes for oxygen and 2 yokes for N2O and 2 pressure gauges for oxygen and 2 for N2O: 2/3 vaporizer confirguation, Tool-free installation selectatic manifold interlock and isolates vaporizers. Number of position:2 or 3, Vaporizers: Halothane, Isoflurane or Sevoflurane. Ventilator: electronically controlled and driven/ pneumatically controlled and driven. Ventilator operating modes- volume control, pressure control with (SIMV/pressure support ,ASB) Tidal volume range: 20 to 1500 mL: One motion from mechanical to manual mode. Ventilator screen with parameter display with continuous monitoring of inspiratory oxygen, breathing frequency, airway pressure, minute volume and tidal volume. Gas supply: Pipeline connection and cylinder input. Power supply: 220 -240V, 50 HZ. Suitable UPS, Power cord with round 3 pin top. | Set | √ | √
1.1.2 | Ventlator: (ICU): Modes of ventilation: Assist/Control (A/C), SIMV, BIPAP, Spontaneous: Breath types: VCV in A/C or SIMV, PCV in A/C or SIMV, PSV SPONT or SIMV, include facility for respiratory mechanic (example... to measure negative inspiratory force), Mandatory Settings: Patient range: adult and pediatric, Ventilation frequency: 2 to 80 bpm. Inspiration time: 0.2 - 10s, Tidal volume: 40 - 2000 ml, Inspiratory flow: 0 - 180 L/min, Inspiratory pressure: 5- 100 cmH2O, PEEP: 0 - 35 cmH2O, Plateau: 0 to 2 seconds, Flow pattern: Square or decelerating ramp, Oxygen concentration: 21% - 100%, Trigger sensitivity: 1 to 20 L/min, Operation-Set Alarm Limits: Heart rate, low insp. pressure, high pressure, low minute volume, high tidal volume, low tidal volume, airway pressure, apnoea alarm, Ventilator Status Indicator: Alarm, caution, normal, VENT INOP, safety valve open, On AC/battery charging, On internal battery, Manual INSP: delivers one mandatory breath, Operating data: Mainpower connection-100-240V, internal battery- approx 60 minutes, Standard accessories: collectors vials, ventilator breathing circuits with water traps:- adult and pediatric, Electronic humidifier with humidification chamber, standard cart, bacterial filters, test lung, and flex arm. Power supply: 220-240V. Suitable UPS, Power cord with round 3 pin top. should have one set of test lung. | Set | √ | √
1.1.3 | Portable ventilator: Mode; Assist/Control, SIMV, Spon. Breath types: VCV, PCV, PSV. Mandatory setting: Respiratory rate- 1 to 80 breath/min, Tidal volume- 50 to 2000 ml, Inspiratory time- 0.2 to 5 seconds, Spontaneous setting- Pressure support: 0 to 70 cmH2O, Sensitivity: 3 to 25 lpm, Common setting: PEEP/CPAP: 0 to 20 cmH2O, Trigger sensitivity: 3 to 25lpm, Oxygen %: 21% to 100%, Operator set alarm limits: Alarms classified as high and medium priority with audible and visual presentations, should have option to monitor respiratory mechanics (negative inspiratiry force, etc) Internal battery: operating time of more than 4 hours and rechargable, Power line 220 - 240 V, 50 Hz, Power cord with round 3 pin top. Accessories: Resuable adult and pediatric circuit and bacterial filters. | Set | √ | √
1.1.4 | Defibrillator machine: ECG Monitoring: 3-lead ECG cable, paddles and 10nos of pacing pads. Screen: High resolution display- heart rate, ECG, SPO2, NIBP, Optional ET CO2measuring, alarm on/off, AED functions and prompts, defibrillator test function, self test functions, pacer functions and delivered energy. Defibrillator: Biphasic waveform. Energy selection-upto 200 joules. Selected using control on sternum paddle or device front pannel. Quick charge time. Monitor display both selected and delivered energy. Charge controls on paddles and on device front pannel. Adult paddles slides off to expose pediatric paddles. AED function- auto-analyze and auto energy level selection, screen prompts and voice prompts. Recorder: Manual or automatic printing modes, on/off control on front pannel and on paddle. With the following modes: NIBP, EtCO2, IBP, Pacemaker, Temperature. Rechargeable battery. Power cord with three round pin plug top. | Set | √ | √ | √
1.1.5 | AED-Automatic External Defibrillator : Fully automatic, L+B11CD screen, Electrodes pads for adult and pediatric patients, Battery powered with alternate AC adapter, Compact unit. Voice prompt instruction- check airway, apply pad, etc. (should supply 10nos of pacing pads each with the AED) Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √ | √ | √ | √
1.1.6 | Intubating Fiberoptic with light source (Adult): Working length of 55 cm to 65 cm flexible probe, 1.4 mm multipurpose channel, 90' field of vision, 2 - 50 mm focal range, 260' tip deflection and 25 mm angulation, and plus to minus diopter adjustment. Accessories: Cleaning brush, light source and adapters & Case. Power supply: 220-240V. Power cord with round 3 pin top. | Set | √
1.1.7 | Intubating Fiberoptic with light source (Pediatric) : Flexible probe with multipurpose channel, 90' field of vision, 2 - 50 mm focal range, 260' tip deflection and 25 mm angulation, and plus to minus diopter adjustment. Accessories: cleaning brush, adapters, light source & case. Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √
1.1.8 | Intubating Fiberoptic with light source (neonate) : Flexible probe with multipurpose channel, 90' field of vision, 2 - 50 mm focal range, 260' tip deflection and 25 mm angulation, and plus to minus diopter adjustment. Accessories: cleaning brush, adapters, light source & case (with AMBUSCOPE FLEXIBLE INTUBATING SCOPE MONITOR). Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √
1.1.9 | Difficult Airway Intubation Aid: (main unit / monitor) aScope 3 system for difficult intubation: should have (1) high resolution reusable monitor, with inbuilt rechargeble battery. Power supply to power the set as well as recharge the battery. (2) 3 size ascope for Adult, Pediatric, and infant, all with connection for monitor, control lever for easy manupulation of the tip, suction port, working channel, etc. |  | √
1.1.10 | scope for Ascope 3 system: (adult) Sterile single use flexible scopes in adult sizes compatible with reusable monitor. Easy to use and light weight. Suitable for intubation, airway inspection, placing and checking double lumen tubes and management of secretions, BAL procedures. | Piece | √
1.1.11 | scope for Ascope 3 system: (pediatric) Sterile single use flexible scopes in pediatrict sizes compatible with reusable monitor. Easy to use and light weight. Suitable for intubation, airway inspection, placing and checking double lumen tubes and bronchial blockers | Piece | √
1.1.12 | Infusion Pump: Pump type: Peristaltic finger type, Control method: Drip rate control, Required infusion set: Standard PVC IV set with 3.4 - 4.5 mm OD tubing, Flow rate range: 1 - 450 ml/h, Purge rate: more than 450 ml/h, Alarm: Door open, Occlusion, Low battery, Internal malfunction, Air in line, Infusion complete, Empty container, Reminder, abnormal drips. Mounting: Pole or Bed frame. Power supply: 220 - 240V =/- 10% 50 Hz.power cord with round 3 pin top | Set | √ | √
1.1.13 | Syringe Pump: Flow rage (ml/h), 0.1 - 200 ml/h 10 ml syringe, 0.1 - 300 ml/h, 5 ml to 50 ml syringe, (0.1 ml/h increments), Rapid flow rate 200 - 900 ml/h depending on size of syringe, Accuracy: Within +/-1 to 2%, Alarm: occlusion, near empty, end, system malfunction, syringe disengage, reminder, low battery, Power source: 240 V, 50Hz and power cord with round 3 pin top. Built in battery, Mounting: Pole or bed frame. | Set | √ | √ | √ | √
1.1.14 | Patient warming system with drape warming blankets : Fast warm up, Temperature range Low 32', medium 38', and high 40' - 44'C, Multiple mounting options: cart, bedrail, IV pole or floor, Drape warming blankets: for upper and lower part of body (adult) & pediatric warming blanket ( set of 10 blankets each). AC power 220 - 240V, 50 Hz, Hose 5 ft. with nozzle clip. Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √ | √
1.1.15 | Blood/Fluid Warming Unit: The warming device should apply advanced heat technology to obtain the right fluid temperature.Should have a wide range of temperature for all routine clinical situations.The device should be able to warm blood/fluid at the rate of 125ml/minute to 300m/minute. Delivers normothermic fluids at flow rates from 15ml/min to 250ml/minute. Wide range of temperature-from 28'C to 42'C. Pole mounting and connects to standard electrical (220V - 240V), Easy maintenance and self calibrating, Accessories as required. Power supply: 220-240V. Power cord with round 3 pin top. | Set | √ | √ | √
1.1.16 | Patient monitor for OT: 15 -17 inch color TFT with maximum 8 waveforms. Standard configuration- NIBP, SPO2, 3/5 lead ECG, Pulse rate, Temperature, dual IBP, resp. & ETCO2 & sidestream, multi-gas/O2 measuring. Large font display mode.TOUCH SCREEN. Comprehensive arrhythmia review. Adjustable alarm sound. Battery and recorder. Power AC 220 - 240V. Patient range: adult, pediatric and neonate. Power cord with round 3 pin top. Accessories to be included: 10 numbers of CO2 sampling lines with water trap, 10 numbers of IBP transducers, Extra SPO2 probe & BP cuff for adult, child and neonate | Set | √ | √ | √
1.1.17 | Patient monitor for PACU & ICU: 10 - 12 inch color TFT display with 8 waveforms. Standard Configuration: NIBP, SPO2, 3/5 lead ECG, EtCO2, IBP, temprature and pulse rate, Arrhythmia review. Adjustable alarm sound. Patient range:adult, pediatric and neonate. TOUCH SCREEN. Power AC 220 - 240V. Battery and recorder. Power cord with round 3 pin top. Accessories to be included: 10 numbers of CO2 sampling lines with water trap,10 numbers of IBP transducers, Extra SPO2 probe & BP cuff for adult, child and neonate (with wall mount) | Set | √ | √
1.1.18 | Patient monitor for Wards : 8-10" color TFT display with 4 - 6 waveforms. Standard configuration; ECG, RESP. Dual-TEMP. NIBP, SPO2, Pulse rate. Adjustable alarm sound. Patient range; adult, pediatric and neonate. Power AC 220 - 240V. Battery and recorder. Accessories to be inlcuded :SPO2 probe & BP cuff for adult, child and neonate. | Set | √ | √ | √
1.1.19 | Vital Sign Monitor: Standard Configuration- NIBP and SPO2, adult and pediatric sensor probe. Power supply: AC 220-240V. Power cord with round 3 pin top | Set | √ | √ | √ | √
1.1.20 | Pulse Oximeter: SPO2 measurement range of 1% - 100%, Accuracy saturation (SpO2 =/- 1 SD) inbuilt rechargeable back up battery, High quality LCD graphical screen, Plethysmographic waveform or magnified view with pulse amplified indicator,visual indicators: Pulse search, audible alarm, silenced or off, artifact indicator, low battery and battery charging, sensor off, Alarm: audible and visual alarm for high and low saturation and pulse rate, audible and visual warning indicator for low battery, display saturation and pulse rate alarm limits. Standard accessories: Extension cable, sensor probe for adult, pediatric and neonate, power cord, operator manual.Variable pitch beep tone, Selection of adult and neonate alarm. Power supply: AC 220-240V, 50 Hz. Power cord with round 3 pin top | Set | √ | √ | √
1.1.21 | GlideScope-Video Laryngoscopes: Features anti-fogging mechanism, integrated high resolution camera and color monitor & video output/recording features, Video laryngoscope with baton for 1. adult, 2. pediatric & 3. neonate size. Power AC 220V- 240V with power power cord of round three pin top, with inbuilt rechargeble battery backup | Set | √ | √
1.1.22 | Digital Ultrasonic Diagnostic Imaging System and for nerve block (Laptop Type): High resolution colour monitor 10 - 15 inch, High resolutions imaging from 2 to 18 MHz, upto 2-4 hours of plug free imaging, Transducer button operation: high speed linear array 6-18 MHz, curved array 2- 6 MHZ. Hockey stick transducer probe, frequency 5-10 MHz. Power supply: 220-240V. Power cord with round 3 pin top | Set | √
1.1.23 | Manual Jet Ventilator: with regulator and Gauge with jet ventilation catheter, 100 cm connecting tube with leur lock. | Set | √
1.1.24 | ICU Bed with mattress: Split safety sides with built-in patient and nurse control, Folding safety sides with control, Electrically operated back rest, knee-break and height adjustment, Electrically operated calf section adjustment, Electrically operated trendelenberg/reverse trendelenberg tilt, Auto-CPR button, Battery backup, Built-in bed extension with mattress retainer, Removal head and foot panels, Linked braking/stering system with four braking castors, Over all length and width 230 - 105 cm approximately, Power supply: 220-240V. Power cord with round 3 pin top | set | √ | √
1.1.25 | Arterial blood gas machine (ABG) : Advanced Analytics, Fast Stat Results, All Liquid Calibration Cartidges, Automates Quality Control, Automatic Quality Analysis, Automatic Data Storage and Reporting, Small Sample Volume, Easy, Safe Capillary or Syringe Use. Standard Measured Test, Calculated Results, Special Calculated Test, Measurement Range. including facility for calculation of anion gap and printer. Calibration: User-selectable, fully automatic two-point calibration, user-selectable single point calibration, manual calibration initiated at any time. Power: 220 - 240V, 50 Hz, Power cord of round 3 pin top. Data port. | Set | √ | √
1.1.26 | Patient Controlled Analgesia Machine: Programmable, work on a wide ranges of syringes/infusion pump type, comprehensive alarm packages, with various modes of infusion, bolus with patient control extension with button,, continuous infusion, lock out, electronic security key, inbuilt rechargeable battery back facility, Power supply: 220-240V, 50 Hz, Power cord with round 3 pin top and suitable voltage stabilizer. | Piece/set | √
1.1.27 | Laryngoscope: Standard Macintosh blade- 4 blades Mac 1,2,3,4 with battery, handle type C. (stainless steel blades) with Fibreoptic light bulbs | Set | √ | √ | √ | √
1.1.28 | Laryngoscope :Standard Millers blade- MIL 1, 2, 3, 4 with battery, handle type C. (stainless steel blades) with Fibreoptic light bulbs | Set | √ | √ | √ | √
1.1.29 | Laryngoscope :Standard Millers blade- MIL 00, 0, 1 with battery, handle type AA. (stainless steel blades) with Fibreoptic light bulbs | Set | √ | √ | √ | √
1.1.30 | McCoy FO blade : Flexible blade with integrated F.O /LED light conductor, with bendable distal end, blade size adult,pediatric and neonate, stainless steel blade and battery handle type C. | Set | √
1.1.31 | Airway Management Stimulator :Head and Neck with Thyroid cartilage and bronchial tree fixed to base, exchangeable skin and thyroid cartilage, carina with ballons/bronchial tree | Set | √
1.1.32 | Manniquin: CPR Torso with light Indicator: Papable anatomical landmarks, Adult CPR training with optional child spring, Fully articulated head, neck and jaw, Palpable carotid pulse, Adult/child mode, Proper compression depth, proper ventilation volume, improper hand placement indicated by yellow, green and red light. Accessories; Mouth and nose pieces, child spring, disposable lower airway, light controller with batteries, disposable tracheal airway washable suit and carrying case.Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √
1.1.33 | CPR Torso Manniquin: Containing adult, child and neonate torso, mouth protection and carrying bag. | Set | √
1.1.34 | ECG Simulator with Arrhythmia | Piece | √
1.1.35 | Recovery Trolley: Light weight trolley, Swing type side rail protector, Provision for IV stand, Oxygen cylinder cage, Adjustable Back rest lifting 0 - 80 degrees, Adjustable height and facility to put in trendelenberg position. Frame made of stainless steel, Four 15 cm antistatic castor, lockable wheel, Foam mattress covered with synthetic leather. | Set | √ | √ | √
1.1.36 | Crash Trolley: Tubular framework, 6 removable bins, two lockable storage units with three drawers each, each cylinder carrier, stainless steel saline rod, mounted on 125 mm DB castor, two brake provision. | Set | √ | √ | √
1.1.37 | Peripheral Nerve Stimulator: Pulse amplitude adjustable from 2 to 100 mA. Modes: TWT, TET, TOF, DBS, PTC. Battery operated. | Set | √
1.1.38 | Cuff Pressure Gauge : 50 mm scale, release valve to allow sensitive adjustment of the desired cuff pressure, scale with cuff range from 20 to 30 cm H20. | Piece | √
1.1.39 | Patient transfer system: For patient transfer between trolley, OT table, ICU bed etc : Easy move or patslide, full body length, compatible with 200 kg patient, with smooth edges. | Piece | √ | √
1.1.40 | Wright's Respirometer: Tidal volume measurement with minute volume facility. Small inner dial records 0 - 100 liters, large outer dial measures 0 - 1 liter, Connector: inlet/outlet, 22 mm male BS/ISO taper. | Piece | √
1.1.41 | Sequential compression device (SCD) for prevention of DVT in ICU. Sleeves for the device: Knee length and thigh length ( Sizes: Small, Medium & Large) (include atleast 10 of each sleeves in the quote) | Set | √
1.1.42 | Fluid warming cabinet for warming IV and irrigation fluids :upto 4.2.cu ft.capacity.compact size for freedom of placement within the facility .customable depth,door material,premium features,base and installation. | Piece | √
1.1.43 | BiPAP machine: - portable- Autoset algorithym,- With humidifier, IPAP: 4-30 ppressure cm H2O, EPAP: 4-30 preaaure cm H2O,- With at least 3 modes: and port for oxygent inlet . a)CPAP, b) Bipap, c) spontaneous mode Should have three sets of following 1. CPAP/BiPAP mask, full face, adult size, reusable, with head strap 2. CPAP/ BiPAP nasal mask, adult size, reusable with head strap | set | √
1.1.44 | Flexible bronchoscopy/fibreoptic bronchoscope: with xenon light source, forceps for bronchoscopic biopsy, channel for biopsy, and suction, biopsy channel diameter of 2.2mm, suction port, and cleaning brush, viewing angulation range of 130 /180o -video bronchoscope would be preferable | set | √
1.1.45 | Hospital Emergency Crash cart: 1. ABS trolley body 2. Plastic-steel columns 3. Adjustable height IV pole attached to unit 4. Removable Dust basket 5. Utility container 6. File bag 7. Defibrillator shelf, power outlet & hooks 8. Sliding side shelf 9. Stainless steel guard rail 10.Concave top covered transparent soft plastic glass 11.Centralized lock 12.Two small & Three big drawer with adjustable partition which can be well organized inside the each drawer. 13.Drawers repositioned eaily 14.Luxurious noiseless castersinner with partitions can be well organized 15.ABS basket & drawer frame can be repositoned easily 16.Luxuriouos casters wheels which enables easy trolleying / movement, with two with brakes 17. CPR board 18. Oxygen tank holder 19. Needle disposal bottle/bin container | set | √
1.2 | Consumables
1.2.1 | Oropharyngeal Airway : Size: 00 | Piece | √ | √ | √ | √
1.2.2 | Oropharyngeal Airway : Size: 0 | Piece | √ | √ | √ | √
1.2.3 | Oropharyngeal Airway : Size: 1 | Piece | √ | √ | √ | √
1.2.4 | Oropharyngeal Airway : Size: 2 | Piece | √ | √ | √ | √
1.2.5 | Oropharyngeal Airway : Size: 3 | Piece | √ | √ | √ | √
1.2.6 | Oropharyngeal Airway : Size: 4 | Piece | √ | √ | √ | √
1.2.7 | Sleeves for sequential compression device. Knee length | Piece | √ | √ | √ | √
1.2.8 | sleeves for sequential compression device. Thigh length | Piece | √ | √ | √ | √
1.3 | Face mask
1.3.1 | Face mask transparent disposible : Size 00 (inflatable cushion) | Piece | √ | √ | √
1.3.2 | Face mask transparent disposible : Size 0 ( inflatable cushion) | Piece | √ | √ | √
1.3.3 | Face mask transparent disposible : Size 1 (Resuable inflatable cushion) | Piece | √ | √ | √
1.3.4 | Face mask transparent disposible : Size 2 (Resuable inflatable cushion) | Piece | √ | √ | √
1.3.5 | Face mask transparent disposible : Size 3 (Resuable inflatable cushion) | Piece | √ | √ | √
1.3.6 | Face mask transparent disposible : Size 4 ( Resuable inflatable cushion) | Piece | √ | √ | √
1.3.7 | BiPAP masks: full face masks, soft cushion with proper seal, head strap, large and medium size compatible with BiPAP machine | Piece | √ | √ | √
1.3.8 | CPAP Respiration Set : CPAP Mask, Harness, PEEP Valve adjustable from 5 - 20cmH2O, and adapter for 22mm tube connection, should be able to connect to ventillator circuit. | Set | √
1.4 | Endotracheal tube cuffed: Oral/Nasal:(O/N)
1.4.1 | Endotracheal tube : PVC, high volume,low pressure (Micro) cuff size I.D 2.5mm (O/N) | Piece | √ | √ | √
1.4.2 | Endotracheal tube :PVC, high volume,low pressure (Micro) cuff size I.D 3 mm (O/N) | Piece | √ | √ | √
1.4.3 | Endotracheal tube :PVC, high volume, low pressure (Micro) cuff size- 3.5mm (O/N) | Piece | √ | √ | √
1.4.4 | Endotracheal tube :PVC, high volume, low pressure (Micro) cuff size- 4 mm (O/N) | Piece | √ | √ | √
1.4.5 | Endotracheal tube :PVC, high volume, low pressure cuff size- 4.5mm (O/N) | Piece | √ | √ | √
1.4.6 | Endotracheal tube :PVC, high volume, low pressure cuff size- 5 mm (O/N) | Piece | √ | √ | √
1.4.7 | Endotracheal tube :PVC, high volume, low pressure cuff size- 5.5 mm (O/N) | Piece | √ | √ | √
1.4.8 | Endotracheal tube :PVC, high volume low pressure cuff - size I.D 6mm (O/N) | Piece | √ | √ | √
1.4.9 | Endotracheal tube :PVC, high volume low pressure cuff - size I.D 6.5mm (O/N) | Piece | √ | √ | √
1.4.10 | Endotracheal tube :PVC, high volume low pressure cuff - size I.D 7mm (O/N) | Piece | √ | √ | √
1.4.11 | Endotracheal tube :PVC, high volume low pressure cuff - size I.D 7.5mm (O/N) | Piece | √ | √ | √
1.4.12 | Endotracheal tube :PVC, high volume low pressure cuff - size I.D 8mm (O/N) | Piece | √ | √ | √
1.4.13 | Endotracheal tube :PVC, high volume low pressure cuff - size I.D 8.5mm (O/N) | Piece | √ | √ | √
1.5 | Endotracheal tube non-cuffed: Oral/Nasal
1.5.1 | Endotracheal tube: PVC, non-cuffed-size I.D 2mm | Piece | √ | √ | √
1.5.2 | Endotracheal tube: PVC, non-cuffed-size I.D 2.5mm | Piece | √ | √ | √
1.5.3 | Endotracheal tube: PVC, non-cuffed-size I.D 3mm | Piece | √ | √ | √
1.5.4 | Endotracheal tube: PVC, non-cuffed-size I.D 3.5mm | Piece | √ | √ | √
1.5.5 | Endotracheal tube: PVC, non-cuffed-size I.D 4mm | Piece | √ | √ | √
1.5.6 | Endotracheal tube: PVC, non-cuffed-size I.D 4.5mm | Piece | √ | √ | √
1.5.7 | Endotracheal tube: PVC, non-cuffed-size I.D 5mm | Piece | √ | √ | √
1.5.8 | Endotracheal tube: PVC, non-cuffed-size I.D 5.5mm | Piece | √ | √ | √
1.6 | RAE tube Cuffed PVC Oral
1.6.1 | RAE tube : PVC, micro cuffed oral- size ID 3mm | Piece | √ | √
1.6.2 | RAE tube : PVC, micro cuffed oral- size ID 3.5mm | Piece | √ | √
1.6.3 | RAE tube : PVC, cuffed oral- size ID 4mm | Piece | √ | √
1.6.4 | RAE tube : PVC, cuffed oral - ID 4.5mm | Piece | √ | √
1.6.5 | RAE tube : PVC, cuffed oral- size ID 5mm | Piece | √ | √
1.6.6 | RAE tube : PVC, cuffed oral- size ID 5.5mm | Piece | √ | √
1.6.7 | RAE tube : PVC, cuffed oral- size ID 6mm | Piece | √ | √
1.6.8 | RAE tube : PVC, cuffed oral- size ID 6.5mm | Piece | √ | √
1.6.9 | RAE tube : PVC, cuffed oral- size ID 7mm | Piece | √ | √
1.6.10 | RAE tube : PVC, cuffed oral- size ID 7.5mm | Piece | √ | √
1.6.11 | RAE tube : PVC, cuffed oral- size ID 8mm | Piece | √ | √
1.7 | RAE Tube Uncuffed Oral
1.7.1 | RAE tube : PVC, non-cuffed oral- size ID 2.5mm | Piece | √ | √
1.7.2 | RAE tube : PVC, non-cuffed oral- size ID 3mm | Piece | √ | √
1.7.3 | RAE tube : PVC, non-cuffed oral- size ID 3.5mm | Piece | √ | √
1.7.4 | RAE tube : PVC, non-cuffed oral- size ID 4mm | Piece | √ | √
1.7.5 | RAE tube : PVC, non-cuffed oral- size ID 4.5mm | Piece | √ | √
1.7.6 | RAE tube : PVC, non-cuffed oral- size ID 5mm | Piece | √ | √
1.7.7 | RAE tube : PVC, non-cuffed oral- size ID 5.5mm | Piece | √ | √
1.8 | RAE tube Cuffed PVC Nasal
1.8.1 | RAE tube : PVC, cuffed nasal- size ID 3.5mm | Piece | √ | √
1.8.2 | RAE tube : PVC, cuffed nasal- size ID 4mm | Piece | √ | √
1.8.3 | RAE tube : PVC, cuffed nasal- size ID 4.5mm | Piece | √ | √
1.8.4 | RAE tube : PVC, cuffed nasal- size ID 5mm | Piece | √ | √
1.8.5 | RAE tube : PVC, cuffed nasal- size ID 5.5mm | Piece | √ | √
1.8.6 | RAE tube : PVC, cuffed nasal- size ID 6mm | Piece | √ | √
1.8.7 | RAE tube : PVC, cuffed nasal- size ID 6.5mm | Piece | √ | √
1.8.8 | RAE tube : PVC, cuffed nasal- size ID 7mm | Piece | √ | √
1.9 | RAE Tube PVC, Uncuffed Nasal
1.9.1 | RAE tube : PVC, non-cuffed nasal- size ID 2.5mm | Piece | √ | √
1.9.2 | RAE tube : PVC, non-cuffed nasal- size ID 3mm | Piece | √ | √
1.9.3 | RAE tube : PVC, non-cuffed nasal- size ID 3.5mm | Piece | √ | √
1.9.4 | RAE tube : PVC, non-cuffed nasal-size ID 4mm | Piece | √ | √
1.9.5 | RAE tube : PVC, non-cuffed nasal- size ID 4.5mm | Piece | √ | √
1.9.6 | RAE tube : PVC, non-cuffed nasal- size ID 5mm | Piece | √ | √
1.9.7 | RAE tube : PVC, non-cuffed nasal- size ID 5.5mm | Piece | √ | √
1.9.8 | RAE tube : PVC, non-cuffed nasal- size ID 6mm | Piece | √ | √
1.9.9 | RAE tube : PVC, non-cuffed nasal- size ID 7mm | Piece | √ | √
1.1 | Double Lumen Endobronchial Tube
1.10.1 | Double Lumen Endobronchial tube set : Left sided- size- 26 Fr | Piece | √
1.10.2 | Double Lumen Endobronchial tube set : Right sided- size 26 Fr | Piece | √
1.10.3 | Double Lumen Endobronchial tube set : Left sided- size 35 Fr | Piece | √
1.10.4 | Double Lumen Endobronchial tube set : Right sided- size 35 Fr | Piece | √
1.10.5 | Double Lumen Endobronchial tube set : Left sided size 37 Fr | Piece | √
1.10.6 | Double Lumen Endobronchial tube set : Right sidedsize 37 Fr | Piece | √
1.10.7 | Double Lumen Endobronchial tube set : Left sided - size 39 Fr | Piece | √
1.10.8 | Double Lumen Endobronchial tube set : Right sided- size 39 Fr | Piece | √
1.10.9 | Double Lumen Endobronchial tube set : Left sided - size 28 Fr | Piece | √
1.11 | Flexo-Metallic Endotracheal Tube cuffed
1.11.1 | Flexo-Metalic Endotracheal Tube : Cuffed- size - I.D 5mm | Piece | √ | √
1.11.2 | Flexo-Metalic Endotracheal Tube : Cuffed- size - I.D 5.5mm | Piece | √ | √
1.11.3 | Flexo-Metalic Endotracheal Tube : Cuffed- size - I.D 6mm | Piece | √ | √
1.11.4 | Flexo-Metallic Endotracheal Tube: Cuffed- size -I.D 6.5 mm | Piece | √ | √
1.11.5 | Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 7mm | Piece | √ | √
1.11.6 | Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 7.5mm | Piece | √ | √
1.11.7 | Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 8mm | Piece | √ | √
1.11.8 | Flexo-Metallic Endotracheal Tube: Cuffed- size - I.D 8.5mm | Piece | √ | √
1.12 | Microlaryngeal Tube
1.12.1 | Microlaryngeal tube : Size - I.D 5mm | Piece | √
1.12.2 | Microlaryngeal tube : Size - I.D 6mm | Piece | √
1.13 | Laryngeal Mask Airway: (LMA)
1.13.1 | Laryngeal mask airway : Size 1, reusable, (curved/classic) with apparture bar | Piece | √ | √
1.13.2 | Laryngeal mask airway : Size 1.5, reusable (curved/classic) with apparture bar | Piece | √ | √
1.13.3 | Laryngeal mask airway : Size 2, reusable (curved/classic) with apparture bar | Piece | √ | √
1.13.4 | Laryngeal mask airway : Size 2.5, reusable (curved/classic) with apparture bar | Piece | √ | √
1.13.5 | Laryngeal mask airway : Size 3, reusable (curved/classic) with apparture bar | Piece | √ | √
1.13.6 | Laryngeal mask airway : Size 4, reusable (curved/classic) with apparture bar | Piece | √ | √
1.13.7 | Laryngeal mask airway : Size 5, reusable (curved/classic) with apparture bar | Piece | √ | √
1.14 | Laryngeal Mask Airway, Flexible, Reinforced Tube
1.14.1 | Laryngeal mask airway : Flexible, reinforced tube, latex free, size 2 | Piece | √
1.14.2 | Laryngeal mask airway : Flexible, reinforced tube, latex free, size 2.5 | Piece | √
1.14.3 | Laryngeal mask airway : Flexible, reinforced tube, latex free, size 3 | Piece | √
1.14.4 | Laryngeal mask airway : Flexible, reinforced tube, latex free, size 4 | Piece | √
1.15 | Laryngeal Mask Airway: Proseal reusable
1.15.1 | Laryngeal Mask Airway: Proseal reusable: Size 2 | Piece | √
1.15.2 | Laryngeal Mask Airway: Proseal reusable: Size 2.5 | Piece | √
1.15.3 | Laryngeal Mask Airway: Proseal reusable: Size 3 | Piece | √
1.15.4 | Laryngeal Mask Airway: Proseal reusable: Size 4 | Piece | √
1.15.5 | Laryngeal Mask Airway: Proseal reusable: Size 5 | Piece | √
1.15.6 | Laryngeal Mask Airway: Proseal reusable: Size 1.5 | Piece | √
1.16 | Intubating LMA Airway
1.16.1 | Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 6, latex free. | Piece | √
1.16.2 | Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 6.5, latex free. | Piece | √
1.16.3 | Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 7, latex free. | Piece | √
1.16.4 | Intubating LMA airway : Size 3, with handle, epiglottic elevating bar, with tracheal tube size 7.5, latex free. | Piece | √
1.16.5 | Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 6, Latex free | Piece | √
1.16.6 | Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 6.5, Latex free | Piece | √
1.16.7 | Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 7, Latex free | Piece | √
1.16.8 | Intubating LMA airway : Size 4, with handle, epiglottic elevating bar. With tracheal tube size 7.5, Latex free | Piece | √
1.17 | Laryngeal Tube
1.17.1 | Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone) - reusable size 2 | Piece | √ | √
1.17.2 | Laryngeal Tube with Drain Tube: Color code syringe and one biteblock (silicone)- reusable size 2.5 | Piece | √ | √
1.17.3 | Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone)- Reusable Size 3 | Piece | √ | √
1.17.4 | Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone)- Reusable Size 4 | Piece | √ | √
1.17.5 | Laryngeal Tube with Drain Tube: Color coded syringe and one biteblock (silicone)- Reusable size 5 | Piece | √ | √
1.17.6 | Esophageal/Tracheal Double Lumen Airway (Combitube): Size: 37 ID | Piece | √ | √
1.18 | Emergency Cricothyrotomy set
1.18.1 | Emergency Cricothyrotomy set : ID 2mm (child) | Piece | √ | √
1.18.2 | Emergency Cricothyrotomy set : ID 4mm (adult) | Piece | √ | √
1.18.3 | Percutaneous dilatational tracheostomy set (Adult) | Piece | √ | √
1.18.4 | Percutaneous dilatational tracheostomy set (pediatric) | Piece | √ | √
1.18.5 | Percutaneous dilatational tracheostomy set (Infant) | Piece | √ | √
1.19 | Stylets
1.19.1 | Stylet : Size: OD Fr 6 (small) | Piece | √ | √ | √
1.19.2 | Stylet : Size: OD Fr 10 (medium) | Piece | √ | √ | √
1.19.3 | Stylet : Size: OD 14 (large) | Piece | √ | √ | √
1.19.4 | Gum elastic bougie (adult) | Piece | √ | √ | √
1.19.5 | Gum elastic bougie (child) | Piece | √ | √ | √
1.19.6 | Gum elastic bougie (neonate) | Piece | √ | √ | √
1.19.7 | Tube exchanger with markings and port for oxygenation adult | Piece | √ | √ | √
1.19.8 | Tube exchanger with markings and port for oxygenation pediatric | Piece | √ | √ | √
1.2 | Resuscitation Bag
1.20.1 | Reusable Silicone Resuscitation Bag : Adult 2000 ML, with 40 cm H2O pressure relief, Air Cushion Face Mask # 4, Reservoir Bag 2600 ML with Valve, 2 m O2 tubing. | Set | √ | √ | √ | √
1.20.2 | Reusable Silicone Resuscitation Bag : Child 500 ML, with 40 cm H2O pressure relief, Air Cushion Face Mask # 2, Reservoir bag 1800 ML with Valve, 2m 02 Tubing | Set | √ | √ | √ | √
1.20.3 | Reusable Silicone Resuscitation Bag : Infant 250 ML with 40 cm H20 pressure relief, Mask round # 1, Reservoir Bag 600 ML with valve, 2 m 02 tubing | Set | √ | √ | √ | √
1.21 | Oxygen Face Mask
1.21.1 | Oxygen face mask : With adjustable nose clip, elastic headstrap with reservoir bag with 2 meter tubing (Adult) | Set | √ | √ | √ | √
1.21.2 | Oxygen face mask : With adjustable nose clip, elastic headstrap with reservoir bag- pediatric, with 2 meter tubing (Pediatric) | Set | √ | √ | √ | √
1.21.3 | Oxygen nasal prong : With 2 meter tubing (Adult) | Set | √ | √ | √
1.21.4 | Oxygen nasal prong : With 2 meter tubing (Pediatric) | Set | √ | √ | √ | √
1.22 | Neubilizer Mask Kit
1.22.1 | Adult nebulizer mask kit : Facemask with adjustable nose clip, head strap, nebulizer chamber with 2 meter tubing. | Set | √ | √ | √ | √
1.22.2 | Pediatric nebulizer mask kit : Face mask with adjustable nose clip, head strap, nebulizer chamber with 2 meter tubing. | Set | √ | √ | √ | √
1.23 | Breathing Circuit
1.23.1 | 22 mm adult breathing circuit : Corrugated, flexible and lightweight tubing, 150 cm, 22F-22F ventilator connections, parallel Y piece ports, 90'angled elbow with capno port, 22M/15F patient connections, straight connector 22M-22M/19F, sterile (PP/PE material), should have extension tube for reservior bag connection | Set | √ | √ | √
1.23.2 | 15 mm pediatric breathing circuit : 100 cm, 22F-22F ventilator connections, Y piece port, 90' angled elbow with capno port, 22M/15F patient connections, sterile. (PP/PE material), should have extension tube for reservior bag connection | Set | √ | √ | √
1.23.3 | Jackson REES breathing system : Corrugated, flexible, lightweight, fresh gas corrugated tube 180 cm, vertical gas inlet, 22M/15F angled patient connections with capno port, 15M ventilator connection, tube 50 cm with 22M connector, 1/2 lt open latex free bag, not sterile. | Set | √ | √ | √
1.23.4 | Mapleson D - Bain's circuit : 22 mm Adult PE Anesthesia circuit, 198 cm, 22M/15F patient connection, 90' angled elbow at patient side, vertical gas inlet, APL valve with 2 liter bag, tube length 200 - 250 cm. Fresh gas inlet at patient end. | Set | √ | √
1.23.5 | T-piece for ICU and recovery | Piece | √ | √
1.24 | Breathing System Bacterial/Viral Filter for Ventilator
1.24.1 | Breathing System Filter for anesthesia circuit (Adult) : Filter material: Pleated hydrophobic membrane. Liquid borne bacterial/viral removal efficiency: 100%. Air borne bacterial/viral removal efficiency: >99.99%. Connections: Patient side 15 mm ID and 22 mm OD, breathing system side 15 mm OD 22 mm ID. Filter Volume and filter weight: Low filter volume and low weight. Construction: Transparent, non-conductive material. Port: CO2 monitoring port. (for OT only) | Piece | √ | √ | √
1.24.2 | Breathing System Filter for anesthesia circuit (pediatric) : Filter material: Pleated hydrophobic membrane. Liquid borne bacterial/viral removal efficiency: 100%. Air borne bacterial/viral removal efficiency: >99.99%. Connections: Patient side 15 mm ID and 22 mm OD, breathing system side 15 mm OD 22 mm ID. Filter Volume and filter weight: Low filter volume and low weight. Construction: Transparent, non-conductive material. Port: CO2 monitoring port. (for OT only) | Piece | √ | √
1.24.3 | HMEF Filter for icu - Adult : 100% bacterial Viral Efficiency, HME Membrane with CO2 port, 22M/15F ISO Connectors. (single use) ( for ICU only) | Piece | √ | √
1.24.4 | HMEF Filter for ICU - Pediatric : 100% bacterial Viral Efficiency, HME Membrane with CO2 port, 22M/15F ISO Connectors.(single use) ( for ICU only) | Piece | √ | √
1.24.5 | HMEF Filter - Neonate: 100% Bacterial Viral Efficiency, HME Membrane with CO2 port, 22M/15F connector.(single use) ( for ICU only) | Piece | √ | √
1.24.6 | 10 mm neonate breathing circuit: Disposable, lightweight,100cm, angled Y piece with Luer lock (capnoport).22mm connectors | Set | √ | √
1.24.7 | 22 mm adult breathing circuit :MRI COMPATIBLE Corrugated, flexible and lightweight tubing,300cm, 22F-22F ventilator connections, parallel Y piece ports, 90'angled elbow with capno port, 22M/15F patient connections, straight connector 22M-22M/19F, not sterile (PP/PE m | Set | √ | √
1.24.8 | 15 mm pediatric breathing circuit : MRI COMPATIBLE 300 cm, 22F-22F ventilator connections, Y piece port, 90' angled elbow with capno port, 22M/15F patient connections, not sterile. (PP/PE material) | Set | √ | √
1.24.9 | 10 mm neonate breathing circuit: MRI COMPATIBLE Disposable, lightweight,300cm, angled Y piece with Luer lock (capnoport) .22mm connectors | Set | √ | √
1.24 | Smooth bore PVC Circuit (ICU Ventilator)
1.24.1 | Smoothbore PVC circuits (with accessories) : 22 mm adult PVC ICU circuit, standard Y-Piece with port, 22M/15F patient connections, inspiratory line 160 cm with adult water trap, expiratory line 160 cm with adult water trap, extra limb 100 cm, 22Flex-22Flex ventilator connections, sterile. | Set | √ | √
1.24.2 | Smoothbore PVC circuits (with accessories) : 22 mm Pediatric PVC ICU circuit, standard Y-Piece with port, 22M/15F patient connections, inspiratory line 160 cm with pediatric water trap, expiratory line 160 cm with pediatric water trap, extra limb 80 cm, 22Flex-22Flex ventilator connections, sterile. | Set | √ | √
1.25 | Breathing Bag/Reservoir Bag
1.25.1 | Breathing Bags (Latex Free) : 0.5L breathing bag, latex free, 22F connector, not sterile | Piece | √ | √ | √
1.25.2 | Breathing Bags (Latex Free) : 1L breathing bag, latex free, 22F connector, not sterile | Piece | √ | √ | √
1.25.3 | Breathing Bags (Latex Free) : 2L breathing bag, latex free, 22F connector, not sterile | Piece | √ | √ | √
1.26 | Pressure Infusion Cuff
1.26.1 | Pressure Infusion Cuff: With hand inflator and manometer fixed to the cuff: wrap around, reusable. 500 ml | Set | √ | √ | √
1.26.2 | Pressure Infusion Cuff: With hand inflator and manometer fixed to the cuff: wrap around, reusable. 1000ml | Set | √ | √
1.26.3 | CPAP Respiration Set : CPAP Mask, Harness, PEEP Valve adjustable from 5 - 20cmH2O, and adapter for 22mm tube connection | Set | √
1.27 | Close Suction System, Replaceable Catheters, Endotracheal Length, MDI Port
1.27.1 | Size: 6Fr/Ch | Piece | √ | √
1.27.2 | Size: 8Fr/Ch | Piece | √ | √
1.27.3 | Size: 10Fr/Ch | Piece | √ | √
1.27.4 | Size: 12Fr/Ch | Piece | √ | √
1.27.5 | Size: 14Fr/Ch | Piece | √ | √
1.27.6 | Size: 16Fr/Ch | Piece | √
1.27.7 | Size: 18Fr/Ch | Piece | √
1.28 | Biteblock
1.28.1 | Biteblock : Autoclavable- size 1 length 2.5 cm | Piece | √ | √
1.28.2 | Biteblock : Autoclavable- size 3 length 3 cm | Piece | √ | √
1.28.3 | Biteblock : Autoclavable- size 5 length 3.5 cm | Piece | √ | √
1.29 | Central Venous Catheter
1.29.1 | CVP Line cava fix 45 cm | Set | √ | √
1.29.2 | CVP Line cava fix 70 cm | Set | √ | √
1.29.3 | CVP manometer : With connector (sterile) | Set | √ | √
1.29.4 | Triple lumen central venous cather sets : 7F, 18G x 150 - 200 mm catheter length. | Set | √ | √
1.29.5 | Triple lumen central venous cather sets : 5F, 20G x 100 - 150 mm catheter length. | Set | √ | √
1.29.6 | Double lumen central venous cather sets : 4F, 20G, 100- 150mm catheter length. | Set | √ | √
1.29.7 | Single lumen Central venous catheter : 3F, 22G, 100-150 mm catheter lendth | Set | √ | √
1.29.8 | Single lumen Central venous catheter : 2.5F, 22G, 60 - 100 mm catheter lendth | Set | √ | √
1.3 | Intra-arterial catheter
1.30.1 | Arterial Catheter: Radial artery set 2.5F, lenght 60 mm | Set | √
1.30.2 | Arterial Catheter: Radial artery set 3F, lenght 60 mm | Set | √
1.30.3 | Closed loop arterial blood sampling disposable transducer set for the invasive Monitoring of Haemodynamics Pressure. | Set | √
1.31 | Extension Tubes/Lines(sterile)
1.31.1 | Extension tube: With 3-way stop cock with leur lock: Length 50 cm | Piece | √ | √
1.31.2 | Extension tube: With 3-way stopcock with leur lock, length 100 cm (low volume) | Piece | √ | √
1.31.3 | Extension tube: With 3-way stopcock with leur lock, length 150 cm | Piece | √ | √
1.31.4 | Extension tube: Length 100-150mm Male and Female Leur Lock (Low volume for syringe pump) | Piece | √ | √
1.32 | Epidural Needles and Catheters
1.32.1 | Epidural needles : 18 G with epidural needle 16 G, with accessories,(loss of air resistence syringe10ml, epidural catheter, sterile pack set with plastic tray | Set | √ | √
1.32.2 | Epidural needles: 18G, with Loss of resistence to air syringe 10ml , With plastic tray in sterile pack | Piece | √ | √
1.32.3 | Epidural needles: 20G, sterile pack. | Piece | √ | √
1.32.4 | Pediatric epidural set 18/20G, 50 mm with 5mm marking with LOR syringe with plastic tray | Piece | √ | √
1.33 | Spinal Needles
1.33.1 | Spinal needle (pencil point) : Disposable, size - 25 G-sterile ,with introducer needle | Piece | √ | √ | √
1.33.2 | Spinal needle (pencil point) : Disposable, size - 26 G- sterile ,with introducer needle | Piece | √ | √ | √
1.33.3 | Spinal needle (pencil point) : Disposable, size - 27 G- sterile with introducer needle | Piece | √ | √ | √
1.33.4 | Spinal needle (Quinckes): size 22, length 15 cm- sterile (for nerve block) | Piece | √
1.33.5 | Spinal needle (Quinckes) : Disposable, size 22 G- sterile | Piece | √ | √ | √
1.33.6 | Spinal needle (Quinckes) : Disposable, size 23 G- sterile | Piece | √ | √ | √
1.33.7 | Spinal needle (Quinckes) : Disposable, size 25 G- sterile | Piece | √ | √ | √
1.33.8 | Spinal needle (Quinckes) : Disposable, size 27G- sterile | Piece | √ | √ | √
1.33.9 | Sodalime : Granule size 4 - 8 mesh or granules 1/4 - 1/8 inch in diameter, PH indicator turing from pink to white when exhausted. | Packet | √ | √ | √
1.34 | Accessories for Medical Cylinders
1.34.1 | Regulator with Flow Meter & Humidifier Chamber for Oxygen cylinder : B/D types of Cylinders | Piece | √ | √ | √ | √
1.34.2 | Regulator for Oxygen cylinder : D types of Cylinders for Anesthesia | Piece | √ | √ | √
1.34.3 | Regulator for Nitrous oxide cylinder : D type | Piece | √ | √ | √
1.34.4 | Key and Wrench for Nitrous oxide cylinder: B & D-Type | Piece | √ | √ | √
1.34.5 | Key and Wrench for Oxygen cylinder : B & D-Type | Piece | √ | √ | √ | √
1.34.6 | Oxygen key: For A type of oxygen cylinders | Piece | √ | √ | √
1.34.7 | High pressure oxygen tubing : For extended connection from Oxygen cylinder to Anesthesia Machine. 5 metre length | Piece | √ | √ | √
1.34.8 | High pressure nitrous oxide tubing : For extended connection from Nitrous oxide cylinder to Anesthesia Machine. 5 metre length | Piece | √ | √ | √
1.35 | Forceps
1.35.1 | Forceps MaGill - Adult | Piece | √ | √ | √
1.35.2 | Forceps MaGill - Pediatric | Piece | √ | √ | √
1.36 | Consumables for EPOC - Hand held (Portable)ABG (Arterial Blood Gas) Analyzer
1.36.1 | Cartridge (Test Card) | 50 test/pkt | √
1.36.2 | Control Fluid (Ampoules) | 10pcs/pkt | √
1.36.3 | Epoc ABG report printing paper roll | Roll/packet | √
1.37 | Sleeves for SCD
1.37.1 | Sleeves for sequential compression device. Knee length, medium size | Piece/pack | √ | √ | √ | √
1.37.2 | sleeves for sequential compression device. Thigh length, medium size | Piece/pack | √ | √ | √ | √
1.38 | Medical Gases (Gas Refilling)
1.38.1 | Hydrogen Refilling-D Type | Per Cylinder | √ | √ | √
1.38.2 | Carbon Dioxide Refilling- B Type | Per Cylinder | √ | √ | √
1.38.3 | Carbon Dioxide Refilling- D Type | Per Cylinder | √
1.38.4 | Oxygen Gas Refilling- AA Type | Per Cylinder | √ | √ | √
1.38.5 | Oxygen Gas Refilling- A Type | Per Cylinder | √ | √ | √
1.38.6 | Oxygen Gas Refilling- B Type | Per Cylinder | √ | √ | √ | √
1.38.7 | Oxygen Gas Refilling- C Type | Per Cylinder | √ | √ | √
1.38.8 | Oxygen Gas Refilling- D Type | Per Cylinder | √ | √ | √
1.38.9 | Nitrogen Gas Refilling- D Type | Per Cylinder | √ | √
1.38.10 | Nitrous Oxide Gas Refilling- D Type | Per Cylinder | √ | √ | √
1.38.11 | Nitrous Oxide Gas Refilling- A Type | Per Cylinder | √ | √ | √
1.38.12 | Nitrous Oxide Gas Refilling- B Type | Per Cylinder | √ | √ | √
1.38.13 | IV drip set compatible with infusion pump | Piece | √ | √
1.38.14 | BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 5cm (find out he length) | Piece | √ | √
1.38.15 | BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 8cm (find out he length) | Piece | √ | √
1.38.16 | BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 10cm | Piece | √ | √
1.38.17 | Hospital Emergency Crash cart: 1. ABS trolley body 2. Plastic-steel columns 3. Adjustable height IV pole attached to unit 4. Removable Dust basket 5. Utility container 6. File bag 7. Defibrillator shelf, power outlet & hooks 8. Sliding side shelf 9. Stainless steel guard rail 10.Concave top covered transparent soft plastic glass 11.Centralized lock 12.Two small & Three big drawer with adjustable partition which can be well organized inside the each drawer. 13.Drawers repositioned eaily 14.Luxurious noiseless castersinner with partitions can be well organized 15.ABS basket & drawer frame can be repositoned easily 16.Luxuriouos casters wheels which enables easy trolleying / movement, with two with brakes 17. CPR board 18. Oxygen tank holder 29. Needle disposal bottle | set | √
1.39 | Surfactant connector
1.39.1 | Surfactant connector attaches to ET 3/3.5 with side port | Piece | √ | √
1.4 | Mucosal atomizer device
1.40.1 | Laryngo-tracheal mucosal atomizer device without syringe | Piece | √ | √
1.4 | Pressure monitoring lines
1.40.1 | Non-expanding PMO line 200cm with luer lock
1.4 | Connectors
1.40.1 | Male to male luer lock connector
1.40.2 | Male luer lock to female connector
1.40.3 | Disposable straight connector 22F-22F
1.40.4 | Disposable straight connector 22M-22F/15F
1.40.5 | Disposable straight connector 22M/15F-22M/15F
1.41 | Forceps
1.41.1 | Forceps MaGill - Adult | Piece | √ | √ | √
1.41.2 | Forceps MaGill - Pediatric | Piece | √ | √ | √
1.41.3 | Capnography (etCO2 ) Sampling Line with dry water trap (compitable with mindray monitor) standard 200 tom400 cm line length, with leur lock male type at both end, single patient use, donot contain natural rubber latex, intended to be used in intubated, or face mask, adaptor dead space is <6.6cc | Piece | √ | √ | √
1.42 | Consumables for EPOC - Hand held (Portable) ABG (Arterial Blood Gas) Analyzer
1.42.1 | Cartridge (Test Card) | 50 test/pkt | √
1.42.2 | Control Fluid (Ampoules) | 10pcs/pkt | √
1.43 | Medical Gases (Gas Refilling)
1.43.1 | Hydrogen Refilling-D Type | Per Cylinder | √ | √ | √
1.43.2 | Carbon Dioxide Refilling- B Type | Per Cylinder | √ | √ | √
1.43.3 | Carbon Dioxide Refilling- D Type | Per Cylinder | √
1.43.4 | Oxygen Gas Refilling- AA Type | Per Cylinder | √ | √ | √
1.43.5 | Oxygen Gas Refilling- A Type | Per Cylinder | √ | √ | √
1.43.6 | Oxygen Gas Refilling- B Type | Per Cylinder | √ | √ | √ | √
1.43.7 | Oxygen Gas Refilling- C Type | Per Cylinder | √ | √ | √
1.43.8 | Oxygen Gas Refilling- D Type | Per Cylinder | √ | √ | √
1.43.9 | Nitrogen Gas Refilling- D Type | Per Cylinder | √ | √
1.43.10 | Nitrous Oxide Gas Refilling- D Type | Per Cylinder | √ | √ | √
1.43.11 | Nitrous Oxide Gas Refilling- A Type | Per Cylinder | √ | √ | √
1.43.12 | Nitrous Oxide Gas Refilling- B Type | Per Cylinder | √ | √ | √
1.43.13 | IV drip set compatible with infusion pump | Piece | √ | √
1.43.14 | BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 5cm (find out he length) | Piece | √ | √
1.43.15 | BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 5cm (find out he length) | Piece | √ | √
1.43.16 | BLOCK NEEDLES echogenic single shot nerve block needle with 30 degree bevel and extension set, 22 G 8cm | Piece | √ | √
1.43.17 | Hospital Emergency Crash cart: 1. ABS trolley body 2. Plastic-steel columns 3. Adjustable height IV pole attached to unit 4. Removable Dust basket 5. Utility container 6. File bag 7. Defibrillator shelf, power outlet & hooks 8. Sliding side shelf 9. Stainless steel guard rail 10.Concave top covered transparent soft plastic glass 11.Centralized lock 12.Two small & Three big drawer with adjustable partition which can be well organized inside the each drawer. 13.Drawers repositioned eaily 14.Luxurious noiseless castersinner with partitions can be well organized 15.ABS basket & drawer frame can be repositoned easily 16.Luxuriouos casters wheels which enables easy trolleying / movement, with two with brakes 17. CPR board 18. Oxygen tank holder 29. Needle disposal bottle | set | √
1.44 | Surfactant connector
1.44.1 | Surfactant connector attaches to ET 3/3.5 with side port | Piece | √ | √
1.45 | Mucosal atomizer device
1.45.1 | Laryngo-tracheal mucosal atomizer device without syringe | Piece | √ | √
1.46 | Pressure monitoring lines
1.46.1 | Non-expanding PMO line 200cm with luer lock
1.47 | Connectors
1.47.1 | Male to male luer lock connector
1.47.2 | Male luer lock to female connector
1.47.3 | Disposable straight connector 22F-22F
1.47.4 | Disposable straight connector 22M-22F/15F
1.47.5 | Disposable straight connector 22M/15F-22M/15F

## Chemotherapy
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Chemotherapy
1 | Equipments
1.1 | Negative pressure Isolator for chemotherapy drug mixing; Two glove ports, tw+A1o transfer chambers, exhaust spigot for connection to dedicated exhaust system (Ducted type). Should also include remote fans for the exhaust system. [OR Four glove ports, two transfer chambers, exhaust spigot for connection to dedicated exhaust system (Ducted type). Should also include remote fans for the exhaust system]. Additional features: Negative pressure for operation and safety. EU GMP Grade A unidirectional airflow work area for aseptic work and product protection. Transfer chambers with electromagnetic timed interlocks and high air change rates for total suppression of airborne contamination during material transfers in and out, and for rapid evaporation of disinfecting agents. Digital display for isolator functions and transfer chamber operation. Automated pressure decay leak test for indicating results and monitoring trends. PLC controls for easy set up, operation and maintenance. Good access to all potentially contaminated areas to facilitate cleaning. Safe to change main (primary) filter. 316L stainless steel polished internally and externally. 13 amp single splash proof socket. Service taps. Glove sleeve leak test disc with digital manometer. Remote fan for the exhaust system. Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √
2 | Consumables
2.1 | Gown with cuffs and back closure; impermeable fabric | Set | √
2.2 | Gloves; for protection | Set | √
2.3 | Goggles | Pair | √
2.4 | Shoe cover; disposable | Pair | √
2.5 | Mask; disposable | Piece | √
2.6 | Absorbable towel | Piece | √
2.7 | Mask (N-95) | Piece | √
2.8 | Disposable sponges | Piece | √
2.9 | Container; punture and leak proof | Piece | √
3 | Container of desiccant powder or granules to absorb wet contents | Piece | √
3.1 | Disposable dust pan | Piece | √
3.2 | Plastic scraper | Piece | √
3.3 | Plastic bag (Ziplock ) | Piece | √
3.4 | Butterfly needle (scalp vein set), Size 23 | Piece | √
3.5 | Butterfly needle (scalp vein set); Size 24 | Piece | √
3.6 | Butterfly needle (scalp vein set); Size 25 | Piece | √
3.7 | Water for injection - plastic 50ml | Piece | √
3.8 | Water for injection-plastic 100ml | Piece | √
3.9 | Syringe with Luer lock cap 2ml (disposable) | Piece | √
4 | Syringe with Luer lock cap 5ml (disposable) | Piece | √
4.1 | Syringe with Luer lock cap 10ml (disposable) | Piece | √
4.2 | Syringe with Luer lock cap 20ml (disposable) | Piece | √
4.3 | Syringe with Luer lock cap 50ml (disposable) | Piece | √

## CSSD
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
CSSD
1 | Equipments
1.1 | Autoclave horizontal (Large): Chamber size: 500mm×5000mm×950 (W×L×H). Electrical power: 18Kw, 3 phase, 440 V 50Hz AC supply. Working pressure: 1.2 to 2.2 kg/Sq. cm ( upto 21 psi), at 121 degree Celsius. Inner chamber: stainless steel ( thickness 5mm to 10 mm) with argon welding. | Piece | √ | √
1.2 | Autoclave vertical 400x600mm High grade stainless steel SS 304 or equivalent, triple walled, internal joint should be argon arc welded; Internal chamber Size (75 to 80 Ltrs): 400mm (diameter)X600mm (depth) ;supplied with two suitable stackable stainless steel drum for sterilization,waterlevel gauge ( Min and Max), temperature and pressure indicator, safety valve,air release valve, water inlet valve and drainage valve, sillicon gasket, automatic cutoff protection for low water, stainless steel drum stand, . Fully automatic programmable controller with digital display: working pressure 15 to 20 psi , working temperature 115 deg to 132 deg. Power supply: 220-240 AC, 50Hz. Power cord of round 3 pin top. Operating enviromental condition: Humidity > 80%, temperature -10 deg to + 40 deg . Should conform to IEC 60601/ ISI or higher certification (CE, FDA,ISO ). | Set | √ | √ | √ | √
1.3 | Autoclave double door: 304 / 316 / 316L Stainless Steel autoclave horizontal (Large) Double door,Size 1100 X 950 X 950 or 1300x1000x900mm (working pressure: 5psi-20psi; working temperature: 115 deg to 132 deg ; shape:square; working pressure :2.5kgf/cm sq.g(35psi) standard controls through PLC (Programmable Logic Controller): multioperating valve, quick vaccum drying apparatus,accidental vaccum Breaker, safety valve, air release valve, water inlet and drainage valve, pressure display, chamber condensate line, pressure & compound gauge). Accessories: low water protection device, Automatic pressure controsl switch (Pressure control switch, Automatic water cut-off device, digital temperature indicator, timer with alarm system, vacuum pump Power supply : 3 phase AC (440V), 50Hz, Power cord of round 3 pin top. Operational enviromental operating condition: Humidity > 80%, temperature -10 deg to + 40 deg . Should conform to IEC 60601/ ISI or higher certification (CE, FDA,ISO ). | piece | √
1.4 | Programmable sterilizer with high vacuum process including low temperature cycle for heat sensitive goods, leak test, Bowie/Dick tape test, environmental operating condition:-10±40° C, RH>80%,altitude>2000m,ISO/TUV/CE certified, Specify loading trays, Volume>1200lts,Chamber depth>1500mm.3 pin round type plug top | Piece | √ | √
1.5 | Automatic Instrument Washing Machine: (10 DIN): Power supply: 220V-230V, Power cord of round 3 pin round type plug top. | Piece | √ | √ | √ |  | Where there is functional OT .
1.6 | Vaccum, Ultrasonic and plasma cleaning and drying equipment: Power supply: 220V-230V, Power cord of round 3 pin round type plug top. | Piece | √ | √
1.7 | Automatic instrument washer, disinfector and dryer (15DIN trays): Camber volume 500litres ,Basket volume 350litres . Reduction of instrument cleaning cycle up to 30 minutes | Piece | √ | √
1.8 | Perforated stainless steel instrument tray with lid: size 23" x 15" x 5.5" | Piece | √ | √
1.9 | Instrument tray: size 16.5" x 8.5" x 2" | Piece | √
2 | Reagents for Plasma sterrad 100s
2.1 | Sterrad 100S sterilliant cassete: code 10113 | 5pcs/ pkt | √ | √
2.2 | Chemical Indiactor strip: Code 14100 | 1000pcs/pk | √ | √
2.3 | Chemical Indiactor Tape: Code 14202 | 6pcs/roll/pkt | √ | √
2.4 | Cycle Sure Biological Indiactor: Code14324 | 4pcs/roll/box | √ | √
2.5 | Tyvek Roll 4" x 70m: Code 12410 | 6pcs/roll/pkt | √ | √
2.6 | Tyvek Roll 6' x 70m: Code 12415 | 4pcs/roll/pkt | √ | √
2.7 | Tyvek Roll 10'x 70m: Code 12425 | 4pcs/roll/pkt | √ | √
2.8 | Tyvek Roll 14' x 70m: Code 12435 | 2pcs/roll/pkt | √ | √
2.9 | Ribbon catridge red/blk | piece | √ | √
2.1 | Casette collection box | piece | √ | √
2.11 | Vapourizer plate | piece | √ | √
2.12 | Printing paper | roll | √ | √
3 | Consumables
3.1 | Autoclave indicating tape: 20mmx50meter | roll | √ | √ | √ | √
3.2 | Checmical indicators strip for stream sterlization | strips per pkt | √ | √ | √
3.3 | total length 8 inch brush length 2.52 ich - Med 100.46 | piece | √ | √ | √
3.4 | total length 7.2 inch Nylon - Med 100.43 | piece | √ | √ | √
3.5 | laparoscopic brush total length 7 inch Med 100.33 | piece | √
3.6 | Flexible bone reamer cleaning brush total length 30 inch brush 3mm - Med 100.57 | piece | √
3.7 | Suction pipe cleaner, total length 12 inch brush 1/8 inch Med 100.34 | piece | √
3.8 | Suction pipe cleaner, total length 12 inch brush 1/4 inch Med 100.35 | piece | √
3.9 | Utility elbow length gloves size 6.5 | piece | √
4 | Utility elbow length gloves size 7 | piece | √
4.1 | Enzymatic Detergent | litre | √
4.2 | Hinge Free Instrunment lubicant spray | 250 ml | √
4.3 | sterlization Perforated instrument tray 12/10 inch | piece | √
4.4 | sterlization Perforated instrument tray 15/17 inch | piece | √
4.5 | 5 shelf shelving/wire mesh with wheels | piece | √

## Dental
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Dental Department
1 | General
1.1 | Equipments
1.1.1 | Bone Rongeurs Luer 14.5cm | Piece | √
1.1.2 | Browne lingual flap retractor | Piece | √
1.1.3 | Chin Retractor Obwegeser 16.5cm | Piece | √
1.1.4 | Condyle retractor - thin - AK's new | Piece | √
1.1.5 | Cleft lip surgery - set | Set | √
1.1.6 | Cleft plate surgery set | Set | √
1.1.7 | Crown wire twisting forceps: 16cms/6 1/4 inch | Piece | √
1.1.8 | Dingmans mouth gag | Piece | √
1.1.9 | Dissecting Forceps, Adson-brown 12cm | Piece | √
1.1.10 | Dissecting Forceps, Gillies 15.5cm | Piece | √
1.1.11 | Dissecting Forceps, Potts-Smith 16cm | Piece | √
1.1.12 | Dissecting Forceps, Potts-Smith 18cm | Piece | √
1.1.13 | Dissecting Forceps, Potts-Smith 23cm | Piece | √
1.1.14 | Finger protector: langenbeck | Piece | √
1.1.15 | Ferguson jaw opener | Piece | √
1.1.16 | Hammers length 230mm, weight 350g. | Piece | √
1.1.17 | Heister jaw opener - SS | Piece | √
1.1.18 | Jaw Splint - Small | Piece | √
1.1.19 | Jaw Splint - Large | Piece | √
1.1.20 | Lip & Check Retractor - Cawood - Minnesota | Piece | √
1.1.21 | Lip & Check Retractor - Oringer | Piece | √
1.2 | Mandibular Modular Fixation System with following contents:
1.2.1 | Mandible Trauma Module Set (2.0mm) consisting of: | Set | √
1.2.2 | Mandible Trauma Module Set (2.4mm) consisting of: | Set | √
1.3 | Maxillofacial Modular Fixation System with following contents:
1.3.1 | Module 1.5mm | Piece | √
1.3.2 | Module 2.0mm | Piece | √
1.3.3 | Maxillary awl Obwegeser 15cm | Piece | √
1.3.4 | Maxillary awl Obwegeser 17.5cm | Piece | √
1.3.5 | Mouth gag Mckesson - adult | Piece | √
1.3.6 | Mouth gag Mckesson - Child | Piece | √
1.3.7 | Needle holder 16cm | Piece | √
1.3.8 | Needle holder ,Mathieu,170mm | Piece | √
1.3.9 | Needle holder 23cm | Piece | √
1.3.10 | Obwegeser Retractor 24cm | Piece | √
1.3.11 | Osteotomy Chisel 10mm | Piece | √
1.3.12 | Osteotomy Chisel 15mm | Piece | √
1.3.13 | Periosteal Elevator Joseph | Piece | √
1.3.14 | Ramus retractor | Piece | √
1.3.15 | Retractor Adson | Piece | √
1.3.16 | Rowe's Mandible bone holding forceps | Piece | √
1.3.17 | Rowe's Maxilla disimpaction clamp - left | Piece | √
1.3.18 | Rowe's Maxilla disimpaction clamp - right | Piece | √
1.3.19 | Rowe's zygomatic elevator | Piece | √
1.3.20 | Scissor Iris 11.5cm/ 4.5 inch | Piece | √
1.3.21 | Scissor Metzenbaum - fino 18cm/ 7 inch | Piece | √
1.3.22 | Scissor Tenotomy 115mm | Piece | √
1.3.23 | Sinus Lift curettes 15cm | Piece | √
1.3.24 | Sinus Lift curettes 18.5cm | Piece | √
1.3.25 | Surgical Skin hooks | Piece | √
1.3.26 | Third Molar Retractor Obwegeser 24cm | Piece | √
1.3.27 | Tissue retractor Austin | Piece | √
1.3.28 | Tongue depressor(wider) | Piece | √
1.3.29 | Wire cutter,Straight | Piece | √
1.3.30 | Wire ligature pliers for circum mandibuler | Piece | √
1.3.31 | Wire ligature pliers for circum zygomatic | Piece | √
1.3.32 | Wire pusher Rowe | Piece | √
1.3.33 | Zygoma elevator Gillies 20cm | Piece | √
1.3.34 | Zygomatic bone awl curved/ zygomatic arch awl | Piece | √
1.3.35 | Zygomatic bone hook with T handle | Piece | √
1.3.36 | Maxillofacial Modular Fixation System with following contents: a) Obwegeser Retractor 24cm; b)Third Molar Retractor Obwegeser 24cm | Set | √
1.3.37 | Elcomed, contra angle, straight HP, motor cable and motor holder, surgical hand piece (W&H) | Set | √
1.3.38 | Tungsten carbide burs for surgical- Round (2.35mmX30-40mm) | piece | √
1.3.39 | Tungsten carbide burs for surgical-Taper criss-cut (2.35mmX30-40mm) | piece | √
1.3.40 | Tungsten carbide burs for surgical-Cylindrical (2.35mmx30-40mm) | piece | √
1.3.41 | Drill Bits with stops, for Stryker coupling, length- 44.5mm, diameter-1.55mm | piece | √
1.3.42 | Drilling depth 6mm | piece | √
1.3.43 | Drilling depth 8mm | piece | √
1.3.44 | Drilling depth 12mm | piece | √
1.3.45 | Bur block stainless steel (sterilizable) for surgical | Piece | √
1.3.46 | Surgical drill-Stryker COMMAND2 50K DRILL. | piece | √
1.3.47 | Cord for Surgical Drill Stryker command2 –FISHER 104 | piece | √
1.3.48 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 6 mm for stryker coupling j-latch type | piece | √
1.3.49 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 8 mm for stryker coupling j-latch type | piece | √
1.3.50 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 10 mm for stryker coupling j-latch type | piece | √
1.3.51 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 12 mm for stryker coupling j-latch type | piece | √
1.3.52 | 2 mm Mandible Cortex Screw , Self- tapping,L 6mm,pure titanium , | piece | √
1.3.53 | 2 mm Mandible Cortex Screw , Self- tapping,L 8mm,pure titanium , | piece | √
1.3.54 | 2 mm Mandible Cortex Screw , Self- tapping,L 10mm,pure titanium , | piece | √
1.3.55 | 2 mm Mandible Cortex Screw , Self- tapping,L 12mm,pure titanium , | piece | √
1.3.56 | 2.4 mm Emergency Screw length 6mm,pure titanium | piece | √
1.3.57 | 2.4 mm Emergency Screw length 8mm,pure titanium | piece | √
1.3.58 | 2.4 mm Emergency Screw length 10mm,pure titanium | piece | √
1.3.59 | 2.4 mm Emergency Screw length 12mm,pure titanium | piece | √
1.3.60 | Mandible Mini Plates 2.0,centre space, 4holes,length 28mm,pure titanium | piece | √
1.3.61 | Mandible Mini Plates 2.0,centre space, 6holes,length 40mm,pure titanium | piece | √
1.3.62 | Mandible Mini Plates 2.0,centre space, 20holes,length 119mm,pure titanium | piece | √
1.3.63 | Chin plates 2.0,double bend,5 holes,pure titanium , | piece | √
1.3.64 | Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece | √
1.3.65 | Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece | √
1.3.66 | 1.0mm MF Cortex Screw,Self tapping, Length 4mm,pure titanium, | piece | √
1.3.67 | 1.0mm MF Cortex Screw,Self tapping, Length 5mm,pure titanium, | piece | √
1.3.68 | 1.0mm MF Cortex Screw,Self tapping, Length 6mm,pure titanium, | piece | √
1.3.69 | Double Drill Sleeve 2.4/1.8 | piece | √
1.3.70 | Depth Gauge for 1.5 to 2.0 screws | piece | √
1.3.71 | IMF Screw Set with stainless steel implants with lid, 1. Wire twister | piece | √
1.3.72 | IMF Screw Set with stainless steel implants with lid, 1. Wire Cutter | Piece | √
1.3.73 | Dental Amalgamator for silver alloy powder & mercury,Power 220-240V, power cord with round 3 pin top. | No | √ | √ | √
1.3.74 | Acrylizer bath electrical with temperature regulator. Temperature range 0- 100°C, Voltage 220-240,power cord with round 3 pin top. | No | √ | √ | √
1.3.75 | Dental Air Compressor, oil free, 1-4 hp, pressure upto 8 bar, noise free, capicity 25 - 30L, Voltage 220-240V, power cord with round 3 pin top. | No | √ | √ | √
1.3.76 | Dental Chair, High end complete set( with low and high speed Hand pieces, Scaler, Air Compressor unit, suction machine, dentsit and assisat chairs stools, three way syringe, light cure unit, x ray viewer,oral camera with screen) Voltage 220-240v, 50Hz. Suitable Voltage stabilizer. All power cords of round 3 pin plug tops | Set | √
1.3.77 | Dental Chair Standard, Complete set (with low and high speed hand pieces, air compressor unit, x ray viewer,two stools, three way syringe, light cure unit), Suitable Voltage stabilizer having high/low cut off with delay timer, Power cord with round 3 pin tops. | Set | √ | √ | √
1.3.78 | Water distilllation flask, Capacity 5 Litres, Voltage supply 220-240v, 50Hz. Power cord with round 3 pin top | No | √ | √ | √
1.3.79 | Dental iontophoresis device, direct current/pulse,0.4mV current,12V, Input volatge supply 220-240, three pin round tops, | No | √ | √
1.3.80 | Light curing unit (cordless) with light guide, Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
1.3.81 | Lathe polishing machine with builtin burchuck & eye shield double arm. Power supply 220-240v, 50Hz. Power cord with round 3 pin top,Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. | No | √ | √ | √
1.3.82 | Lathe machine Trimmer , nonstop 220v builtin burchuck, inbuilt light-system, suction chamber & eye shield. Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
1.3.83 | Micromotar, rpm 1000-50,000, Electrical, 220V-240V, table top, with CHP,50Hz & foot control, with oil,Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
1.3.84 | Micromotor , rpm 1000-50,000, with Cable for SHP attachment. Power Supply 220-230v, 50Hz, Power cord with round 3 pin top. | no | √ | √ | √
1.3.85 | OPG macine with Cephalometric CDR program,Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | No | √ | √ | √
1.3.86 | Vacuum alginate/plaster/stone mixer, Table top with inbuilt Vacuum pump/gauze digital display along mixing 170 & 560 cc Bowls/covers/blades. Power supply 220-240v, 50Hz. Power cord with round 3 pin top. | No | √ | √ | √
1.3.87 | Dental Vibrator, Power supply 220-240v, 50Hz. Power cord with round 3 pin top. | No | √ | √ | √
1.3.88 | Absorbable gelatin haemostatic sponge | 10pcs/ pkt | √ | √ | √
1.3.89 | Dispensing gun, periodontal dressing material. | Set | √ | √ | √
1.3.90 | Disposable suction tube, standard zsize,dental | 100pcs/ pkt | √ | √
1.3.91 | Dycal cement, Calcium Hydroxide Composition, (Base 13 gm, catalyst 11 gm with mixing pad) | Packet/set of Base 13 gm and catalyst 11 gm | √ | √
1.3.92 | Cocoa-butter tube for glass ionomer | Set | √ | √ | √
1.3.93 | Glass Ionomer, Light cured material set | Set | √ | √
1.3.94 | Miracle Mix, restorative | Set | √ | √ | √
1.3.95 | Mortar and pestle, Glass/porcelain | Set | √ | √ | √
1.3.96 | Silver alloy powder | 30gm | √ | √ | √
1.3.97 | Iodoform paste, Paedo root canal sealant | Set | √ | √ | √
1.3.98 | X-ray lead apron with thyroid collar- flexible | piece | √ | √ | √
1.3.99 | ZnOE Impression (base-catalyst) paste tube | Set | √ | √ | √
1.3.100 | Forceps Cow-Horn, Lower | Piece | √ | √ | √
1.3.101 | Forceps Cow-Horn, Upper | Piece | √ | √ | √
1.3.102 | A.R.T. (autraumatic filling) Hand instrument set of 8 with head lamp. | Set | √ | √ | √
1.3.103 | CHP Hand piece for Micromotor, compatible with exisiting machine | Piece | √ | √ | √
1.3.104 | Spot welder (Dental use) | Piece | √ | √ | √
1.3.105 | Whitening Accelerator (UV light) for bleaching | Piece | √ | √
1.3.106 | Scissor Crown curved, | Piece | √ | √ | √
1.3.107 | Universal Wire Cutter Bandage Scissor, Size 120 mm, SS with golden handle. | piece | √ | √ | √
1.3.108 | Osada Model XL-S30 (Oral surgery),Oral surgery and implant, Handpiece system,( All electric High torque/low speed), Micromotor handpiece and built-in Peristaltic irrigation system. Micromotor with AMC: LVSA,Hand piece:Speed range 1600-60,000 rpm. SH-31 SP Double speed Twist Chuck Handpiece,LVSA micromotor/MCC 8-8 ft coiled cord | Piece | √
1.3.109 | Portable All-in-One PEDO-30W,Power Console ( Floor Unit ), Variable Speed ( 800-30,000 rpm), Standard Podiatric Handpiece: SH-29 S/ LVS /MC S8 on HHR Rubber,Output Speed Range 550-11,000 rpm | Set | √
2 | Consumables
2.1 | Absorbable gelatin haemostatic sponge | 10pcs/ pkt | √ | √ | √
2.2 | Acrylic powder, veined, heat cure with liquid Cold mold seal, separating sheet | Set | √ | √ | √
2.3 | Acrylic, powder, clear, heat cure with liquid, cold mold seal, separating sheet | Set | √ | √ | √
2.4 | Acrylic, powder, pink, heat cure with liquid, cold mold seal, separating sheet | Set | √ | √ | √
2.5 | Acrylic, powder, pink, cold cure with liquid | Set | √ | √ | √
2.6 | Acrylic, powder, clear, cold cure with liquid | Set | √ | √ | √
2.7 | Acrylic trimmer, SHP,steel, round – small, std size | Piece | √ | √ | √
2.8 | Acrylic trimmer, steel, SHP round – medium, std size | Piece | √ | √ | √
2.9 | Acrylic trimmer, steel, SHP,round – large, std size | Piece | √ | √ | √
3 | Acrylic trimmer, steel, SHPflame shape – small, std size | Piece | √ | √ | √
3.1 | Acrylic trimmer, steel, SHP flame shape – medium, | Piece | √ | √ | √
3.2 | Acrylic trimmer, steel, SHP flame shape – large, std size | Piece | √ | √ | √
3.3 | Articulating paper, standard-40/60µm | Packet of 12 books of 12 sheets | √ | √ | √
3.4 | Band for matrix retainer, Ivory No.1 | 10pcs/ pkt | √ | √ | √
3.5 | Band for matrix retainer, Tofflemire contra-angle, junior | 10pcs/ pkt | √ | √ | √
3.6 | Band for matrix retainer, Tofflemire contra-angle, adult | 10pcs/ pkt | √ | √ | √
3.7 | Bristle brush, mounted, CHP | Piece | √ | √ | √
3.8 | Bur holder,revolving 7” diameter 4” high(400x343) | Piece | √ | √ | √
3.9 | Bur cleaning brush | Piece | √ | √ | √
4 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 006 | 5pcs/pkt | √ | √ | √
4.1 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 008 | 5pcs/pkt | √ | √ | √
4.2 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 010 | 5pcs/pkt | √ | √ | √
4.3 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 012 | 5pcs/pkt | √ | √ | √
4.4 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 014 | 5pcs/pkt | √ | √ | √
4.5 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 016 | 5pcs/pkt | √ | √ | √
4.6 | Burs Diamond CHP friction grip for High speed turbine HP, Inverted, size ISO 018 | 5pcs/pkt | √ | √ | √
4.7 | Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 006 | 5pcs/pkt | √ | √ | √
4.8 | Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 008 | 5pcs/pkt | √ | √ | √
4.9 | Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 010 | 5pcs/pkt | √ | √ | √
5 | Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 012 | 5pcs/pkt | √ | √ | √
5.1 | Burs Diamond CHP friction grip for Turbine/ Airotor HP Tapered fissure size ISO 014 | 5pcs/pkt | √ | √ | √
5.2 | Burs Diamond CHP friction grip for High speed turbine HP, Tapered fissure-long Shank, size ISO 016 | 5pcs/pkt | √ | √ | √
5.3 | Burs Diamond CHP friction grip for High speed turbine HP, Tapered fissure-Long Shank, size ISO 018 | 5pcs/pkt | √ | √ | √
5.4 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 006 | 5pcs/pkt | √ | √ | √
5.5 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 008 | 5pcs/pkt | √ | √ | √
5.6 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 010 | 5pcs/pkt | √ | √ | √
5.7 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 012 | 5pcs/pkt | √ | √ | √
5.8 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 014 | 5pcs/pkt | √ | √ | √
5.9 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 016 | 5pcs/pkt | √ | √ | √
6 | Burs Diamond CHP friction grip for High speed turbine HP, Straight, fissure size ISO 018 | 5pcs/pkt | √ | √ | √
6.1 | Carborundum discs | 100disc/ pkt | √ | √ | √
6.2 | Cellophane/ Mylar strips | 100dics/ pkt | √ | √ | √
6.3 | Chamois leather for amalgam | Piece | √ | √ | √
6.4 | Clove oil | 500 ml | √ | √ | √
6.5 | Dispensing gun, periodontal dressing material. | Set | √ | √ | √
6.6 | Cold mold seal of Sodium alginate solution (450 ml) | Bottle | √ | √ | √
6.7 | Composite light cure material, complete set | Packet/ set | √ | √ | √
6.8 | Composite finishing kit | Packet/ kit | √ | √ | √
6.9 | Cotton buff wheel, Large for Lathe,4½” | Piece | √ | √ | √
7 | Cotton buff wheel, Small for SHP, mounted,3” | Piece | √ | √ | √
7.1 | Crown preparation set | set | √ | √ | √
7.2 | Dappen glass, different colours | Piece | √ | √ | √
7.3 | Dental wedges wooden, readymade, assorted of 4 different colours. | 400pcs pkt | √ | √ | √
7.4 | Dental floss | Piece | √ | √ | √
7.5 | Dental Plaque disclosing Solution/ gel | 45ml | √ | √ | √
7.6 | Denture adhesive powder | 15gm | √ | √ | √
7.7 | Denture cleaning powder | 15gm | √ | √ | √
7.8 | Disposable suction tube, standard zsize,dental | 100pcs/ pkt | √ | √ | √
7.9 | Dry socket paste | set | √ | √ | √
8 | Dycal cement, Calcium Hydroxide Composition, (Base 13 gm, catalyst 11 gm with mixing pad) | Packet/set | √ | √ | √
8.1 | Endo: Broaches, barbed, 15-40, 21 mm length | Pkt | √ | √ | √
8.2 | Endo: Broaches, barbed, 15-40, 25mm length | Pkt | √ | √ | √
8.3 | Endo: EDTA gel for canal irrigation, 5gm | Piece | √ | √ | √
8.4 | Endo: Finger spreader assorted, 15-40, | Pkt | √ | √ | √
8.5 | Endo: Forma cresol solution, | 10 ml | √ | √ | √
8.6 | Endo: GP points, assorted pkt, size 15-40 | Pkt | √ | √ | √
8.7 | Endo: GP points, assorted pkt, size 45-80 | Pkt | √ | √ | √
8.8 | Endo: Hedstroem file NiTi, 25mm, assorted pkt, size 45 - 80 | Pkt | √ | √ | √
8.9 | Endo: Hedstroem file NiTi, 25mm, assorted pkt, size 15 - 40 | Pkt | √ | √ | √
9 | Endo: Hedstroem file NiTi, 21mm, assorted pkt, size 15 - 40 | Pkt | √ | √ | √
9.1 | Endo: K- file, NiTi, 25mm, assorted pkt, size 15- 40 | Pkt | √ | √ | √
9.2 | Endo: K- file, NiTi, 21mm, assorted pkt, size 15- 40 | Pkt | √ | √ | √
9.3 | Endo: K- file, NiTi, 21mm, assorted pkt, size 15- 40 | Pkt | √ | √ | √
9.4 | Endo: K - file, NiTi, 25mm, assorted pkt, size 45 - 80 | Pkt | √ | √ | √
9.5 | Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 15 - 40 | Pkt | √ | √ | √
9.6 | Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 45 -80 | Pkt | √ | √ | √
9.7 | Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 15 - 40, Lock grip assorted | Pkt | √ | √ | √
9.8 | Endo: K - Reamer, NiTi, 25mm, assorted pkt, size 45 -80, lock grip assorted | Pkt | √ | √ | √
9.9 | Endo: Lentulo - Spiral, NiTi, 25mm, Lock grip, assorted, | Pkt | √ | √ | √
10 | Endo: Lentulo-spiral, assorted,19mm lock grip | Pkt | √ | √
10.1 | Endo: Orifice opener | Pkt | √ | √ | √
10.2 | Endo: Paper point, assorted packet, size 15-40 | Pkt | √ | √ | √
10.3 | Endo: Paper point, assorted packet, size 45 - 80 | Pkt | √ | √ | √
10.4 | Endo: Probe, periodontal (graded) | Pkt | √ | √ | √
10.5 | Endo: RCT obturation paste, Endomethasone N, Powder 14 gm and liquid 10 ml | Set | √ | √ | √
10.6 | Endo: RCT plugger | Piece | √ | √ | √
10.7 | Endo: Root canal filling solvent | 50ml | √ | √ | √
10.8 | Endo: Sodium hypochlorite 3% | 450ml | √ | √ | √
10.9 | Endo: pulp devitilizer(Paraformaldehyde, Phenol, Lidocaine Base) | Ml | √ | √ | √
11 | Endo RCT sealer Endomethasone powder and liquid | Set | √ | √ | √
11.1 | Endobox, aluminuim, with stand, small (104x44x40mm) | Piece | √ | √ | √
11.2 | Endobox, aluminuim, with stand, large (204x105x54mm) | Piece | √ | √ | √
11.3 | Endo: RCT measurement kit | Piece | √ | √ | √
11.4 | Erich ARCH bar, 12 inch | Piece | √ | √
11.5 | Eye-shield Plastic transparent, spectacle type | Piece | √ | √ | √
11.6 | Felt cone, large (1”x4”) | Piece | √ | √ | √
11.7 | Felt cone, medium (1”x3”) | Piece | √ | √ | √
11.8 | Glass Ionomer, Fuji VII,VII, IX | Set | √ | √ | √
11.9 | Cocoa-butter tube for glass ionomer | Set | √ | √ | √
12 | Glass Ionomer, Light cured material set | Set | √ | √
12.1 | Glass Ionomer, Pit & Fissure Sealant material | Set | √ | √ | √
12.2 | Miracle Mix, restorative | Set | √ | √ | √
12.3 | Glass mixing slab, 6"x 3" x 1/2" | Piece | √ | √ | √
12.4 | Gum paint, Gumex | 60ml | √ | √ | √
12.5 | Handpiece lubricating oil spray | Cane | √ | √ | √
12.6 | Impression Compound | 5pcs/pkt | √ | √ | √
12.7 | Impression material Sodium-Alginate, normal setting | 450gm | √ | √ | √
12.8 | Lathe stone wheel – small (2”x¼”) | Piece | √ | √ | √
12.9 | Lathe stone wheel - medium(3”x¼”) | Piece | √ | √ | √
13 | Lathe stone wheel - large(4”x¼”) | Piece | √ | √ | √
13.1 | Mandrel for Sand paper SHP | Piece | √ | √ | √
13.2 | Mandril for disc, CHP, L.G. | Piece | √ | √ | √
13.3 | Mandril for disc, SHP | Piece | √ | √ | √
13.4 | Mercury liquid, 30gm (100%) | Piece | √ | √ | √
13.5 | Mortar and pestle, Glass/porcelain | Set | √ | √ | √
13.6 | Mould guide, rubber, lower | Piece | √ | √ | √
13.7 | Mould guide, rubber, upper | Piece | √ | √ | √
13.8 | Mouth gag , rubber adult | Piece | √ | √ | √
13.9 | Plaster of Paris | 50kg/ drum | √ | √ | √
14 | Plastic disc, soflex polishing | 100dics/ pkt | √ | √ | √
14.1 | Plastic face shield | 10pcs/ pkt | √ | √ | √
14.2 | Polishing rubber cups, CHP, L.G. | Piece | √ | √ | √
14.3 | Polishing brush for lathe | Piece | √ | √ | √
14.4 | Polysulfide impression Paste, Base/Catalyst set/kit, Regular: Kerr | Set | √ | √ | √
14.5 | Polysulfide rubber base tray Adhesive: Kerr | Set | √ | √
14.6 | Prophylaxis paste, Dose Assorted | Piece | √ | √ | √
14.7 | Provi-paste filling material | 20gm | √ | √ | √
14.8 | Pumice powder | 5kg/pkt | √ | √ | √
14.9 | Rubber bowl – large (12.5cm) | Piece | √ | √ | √
15 | Rubber bowl – small (9.5cm) | Piece | √ | √ | √
15.1 | Sand paper disc, assorted | 100dics/ pkt | √ | √ | √
15.2 | Sharpening stone, Arkansas, double side | Piece | √ | √ | √
15.3 | Shellac Base Plate - Lower | 50pcs/ pkt | √ | √ | √
15.4 | Shellac Base Plate - Upper | 50pcs/ pkt | √ | √ | √
15.5 | Silver alloy powder | 30gm | √ | √ | √
15.6 | Stone mounted, SHP, cylindrical – Small,std size | Piece | √ | √ | √
15.7 | Stone mounted, SHP, cylindrical – Medium std size, | Piece | √ | √ | √
15.8 | Stone mounted, SHP, cylindrical – Large, std size | Piece | √ | √ | √
15.9 | Stone mounted, SHP, flame shaped – Small, std size | Piece | √ | √ | √
16 | Stone mounted, SHP, flame shaped – Medium, std size | Piece | √ | √ | √
16.1 | Stone mounted, SHP, flame shaped – Large, std size | Piece | √ | √ | √
16.2 | Stone mounted, SHP, round – Small, std size | Piece | √ | √ | √
16.3 | Stone mounted, SHP, round – Medium, std size | Piece | √ | √ | √
16.4 | Stone mounted, SHP, round – Large, std size | Piece | √ | √ | √
16.5 | Stone mounted, SHP, tapered – Small, std size | Piece | √ | √ | √
16.6 | Stone mounted, SHP, tapered – Medium, std size | Piece | √ | √ | √
16.7 | Stone mounted, SHP, tapered – Large, std size | Piece | √ | √ | √
16.8 | Stone plaster | 1kg/pkt | √ | √ | √
16.9 | Teeth, acrylic cross linked, shade - 24, lower anterior small, set of 6 | Set | √ | √ | √
17 | Teeth, acrylic cross linked, shade - 24, lower anterior medium, set of 6 | Set | √ | √ | √
17.1 | Teeth, acrylic cross linked, shade - 24, lower anterior large, set of 6 | Set | √ | √ | √
17.2 | Teeth, acrylic cross linked, shade - 24, upper & lower posterior small set of 16 | Set | √ | √ | √
17.3 | Teeth, acrylic cross linked, shade - 24, upper & lower posterior medium set of 16 | Set | √ | √ | √
17.4 | Teeth, acrylic cross linked, shade - 24, upper & lower posterior large set of 16 | Set | √ | √ | √
17.5 | Teeth, acrylic cross linked, shade - 24, set of 28 upper & lower full set small | Set | √ | √ | √
17.6 | Teeth, acrylic cross linked, shade - 24, set of 28 upper & lower full set medium | Set | √ | √ | √
17.7 | Teeth, acrylic cross linked, shade - 24, set of 28 upper & lower full set large | Set | √ | √ | √
17.8 | Teeth, acrylic cross linked, shade - 24, upper anterior small, set of 6 | Set | √ | √ | √
17.9 | Teeth, acrylic cross linked, shade - 24, upper anterior medium, set of 6 | Set | √ | √ | √
18 | Teeth, acrylic cross linked, shade - 24, upper anterior large, set of 6 | Set | √ | √ | √
18.1 | Teeth, acrylic cross linked, shade - 25, lower anterior small, set of 6 | Set | √ | √ | √
18.2 | Teeth, acrylic cross linked, shade - 25, lower anterior medium, set of 6 | Set | √ | √ | √
18.3 | Teeth, acrylic cross linked, shade - 25, lower anterior large, set of 6 | Set | √ | √ | √
18.4 | Teeth, acrylic cross linked, shade - 25, upper & lower posterior small set of 16 | Set | √ | √ | √
18.5 | Teeth, acrylic cross linked, shade - 25, upper & lower posterior medium set of 16 | Set | √ | √ | √
18.6 | Teeth, acrylic cross linked, shade - 25, upper & lower posterior large set of 16 | Set | √ | √ | √
18.7 | Teeth, acrylic cross linked, shade - 25, set of 28 upper & lower full set small | Set | √ | √ | √
18.8 | Teeth, acrylic cross linked, shade - 25, set of 28 upper & lower full set medium | Set | √ | √ | √
18.9 | Teeth, acrylic cross linked, shade - 25, set of 28 upper & lower full set large | Set | √ | √ | √
19 | Teeth, acrylic cross linked, shade - 25, upper anterior small, set of 6 | Set | √ | √ | √
19.1 | Teeth, acrylic cross linked, shade - 25, upper anterior medium, set of 6 | Set | √ | √ | √
19.2 | Teeth, acrylic cross linked, shade - 25, upper anterior large, set of 6 | Set | √ | √ | √
19.3 | Teeth, acrylic cross linked, shade - 26, lower anterior small, set of 6 | Set | √ | √ | √
19.4 | Teeth, acrylic cross linked, shade - 26, lower anterior medium, set of 6 | Set | √ | √ | √
19.5 | Teeth, acrylic cross linked, shade - 26, lower anterior large, set of 6 | Set | √ | √ | √
19.6 | Teeth, acrylic cross linked, shade - 26, upper & lower posterior small set of 16 | Set | √ | √ | √
19.7 | Teeth, acrylic cross linked, shade - 26, upper & lower posterior medium set of 16 | Set | √ | √ | √
19.8 | Teeth, acrylic cross linked, shade - 26, upper & lower posterior large set of 16 | Set | √ | √ | √
19.9 | Teeth, acrylic cross linked, shade - 26, set of 28 upper & lower full set small | Set | √ | √ | √
20 | Teeth, acrylic cross linked, shade - 26, set of 28 upper & lower full set medium | Set | √ | √ | √
20.1 | Teeth, acrylic cross linked, shade - 26, set of 28 upper & lower full set large | Set | √ | √ | √
20.2 | Teeth, acrylic cross linked, shade - 26, upper anterior small, set of 6 | Set | √ | √ | √
20.3 | Teeth, acrylic cross linked, shade - 26, upper anterior medium, set of 6 | Set | √ | √ | √
20.4 | Teeth, acrylic cross linked, shade - 26, upper anterior large, set of 6 | Set | √ | √ | √
20.5 | Teeth, acrylic cross linked, shade - 27, lower anterior small, set of 6 | Set | √ | √ | √
20.6 | Teeth, acrylic cross linked, shade - 27, lower anterior medium, set of 6 | Set | √ | √ | √
20.7 | Teeth, acrylic cross linked, shade - 27, lower anterior large, set of 6 | Set | √ | √ | √
20.8 | Teeth, acrylic cross linked, shade - 27, upper & lower posterior small set of 16 | Set | √ | √ | √
20.9 | Teeth, acrylic cross linked, shade - 27, upper & lowerposterior medium set of 16 | Set | √ | √ | √
21 | Teeth, acrylic cross linked, shade - 27, upper & lower posterior large set of 16 | Set | √ | √ | √
21.1 | Teeth, acrylic cross linked, shade - 27, set of 28 upper & lower full set small | Set | √ | √ | √
21.2 | Teeth, acrylic cross linked, shade - 27, set of 28 upper & lower full set medium | Set | √ | √ | √
21.3 | Teeth, acrylic cross linked, shade - 27, set of 28 upper & lower full set large | Set | √ | √ | √
21.4 | Teeth, acrylic cross linked, shade - 27, upper anterior small, set of 6 | Set | √ | √ | √
21.5 | Teeth, acrylic cross linked, shade - 27, upper anterior medium, set of 6 | Set | √ | √ | √
21.6 | Teeth, acrylic cross linked, shade - 27, upper anterior large, set of 6 | Set | √ | √ | √
21.7 | Teeth, acrylic cross linked, shade - 28, lower anterior small, set of 6 | Set | √ | √ | √
21.8 | Teeth, acrylic cross linked, shade - 28, lower anterior medium, set of 6 | Set | √ | √ | √
21.9 | Teeth, acrylic cross linked, shade - 28, lower anterior large, set of 6 | Set | √ | √ | √
22 | Teeth, acrylic cross linked, shade - 28, upper & lower posterior small set of 16 | Set | √ | √ | √
22.1 | Teeth, acrylic cross linked, shade - 28, upper & lower posterior medium set of 16 | Set | √ | √ | √
22.2 | Teeth, acrylic cross linked, shade - 28, upper & lower posterior large set of 16 | Set | √ | √ | √
22.3 | Teeth, acrylic cross linked, shade - 28, upper anterior small, set of 6 | Set | √ | √ | √
22.4 | Teeth, acrylic cross linked, shade - 28, upper anterior medium, set of 6 | Set | √ | √ | √
22.5 | Teeth, acrylic cross linked, shade - 28, upper anterior large, set of 6 | Set | √ | √ | √
22.6 | Teeth, acrylic cross linked, shade - 29, lower anterior small, set of 6 | Set | √ | √ | √
22.7 | Teeth, acrylic cross linked, shade - 29, lower anterior medium, set of 6 | Set | √ | √ | √
22.8 | Teeth, acrylic cross linked, shade - 29, lower anterior large, set of 6 | Set | √ | √ | √
22.9 | Teeth, acrylic cross linked, shade - 29, upper & lower posterior small set of 16 | Set | √ | √ | √
23 | Teeth, acrylic cross linked, shade - 29, upper & lower posterior medium set of 16 | Set | √ | √ | √
23.1 | Teeth, acrylic cross linked, shade - 29, upper & lower posterior large set of 16 | Set | √ | √ | √
23.2 | Teeth, acrylic cross linked, shade - 29, upper & lower posterior large set of 17 | Set | √ | √ | √
23.3 | Teeth, acrylic cross linked, shade - 29, upper anterior small, set of 6 | Set | √ | √ | √
23.4 | Teeth, acrylic cross linked, shade - 29, upper anterior medium, set of 6 | Set | √ | √ | √
23.5 | Teeth, acrylic cross linked, shade - 29, upper anterior large, set of 6 | Set | √ | √ | √
23.6 | Iodoform paste, Paedo root canal sealant | Set | √ | √ | √
23.7 | Temporary SS crown standard set | Set | √ | √ | √
23.8 | Topical fluoride gel (1 ltr Jar/Bottle) | 1ltr/bottle | √ | √ | √
23.9 | Tracing green stick | 10pcs/pkt | √ | √ | √
24 | Wax modeling | 10pcs/pkt | √ | √ | √
24.1 | Wax sticky | 10pcs/pkt | √ | √ | √
24.2 | Whitening powder | 1kg/pkt | √ | √ | √
24.3 | Wire spool, SS, G -19 | Piece | √ | √ | √
24.4 | Wire spool, SS, G -20 | Piece | √ | √ | √
24.5 | Wire spool, SS, G -21 | Piece | √ | √ | √
24.6 | Wire spool, SS, G -22 | Piece | √ | √ | √
24.7 | Wire spool, SS, G -23 | Piece | √ | √ | √
24.8 | Wire spool, SS, G -27 | Piece | √ | √ | √
24.9 | Wire spool, SS, G -28 | Piece | √ | √ | √
25 | X- ray Developer (powder) | 1kg/pkt | √ | √ | √
25.1 | X- ray Fixer (Powder) | 1kg/pkt | √ | √ | √
25.2 | X- ray film, Occlusal, 57x76mm (2 ¼” x 3”) No.4,E-speed, minimum 2yrs of expiry date | 50pcs/ pkt | √ | √ | √
25.3 | X- ray film, IOPA, No. 0,E-speed, minimum 2yrs of expiry date | 50pcs/ pkt | √ | √ | √
25.4 | X-ray film, IOPA, No. 2, E-speed, minimum 2yrs of expiry date | 50pcs/ pkt | √ | √ | √
25.5 | X-ray film Bitewing holder | 50pcs/ pkt | √ | √ | √
25.6 | X-ray film lateral ceph, Size - 8x10 inch' | 50pcs/ pkt | √ | √ | √
25.7 | X-ray lead apron with thyroid collar- flexible | piece | √ | √ | √
25.8 | X-ray OPG film, 6x12" | 50pcs/ pkt | √ | √ | √
25.9 | Zinc Phosphate cement , podwder-liquid set | Set | √ | √ | √
26 | ZnOE cement, powder & liquid set | Set | √ | √ | √
26.1 | ZnOE Impression (base-catalyst)paste tube | Set | √ | √ | √
26.2 | Needdle,1”,27G (0.40x21 mm) | 100pcs/box | √ | √ | √
26.3 | Cerclage Wire, 0.6mm dia, without eye, length 175mm | Piece | √
26.4 | Surgical Cutting Burs, SS, HM 141 204 040 | Piece | √
26.5 | Surgical Cutting Burs, SS, 165 205 023 | Piece | √
26.6 | Tungsten carbide SS, HM 141 205 040 | Piece | √
26.7 | Tungsten carbide SS, HM 141 205 050 | Piece | √
26.8 | Tungsten carbide, HM 33T 104 016 | Piece | √
26.9 | Tungsten carbide, HM 151 316 016 | Piece | √
27 | Tungsten carbide, HM 152 315 014 | Piece | √
27.1 | Tungsten carbide, HM 162 104 016 | Piece | √
3 | Instrunment
3.1 | Amalgam burnisher, double ended,(1mm,1.1mm) small | Piece | √ | √ | √
3.2 | Amalgam burnisher, double ended, medium(1.8mm, 3mm) | Piece | √ | √ | √
3.3 | Amalgam burnisher, double ended(1.6mm,2.4mm) | Piece | √ | √ | √
3.4 | Amalgam carver, diamond shape | Piece | √ | √ | √
3.5 | Amalgam carver, double ended, ½ Hollenbeck | Piece | √ | √ | √
3.6 | Amalgam condenser, serrated double ended (1.2mm,1.5mm) small | Piece | √ | √ | √
3.7 | Amalgam carrier, plastic, conventional, straight | Piece | √ | √ | √
3.8 | Amalgam carrier, plastic, curved 45˚ | Piece | √ | √ | √
3.9 | Bone Chisel, osteotome 4" pointed curved, triangular | Piece | √ | √ | √
3.1 | Bone chisel, osteotome 4" semi circular, curved | Piece | √ | √ | √
3.11 | Bone files, Small, 6 inch, double ended | Piece | √ | √ | √
3.12 | Bone nibbler dental, Small, 6 | Piece | √ | √ | √
3.13 | Bone Rongeur | Piece | √ | √ | √
3.14 | Cement plastic filling, medium, double ended | Piece | √ | √ | √
3.15 | Cement plastic filling, small, double ended | Piece | √ | √ | √
3.16 | Cement plugger condenser, medium, double ended | Piece | √ | √ | √
3.17 | Cement plugger condenser, small, double ended | Piece | √ | √ | √
3.18 | Cheek retractor blade | Piece | √ | √ | √
3.19 | Coupland, Apexo elevator, set of 2 | Piece | √ | √ | √
3.2 | Dental Cartridge Syringe 1.8ml | Piece | √ | √ | √
3.21 | Dental flask, brass | Piece | √ | √ | √
3.22 | Elevator, Bristows | Piece | √ | √ | √
3.23 | Elevator, cryer's - left | Piece | √ | √ | √
3.24 | Elevator, cryer's - right | Piece | √ | √ | √
3.25 | Elevator, Periosteal, SS - Small | Piece | √ | √ | √
3.26 | Elevator, Periosteal, SS - large | Piece | √ | √ | √
3.27 | Elevator, Rowes | Piece | √ | √ | √
3.28 | Elevators, Coupland, straight - Small | Piece | √ | √ | √
3.29 | Elevators, Coupland, straight - Medium | Piece | √ | √ | √
3.3 | Elevators, Coupland, straight - Large | Piece | √ | √ | √
3.31 | Elevators, warwick jams, set of 3 | Set | √ | √ | √
3.32 | Endo-spoon excavator, double ended - small | Piece | √ | √ | √
3.33 | Endo-Spoon excavator, double ended - large | Piece | √ | √ | √
3.34 | File, volcanite round | Piece | √ | √
3.35 | File, volcanite, 1/2 round | Piece | √ | √ | √
3.36 | Forcep - bone cutting – lower | Piece | √ | √ | √
3.37 | Forcep - bone cutting – upper | Piece | √ | √ | √
3.38 | Forceps - dissecting, tooth, 4 inch, Adson | Piece | √ | √ | √
3.39 | Forceps - plain serrated tissue, 4 inch | Piece | √ | √ | √
3.4 | Forceps - toothed tissue 4 inch | Piece | √ | √ | √
3.41 | Forceps Cow-Horn, Lower | Piece | √ | √ | √
3.42 | Forceps Cow-Horn, Upper | Piece | √ | √ | √
3.43 | Forceps, Extraction, Deciduous Teeth: Lower Anterior, Klein pattern | Piece | √ | √ | √
3.44 | Forceps, Extraction, Deciduous Teeth: Lower Molar, Klein pattern | Piece | √ | √ | √
3.45 | Forceps, Extraction, Deciduous Teeth: Lower Root, Klein pattern | Piece | √ | √ | √
3.46 | Forceps, Extraction, Deciduous Teeth: Upper Anterior, Klein pattern | Piece | √ | √ | √
3.47 | Forceps, Extraction, Deciduous Teeth: Upper Molar, Klein pattern | Piece | √ | √ | √
3.48 | Forceps, Extraction, Deciduous Teeth: Upper Root, Klein pattern | Piece | √ | √ | √
3.49 | Forceps, Extraction, Deciduous Teeth: Upper Universal, Klein pattern | Piece | √ | √ | √
3.5 | Forceps-extraction, permanent teeth Lower Anterior | Piece | √ | √ | √
3.51 | Forceps-extraction, permanent teeth lower molar | Piece | √ | √ | √
3.52 | Forceps-extraction, permanent teeth lower premolar | Piece | √ | √ | √
3.53 | Forceps-extraction, permanent teeth lower root | Piece | √ | √ | √
3.54 | Forceps-extraction, permanent teeth Lower Third Molar | Piece | √ | √ | √
3.55 | Forceps-extraction, permanent teeth Upper Anterior | Piece | √ | √ | √
3.56 | Forceps-extraction, permanent teeth upper molar, left | Piece | √ | √ | √
3.57 | Forceps-extraction, permanent teeth upper molar, right | Piece | √ | √ | √
3.58 | Forceps-extraction, permanent teeth Upper premolar | Piece | √ | √ | √
3.59 | Forceps-extraction, permanent teeth Upper root | Piece | √ | √ | √
3.6 | Forceps-extraction, permanent teeth Upper Third Molar | Piece | √ | √ | √
3.61 | A.R.T. (autraumatic filling) Hand instrument set of 8 with head lamp. | Set | √ | √ | √
3.62 | Heavy wire cutter | Piece | √ | √ | √
3.63 | Lecrons carver | Piece | √ | √ | √
3.64 | Mallet surgical, for nylon headed | Piece | √ | √ | √
3.65 | Mallet prosthetic SS | Piece | √ | √ | √
3.66 | Mirror, mouth Top, sterilizable, no.5, adult | Piece | √ | √ | √
3.67 | Mirror, handle sterilizable, no. 5 | Piece | √ | √ | √
3.68 | Plaster cutter plier type | Piece | √ | √ | √
3.69 | Plaster knife | Piece | √ | √ | √
3.7 | Plaster spatula | Piece | √ | √ | √
3.71 | Pliers Young s universal | Piece | √ | √ | √
3.72 | Pliers, Adams medium, | Piece | √ | √ | √
3.73 | Pliers, Clasp bending | Piece | √ | √ | √
3.74 | Pliers, loop forming | Piece | √ | √ | √
3.75 | Probe, double ended | Piece | √ | √ | √
3.76 | RCT Plugger | Piece | √ | √ | √
3.77 | Scaler, cingulum | Piece | √ | √ | √
3.78 | Scaler, double ended inter-dental | Piece | √ | √ | √
3.79 | Scaler, Jaquette | Piece | √ | √ | √
3.8 | Scaler, sickle - Small | Piece | √ | √ | √
3.81 | Scaler, sickle - Large | Piece | √ | √ | √
3.82 | Scaler, universal | Piece | √ | √ | √
3.83 | Crown Bandage Scissor, Saw edge, Curved, size 110 m, SS with golden handle. | Piece | √ | √ | √
3.84 | Scissors Surgical straight, sharp pointed 5 " | Piece | √ | √ | √
3.85 | Scissors Surgical straight, sharp pointed 6 " | Piece | √ | √ | √
3.86 | Scissors Surgical curved, sharp pointed 5 " | Piece | √ | √ | √
3.87 | Scissors Surgical curved, sharp pointed 6 " | Piece | √ | √ | √
3.88 | Spatula, cement, SS | Piece | √ | √ | √
3.89 | Spirit lamp, steel/brass, | 250 ml | √ | √ | √
3.9 | Sub - gingival Scaler complete set | Piece | √ | √ | √
3.91 | Tweezer dental, fine tip | Piece | √ | √ | √
3.92 | Wax knife double ended | Piece | √ | √ | √
3.93 | Wax spatula double ended | Piece | √ | √ | √
3.94 | Articulator Hanau H2 with 6 pairs of mounting rings | Piece | √ | √ | √
3.95 | Articulator, free plane, mean Value | Piece | √ | √ | √
3.96 | Articulator, hinge | Piece | √ | √ | √
3.97 | Chip blower with nozzle | Piece | √ | √ | √
3.98 | Impression tray, Edentulous, Non preforated, lower, size – 0 | Piece | √ | √ | √
3.99 | Impression tray, Edentulous, Non preforated, lower, size – 1 | Piece | √ | √ | √
3.1 | Impression tray, Edentulous, Non preforated, lower, size – 2 | Piece | √ | √ | √
3.101 | Impression tray, Edentulous, Non preforated, lower, size – 3 | Piece | √ | √ | √
3.102 | Impression tray, Edentulous, Non preforated, lower, size – 4 | Piece | √ | √ | √
3.103 | Impression tray, Edentulous, Non preforated, upper, size – 0 | Piece | √ | √ | √
3.104 | Impression tray, Edentulous, Non preforated, upper, size – 1 | Piece | √ | √ | √
3.105 | Impression tray, Edentulous, Non preforated, upper, size – 2 | Piece | √ | √ | √
3.106 | Impression tray, Edentulous, Non preforated, upper, size – 3 | Piece | √ | √ | √
3.107 | Impression tray, Edentulous, Non preforated, upper, size – 4 | Piece | √ | √ | √
3.108 | Impression tray Dentulous, perforated lower size – 0 | Piece | √ | √ | √
3.109 | Impression tray Dentulous, perforated lower size – 1 | Piece | √ | √ | √
3.11 | Impression tray Dentulous, perforated lower size – 2 | Piece | √ | √ | √
3.111 | Impression tray Dentulous, perforated lower size – 3 | Piece | √ | √ | √
3.112 | Impression tray Dentulous, perforated lower size – 4 | Piece | √ | √ | √
3.113 | Impression tray Dentulous, perforated, upper size – 0 | Piece | √ | √ | √
3.114 | Impression tray Dentulous, perforated, upper size – 1 | Piece | √ | √ | √
3.115 | Impression tray Dentulous, perforated, upper size – 2 | Piece | √ | √ | √
3.116 | Impression tray Dentulous, perforated, upper size – 3 | Piece | √ | √ | √
3.117 | Impression tray Dentulous, perforated, upper size – 4 | Piece | √ | √ | √
3.118 | Impression tray Dentulous, non-perforated, rimlock, lower size – 1 | Piece | √ | √ | √
3.119 | Impression tray Dentulous, non-perforated, rimlock, lower size – 2 | Piece | √ | √ | √
3.12 | Impression tray Dentulous, non-perforated, rimlock, lower size – 3 | Piece | √ | √ | √
3.121 | Impression tray Dentulous, non-perforated, rimlock, lower size – 4 | Piece | √ | √ | √
3.122 | Impression tray Dentulous, non-perforated, rimlock, lower size – 5 | Piece | √ | √ | √
3.123 | Impression tray Dentulous, non-perforated, rimlock, upper size – 1 | Piece | √ | √ | √
3.124 | Impression tray Dentulous, non-perforated, rimlock, upper size – 2 | Piece | √ | √ | √
3.125 | Impression tray Dentulous, non-perforated, rimlock, upper size – 3 | Piece | √ | √ | √
3.126 | Impression tray Dentulous, non-perforated, rimlock, upper size – 4 | Piece | √ | √ | √
3.127 | Impression tray Dentulous, non-perforated, rimlock, upper size – 5 | Piece | √ | √ | √
3.128 | Endodontic HP apex locator/triauto | Piece | √ | √ | √
3.129 | Flask clamp | Piece | √ | √ | √
3.13 | Flask, table clamp | Piece | √ | √ | √
3.131 | Gingivectomy set, complete | Set | √ | √ | √
3.132 | Gum Scissor Iris straight 14cm | Piece | √ | √ | √
3.133 | Hot plate(metal) | Piece | √ | √ | √
3.134 | Matrix retainer ivory no. 1 | Piece | √ | √ | √
3.135 | Matrix retainer toflemire no. 8 | Piece | √ | √ | √
3.136 | Spot welder (Dental use) | Piece | √ | √ | √
3.137 | Whitening Accelerator (UV light) for bleaching | Piece | √ | √
3.138 | Bending Pliers with beak, for plates up to 2.7 | Piece | √
3.139 | Bending-Cutting Pliers 1.0 to 2.0, right | Piece | √
3.14 | Bone Curettes Hemingway 18.5cm | Piece | √
3.141 | Bone Curettes Volkmann 170mm | Piece | √
3.142 | Bone cutting forcep 18cm | Piece | √
3.143 | Bone file 18cm | Piece | √
3.144 | Bone Holding forcep 13cm | Piece | √
3.145 | Bone Holding forcep 14cm | Piece | √
3.146 | Bone Holding forceps 18.5cm | Piece | √
3.147 | Bone Holding forceps, self-centering, length 191mm, softlock. | Piece | √
3.148 | Bone Reduction forcep with points, narrow 161mm softlock | Piece | √
3.149 | Bone Reduction forcep with points, wide 127mm softlock | Piece | √
3.15 | Bone Rongeur 12.5cm | Piece | √
3.151 | Bone Rongeur 16.5cm | Piece | √
3.152 | Bone Rongeurs Luer 14.5cm | Piece | √
3.153 | Browne lingual flap retractor | Piece | √
3.154 | Chin Retractor Obwegeser 16.5cm | Piece | √
3.155 | Condyle retractor - thin - AK's new | Piece | √
3.156 | Cleft lip surgery - set | Set | √
3.157 | Cleft plate surgery set | Set | √
3.158 | Crown wire twisting forceps: 16cms/6 1/4 inch | Piece | √
3.159 | Dingmans mouth gag | Piece | √
3.16 | Dissecting Forceps, Adson-brown 12cm | Piece | √
3.161 | Dissecting Forceps, Gillies 15.5cm | Piece | √
3.162 | Dissecting Forceps, Potts-Smith 16cm | Piece | √
3.163 | Dissecting Forceps, Potts-Smith 18cm | Piece | √
3.164 | Dissecting Forceps, Potts-Smith 23cm | Piece | √
3.165 | Finger protector: langenbeck | Piece | √
3.166 | Ferguson jaw opener | Piece | √
3.167 | Hammers length 230mm, weight 350g. | Piece | √
3.168 | Heister jaw opener - SS | Piece | √
3.169 | Jaw Splint - Small | Piece | √
3.17 | Jaw Splint - Large | Piece | √
3.171 | Lip & Check Retractor - Cawood - Minnesota | Piece | √
3.172 | Lip & Check Retractor - Oringer | Piece | √
3.1 | Mandibular Modular Fixation System with following contents:
3.1.1 | Mandible Trauma Module Set (2.0mm) consisting of: | Set | √
3.1.2 | Mandible Trauma Module Set (2.4mm) consisting of: | Set | √
3.2 | Maxillofacial Modular Fixation System with following contents:
3.2.1 | Module 1.5mm | Piece | √
3.2.2 | Module 2.0mm | Piece | √
3.2.3 | Maxillary awl Obwegeser 15cm | Piece | √
3.2.4 | Maxillary awl Obwegeser 17.5cm | Piece | √
3.2.5 | Mouth gag Mckesson - adult | Piece | √
3.2.6 | Mouth gag Mckesson - Child | Piece | √
3.2.7 | Needle holder 16cm | Piece | √
3.2.8 | Needle holder ,Mathieu,170mm | Piece | √
3.2.9 | Needle holder 23cm | Piece | √
3.2.10 | Obwegeser Retractor 24cm | Piece | √
3.2.11 | Osteotomy Chisel 10mm | Piece | √
3.2.12 | Osteotomy Chisel 15mm | Piece | √
3.2.13 | Periosteal Elevator Joseph | Piece | √
3.2.14 | Ramus retractor | Piece | √
3.2.15 | Retractor Adson | Piece | √
3.2.16 | Rowe's Mandible bone holding forceps | Piece | √
3.2.17 | Rowe's Maxilla disimpaction clamp - left | Piece | √
3.2.18 | Rowe's Maxilla disimpaction clamp - right | Piece | √
3.2.19 | Rowe's zygomatic elevator | Piece | √
3.2.20 | Scissor Iris 11.5cm/ 4.5 inch | Piece | √
3.2.21 | Scissor Metzenbaum - fino 18cm/ 7 inch | Piece | √
3.2.22 | Scissor Tenotomy 115mm | Piece | √
3.2.23 | Sinus Lift curettes 15cm | Piece | √
3.2.24 | Sinus Lift curettes 18.5cm | Piece | √
3.2.25 | Surgical Skin hooks | Piece | √
3.2.26 | Third Molar Retractor Obwegeser 24cm | Piece | √
3.2.27 | Tissue retractor Austin | Piece | √
3.2.28 | Tongue depressor(wider) | Piece | √
3.2.29 | Wire cutter,Straight | Piece | √
3.2.30 | Wire ligature pliers for circum mandibuler | Piece | √
3.2.31 | Wire ligature pliers for circum zygomatic | Piece | √
3.2.32 | Wire pusher Rowe | Piece | √
3.2.33 | Zygoma elevator Gillies 20cm | Piece | √
3.2.34 | Zygomatic bone awl curved/ zygomatic arch awl | Piece | √
3.2.35 | Zygomatic bone hook with T handle | Piece | √
3.2.36 | Maxillofacial Modular Fixation System with following contents: a) Obwegeser Retractor 24cm; b)Third Molar Retractor Obwegeser 24cm | Set | √
3.2.37 | Elcomed, contra angle, straight HP, motor cable and motor holder, surgical hand piece (W&H) | set | √
3.2.38 | Tungsten carbide burs for surgical- Round (2.35mmX30-40mm) | piece | √
3.2.39 | Tungsten carbide burs for surgical-Taper criss-cut (2.35mmX30-40mm) | piece | √
3.2.40 | Tungsten carbide burs for surgical-Cylindrical (2.35mmx30-40mm) | piece | √
3.2.41 | Drill Bits with stops, for Stryker coupling, length- 44.5mm, diameter-1.55mm | piece | √
3.2.42 | Drilling depth 6mm | piece | √
3.2.43 | Drilling depth 8mm | piece | √
3.2.44 | Drilling depth 12mm | piece | √
3.2.45 | Bur block stainless steel (sterilizable) for surgical | Piece | √
3.2.46 | Surgical drill-Stryker COMMAND2 50K DRILL. | piece | √
3.2.47 | Cord for Surgical Drill Stryker command2 –FISHER 104 | piece | √
3.2.48 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 6 mm for stryker coupling j-latch type | piece | √
3.2.49 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 8 mm for stryker coupling j-latch type | piece | √
3.2.50 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 10 mm for stryker coupling j-latch type | piece | √
3.2.51 | Surgical Drill bits with stops a) 1.5 diameter, length 44.5 mm Drilling depth 12 mm for stryker coupling j-latch type | piece | √
3.2.52 | 2 mm Mandible Cortex Screw , Self- tapping,L 6mm,pure titanium , | piece | √
3.2.53 | 2 mm Mandible Cortex Screw , Self- tapping,L 8mm,pure titanium , | piece | √
3.2.54 | 2 mm Mandible Cortex Screw , Self- tapping,L 10mm,pure titanium , | piece | √
3.2.55 | 2 mm Mandible Cortex Screw , Self- tapping,L 12mm,pure titanium , | piece | √
3.2.56 | 2.4 mm Emergency Screw length 6mm,pure titanium | piece | √
3.2.57 | 2.4 mm Emergency Screw length 8mm,pure titanium | piece | √
3.2.58 | 2.4 mm Emergency Screw length 10mm,pure titanium | piece | √
3.2.59 | 2.4 mm Emergency Screw length 12mm,pure titanium | piece | √
3.2.60 | Mandible Mini Plates 2.0,centre space, 4holes,length 28mm,pure titanium | piece | √
3.2.61 | Mandible Mini Plates 2.0,centre space, 6holes,length 40mm,pure titanium | piece | √
3.2.62 | Mandible Mini Plates 2.0,centre space, 20holes,length 119mm,pure titanium | piece | √
3.2.63 | Chin plates 2.0,double bend,5 holes,pure titanium , | piece | √
3.2.64 | Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece | √
3.2.65 | Surgical Drill bit with stops 0.76mm,length 44.5, Drilling depth 5mm,for stryker coupling j-latch type | piece | √
3.2.66 | 1.0mm MF Cortex Screw,Self tapping, Length 4mm,pure titanium, | piece | √
3.2.67 | 1.0mm MF Cortex Screw,Self tapping, Length 5mm,pure titanium, | piece | √
3.2.68 | 1.0mm MF Cortex Screw,Self tapping, Length 6mm,pure titanium, | piece | √
3.2.69 | Double Drill Sleeve 2.4/1.8 | piece | √
3.2.70 | Depth Gauge for 1.5 to 2.0 screws | piece | √
3.2.71 | IMF Screw Set with stainless steel implants with lid, 1. Wire twister | piece | √
3.2.72 | IMF Screw Set with stainless steel implants with lid, 1. Wire Cutter | piece | √
4 | Accessories
4.1 | Ultra sonic scaler insert tips with keys, interdental for sirona scaler - HP, compatible with the (Specify Name; Brand & Model) Ultra sonic scaler. | Piece | √ | √ | √
4.2 | Ultra sonic scaler inserts tips with keys, subgingival for sirona scaler, HP, compatible with the (Specify Name; Brand & Model) Ultra sonic scaler. | Piece | √ | √ | √
4.3 | Turbine hand piece high speed, press button type, 4 holes, Standard | Piece | √ | √ | √
4.4 | Turbine Hand piece high speed, fibre optic, press button, standard head, 4 holes, T1/T2, compatible to the existing Dental chair | Piece | √ | √ | √
4.5 | Turbine hand piece high speed, fibre optic, press button, small head, 4 holes, T1/T2, compatible to the existing Dental chair | Piece | √ | √ | √
4.6 | Fibre optic quick coupling attachment/agent with bulb for hand piece,T1/T2 CHP compatible to the existing Dental Chair | Piece | √ | √ | √
4.7 | Fibre Optic high speed, CHP, Press button, standard head with quick coupling, compatible to the existing Dental chair | Piece | √ | √ | √
4.8 | Exposure / firing switch cat 5154307 with cable for wall mount KODEX, IOPA, X-Ray Machine model no. KODAK 2100, 230 V-240V, 50/60 Hz, 60 Kv, 7 AMP (compatible to the existing X ray machine) | Piece | √ | √ | √
4.9 | CHP Hand piece for Micromotor, compatible with exisiting machine | Piece | √ | √ | √
4.1 | SHP Hand piece for Micromotor, W&H | Piece | √ | √ | √
4.11 | Composite polishing paste | 2 Syringes/ Pkt | √
4.12 | Mineral trioxide aggregate (MTA) powder and liquid | Set | √
4.13 | Tooth forceps comella pattern (to extract wisdom tooth roots) | piece | √
4.14 | Universal Wire Cutter Bandage Scissor, Size 120 mm, SS with golden handle. | piece | √ | √ | √
5 | Oral Surgery and Implants
5.1 | Equipments
5.1.1 | Double Speed Contra Handpiece | Piece | √ | √
5.1.2 | Double Speed Twist Chuck Handpiece/ LVSA Micromotor/MCC8, 8 - Ft.: | Set | √ | √
5.1.3 | Low Speed High Torque Electric Handpiece Systems for Oral Surgery and Implants. | Piece | √ | √
5.2 | Consumables
5.2.1 | Cerclage Wire, 0.6mm dia, without eye, length 175mm | Piece | √ | √
5.2.2 | Surgical Cutting Burs, SS, HM 141 204 040 | Piece | √ | √
5.2.3 | Surgical Cutting Burs, SS, 165 205 023 | Piece | √ | √
5.2.4 | Tungsten carbide SS, HM 141 205 040 | Piece | √ | √
5.2.5 | Tungsten carbide SS, HM 141 205 050 | Piece | √ | √
5.2.6 | Tungsten carbide, HM 33T 104 016 | Piece | √ | √
5.2.7 | Tungsten carbide, HM 151 316 016 | Piece | √ | √
5.2.8 | Tungsten carbide, HM 152 315 014 | Piece | √ | √
5.2.9 | Tungsten carbide, HM 162 104 016 | Piece | √ | √
5.3 | Instrunment
5.3.1 | Bending Pliers with beak, for plates up to 2.7 | Piece | √ | √
5.3.2 | Bending-Cutting Pliers 1.0 to 2.0, right | Piece | √ | √
5.3.3 | Bone Curettes Hemingway 18.5cm | Piece | √ | √
5.3.4 | Bone Curettes Volkmann 170mm | Piece | √ | √
5.3.5 | Bone cutting forcep 18cm | Piece | √ | √
5.3.6 | Bone file 18cm | Piece | √ | √
5.3.7 | Bone Holding forcep 13cm | Piece | √ | √
5.3.8 | Bone Holding forcep 14cm | Piece | √ | √
5.3.9 | Bone Holding forceps 18.5cm | Piece | √ | √
5.3.10 | Bone Holding forceps, self-centering, length 191mm, softlock. | Piece | √ | √
5.3.11 | Bone Reduction forcep with points, narrow 161mm softlock | Piece | √ | √
5.3.12 | Bone Reduction forcep with points, wide 127mm softlock | Piece | √ | √
5.3.13 | Bone Rongeur 12.5cm | Piece | √ | √
5.3.14 | Bone Rongeur 16.5cm | Piece | √ | √
5.3.15 | Bone Rongeurs Luer 14.5cm | Piece | √ | √
5.3.16 | Browne lingual flap retractor | Piece | √ | √
5.3.17 | Chin Retractor Obwegeser 16.5cm | Piece | √ | √
5.3.18 | Condyle retractor - thin - AK's new | Piece | √ | √
5.3.19 | Cleft lip surgery - set | Set | √ | √
5.3.20 | Cleft plate surgery set | Set | √ | √
5.3.21 | Crown wire twisting forceps: 16cms/6 1/4 inch | Piece | √ | √
5.3.22 | Dingmans mouth gag | Piece | √ | √
5.3.23 | Dissecting Forceps, Adson-brown 12cm | Piece | √ | √
5.3.24 | Dissecting Forceps, Gillies 15.5cm | Piece | √ | √
5.3.25 | Dissecting Forceps, Potts-Smith 16cm | Piece | √ | √
5.3.26 | Dissecting Forceps, Potts-Smith 18cm | Piece | √ | √
5.3.27 | Dissecting Forceps, Potts-Smith 23cm | Piece | √ | √
5.3.28 | Finger protector: langenbeck | Piece | √ | √
5.3.29 | Ferguson jaw opener | Piece | √ | √
5.3.30 | Hammers length 230mm, weight 350g. | Piece | √ | √
5.3.31 | Heister jaw opener - SS | Piece | √ | √
5.3.32 | Jaw Splint - Small | Piece | √ | √
5.3.33 | Jaw Splint - Large | Piece | √ | √
5.3.34 | Lip & Check Retractor - Cawood - Minnesota | Piece | √ | √
5.3.35 | Lip & Check Retractor - Oringer | Piece | √ | √
5.4 | Mandibular Modular Fixation System with following contents:
5.4.1 | Mandibular Modular Fixation System with following contents: | Set | √ | √
5.4.2 | Mandible Trauma Module Set (2.4mm) consisting of: | Set | √ | √
5.5 | Maxillofacial Modular Fixation System with following contents:
5.5.1 | Module 1.5mm | Piece | √ | √
5.5.2 | Module 2.0mm | Piece | √ | √
5.5.3 | Maxillary awl Obwegeser 15cm | Piece | √ | √
5.5.4 | Maxillary awl Obwegeser 17.5cm | Piece | √ | √
5.5.5 | Mouth gag Mckesson - adult | Piece | √ | √
5.5.6 | Mouth gag Mckesson - Child | Piece | √ | √
5.5.7 | Needle holder 16cm | Piece | √ | √
5.5.8 | Needle holder ,Mathieu,170mm | Piece | √ | √
5.5.9 | Needle holder 23cm | Piece | √ | √
5.5.10 | Obwegeser Retractor 24cm | Piece | √ | √
5.5.11 | Osteotomy Chisel 10mm | Piece | √ | √
5.5.12 | Osteotomy Chisel 15mm | Piece | √ | √
5.5.13 | Periosteal Elevator Joseph | Piece | √ | √
5.5.14 | Ramus retractor | Piece | √ | √
5.5.15 | Retractor Adson | Piece | √ | √
5.5.16 | Rowe's Mandible bone holding forceps | Piece | √ | √
5.5.17 | Rowe's Maxilla disimpaction clamp - left | Piece | √ | √
5.5.18 | Rowe's Maxilla disimpaction clamp - right | Piece | √ | √
5.5.19 | Rowe's zygomatic elevator | Piece | √ | √
5.5.20 | Scissor Iris 11.5cm/ 4.5 inch | Piece | √ | √
5.5.21 | Scissor Metzenbaum - fino 18cm/ 7 inch | Piece | √ | √
5.5.22 | Scissor Tenotomy 115mm | Piece | √ | √
5.5.23 | Sinus Lift curettes 15cm | Piece | √ | √
5.5.24 | Sinus Lift curettes 18.5cm | Piece | √ | √
5.5.25 | Surgical Skin hooks | Piece | √ | √
5.5.26 | Third Molar Retractor Obwegeser 24cm | Piece | √ | √
5.5.27 | Tissue retractor Austin | Piece | √ | √
5.5.28 | Tongue depressor(wider) | Piece | √ | √
5.5.29 | Wire cutter,Straight | Piece | √ | √
5.5.30 | Wire ligature pliers for circum mandibuler | Piece | √ | √
5.5.31 | Wire ligature pliers for circum zygomatic | Piece | √ | √
5.5.32 | Wire pusher Rowe | Piece | √ | √
5.5.33 | Zygoma elevator Gillies 20cm | Piece | √ | √
5.5.34 | Zygomatic bone awl curved/ zygomatic arch awl | Piece | √ | √
5.5.35 | Zygomatic bone hook with T handle | Piece | √ | √
5.5.36 | Maxillofacial Modular Fixation System with following contents: a) Obwegeser Retractor 24cm; b)Third Molar Retractor Obwegeser 24cm | Set | √ | √
5.5.37 | Elcomed, contra angle, straight HP, motor cable and motor holder, surgical hand piece (W&H) | set | √ | √
5.5.38 | Tungsten carbide burs for surgical- Round (2.35mmX30-40mm) | piece | √ | √
5.5.39 | Tungsten carbide burs for surgical-Taper criss-cut (2.35mmX30-40mm) | piece | √ | √
5.5.40 | Tungsten carbide burs for surgical-Cylindrical (2.35mmx30-40mm) | piece | √ | √
5.5.41 | Drill Bits with stops, for Stryker coupling, length- 44.5mm, diameter-1.55mm | piece | √ | √
5.5.42 | Drilling depth 6mm | piece | √ | √
5.5.43 | Drilling depth 8mm | piece | √ | √
5.5.44 | Drilling depth 12mm | piece | √ | √
5.5.45 | Bur block stainless steel (sterilizable) for surgical | Piece | √ | √
6 | Orthodontics
6.1 | Consumables
6.1.1 | Orthocryl Powder Ploymer, Green | 1 kg/ pkt | √
6.1.2 | Orthocryl liquid Monomer / Green | 50ml | √
6.1.3 | Orthocryl Powder Ploymer, Neon | 1kg/pkt | √
6.1.4 | Orthocryl liquid Monomer / Neon | 50ml | √
6.1.5 | Orthocryl Powder Ploymer, Purple | 1kg/pkt | √
6.1.6 | Orthocryl liquid Monomer / Purple | 50ml | √
6.1.7 | Orthocryl Powder Ploymer , Yellow | 1kg/pkt | √
6.1.8 | Orthocryl liquid Monomer/ Yellow | 50ml | √
6.1.9 | Low fusing silver soldering flux | spool | √
6.1.10 | Anti-flux | 50g | √
6.1.11 | Self-Igniting Butane Micro-Torch (refillable): gas torch that creates a precise flame emitting from a nozzle head, allowing direct heat application to an object. | pcs | √
6.1.12 | Pressurized refill butane Gas for Self-Igniting Butane Micro-Torch | 250ml/can | √
6.1.13 | Mouth gag rubber, pediatric | Piece | √ | √
6.1.14 | Orthocryl Powder Ploymer 1 kg/ pkt, Green | Packet | √ | √
6.1.15 | Orthocryl liquid Monomer / Green | 50ml | √ | √
6.1.16 | Orthocryl Powder Ploymer, Neon | 1kg/pkt | √ | √
6.1.17 | Orthocryl liquid Monomer / Neon | 50ml | √ | √
6.1.18 | Orthocryl Powder Ploymer, Purple | 1kg/pkt | √ | √
6.1.19 | Orthocryl liquid Monomer / Purple | 50ml | √ | √
6.1.20 | Orthocryl Powder Ploymer , Yellow | 1kg/pkt | √ | √
6.1.21 | Petralit cement, (Silico-phosphate cement), powder 60 gm & liquid 22 ml | Set | √ | √ | √
6.2 | Equipments
6.2.1 | Dual wheel model trimmer,ower supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top. | Piece | √
6.2.2 | Essix Vacuum Thermoforming Machine, Power supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top. Appropriate voltage stabilizer should be provided with the unit. | Piece | √
6.2.3 | Sterilization sleave sealer machine, Power supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top. | Piece | √
6.2.4 | 022 Damon brackets with cuspid hook | Set | √
6.2.5 | 3-dimensional screw – (Bertoni) | Piece | √
6.2.6 | APC- adhesive precoated brackets MBT 022 slot | Set | √
6.2.7 | Arch hooks Rectangular | Packet | √
6.2.8 | Ball end clasps 028, 032 | Packet | √
6.2.9 | Bicuspid bands kit, upper right & left, with 022 bracket prewelded | Packet | √
6.2.10 | Bicuspid bands kit, lower right & left, with 022 bracket prewelded | Packet | √
6.2.11 | Bite turbos | Packet | √
6.2.12 | Bondable buttons | Packet | √
6.2.13 | Bondable lingual retainer 3 x 3 | Packet | √
6.2.14 | Bondable lingual retainer 4 x 4 | Packet | √
6.2.15 | Bracket storage box (for individual brackets) | Box | √
6.2.16 | Brass wire | Spool | √
6.2.17 | Clarity Ceramic bracket 5 - 5 (022 slot) | Set | √
6.2.18 | Copper NiTi 017x0.025 upper | Packet | √
6.2.19 | Copper NiTi 017x0.025 lower | Packet | √
6.2.20 | Copper NiTi 0.014 inch upper | Packet | √
6.2.21 | Copper NiTi 0.014 inch lower | Packet | √
6.2.22 | Copper NiTi 0.016 inch upper | Packet | √
6.2.23 | Copper NiTi 0.016 inch lower | Packet | √
6.2.24 | Copper NiTi 0.018 inch upper | Packet | √
6.2.25 | Copper NiTi 0.018 inch lower | Packet | √
6.2.26 | Crimpable hook 019 x 025 | Packet | √
6.2.27 | Cuspid bands kit, upper,left/right 022 bracket prewelded, Kit/s | Pack | √
6.2.28 | Cuspid bands kit, lower,left/right,022 bracket prewelded, Kit/s | Pack | √
6.2.29 | E-chain elastics closed, black | Spool | √
6.2.30 | E-chain elastics closed, blue | Spool | √
6.2.31 | E-chain elastics closed, Clear | Spool | √
6.2.32 | E-chain elastics closed, Green | Spool | √
6.2.33 | E-chain elastics closed, orange | Spool | √
6.2.34 | E-chain elastics closed, pink | Spool | √
6.2.35 | E-chain elastics closed, purple | Spool | √
6.2.36 | E-chain elastics closed, red | Spool | √
6.2.37 | E-chain elastics continuous, Clear | Spool | √
6.2.38 | E-chain elastics continuous, Gray | Spool | √
6.2.39 | E-chain elastics open, Clear | Spool | √
6.2.40 | E-chain elastics open, Gray | Spool | √
6.2.41 | Elastic separators | Stick | √
6.2.42 | Elastic thread 50ft length | Spool | √
6.2.43 | Etchant gel (9.6% hydrofluoric acid) 4 syringes of 1.2ml and 8 pre-bent applicator tips. | Piece (Syringe of 1.2 ml) | √
6.2.44 | Etchant gel (38% phosphoric acid) with 50 pre –bent applicator tips, Thixotropic | Piece (Syringe of 25 ml) | √
6.2.45 | Expansion screw safety keys | Packet | √
6.2.46 | Extraoral elastics 1/2" 8 OZ | Packet | √
6.2.47 | Extraoral elastics 1/2" 14 Oz | Packet | √
6.2.48 | Extraoral elastics ¼” 8 Oz | Packet | √
6.2.49 | Extraoral elastics ¼” 14 Oz | Packet | √
6.2.50 | Extraoral elastics 3/16” 8 Oz | Packet | √
6.2.51 | Extraoral elastics 3/16” 14 Oz | Packet | √
6.2.52 | Extraoral elastics 3/8” 8 Oz | Packet | √
6.2.53 | Extraoral elastics 3/8” 14 Oz | Packet | √
6.2.54 | Extraoral elastics 5/16” 8 Oz | Packet | √
6.2.55 | Extraoral elastics 5/16” 14 Oz | Packet | √
6.2.56 | Fan-type expansion screw | Piece | √
6.2.57 | Flexible Double Ended plastic self-retaining Cheek Retractor, Paedo Small | Piece | √
6.2.58 | Forsus Kit (3M) | Pkt | √
6.2.59 | Gemini Roth 022 brackets with 4 tubes (UT/LD) | Set | √
6.2.60 | Glass Ionomer Band Cement | Pkt | √
6.2.61 | Headgear kit | Pkt | √
6.2.62 | Hyrax expansion screw, 11mm | Piece | √
6.2.63 | Hyrax expansion screw, 13mm | Piece | √
6.2.64 | Hyrax expansion screw, 8mm | Piece | √
6.2.65 | Hyrax expansion screw, 9mm | Piece | √
6.2.66 | Intraoral elastics ¼" 3.5 Oz | Pkt | √
6.2.67 | Intraoral elastics ¼” 6 Oz | Pkt | √
6.2.68 | Intraoral elastics 1/8" 2 Oz; | Pkt | √
6.2.69 | Intraoral elastics 1/8” 3.5 Oz | Pkt | √
6.2.70 | Intraoral elastics 3/16” 3.5 Oz | Pkt | √
6.2.71 | Intraoral elastics 3/8” 2 Oz | Pkt | √
6.2.72 | Intraoral elastics 3/8” 3.5 Oz | Pkt | √
6.2.73 | Intraoral elastics 5/16” 2 Oz | Pkt | √
6.2.74 | Intraoral elastics 5/16” 3.5 Oz | Pkt | √
6.2.75 | Intraoral elastics 5/16” 4.5 Oz | Pkt | √
6.2.76 | Intraoral stops 2 mm length | Pkt | √
6.2.77 | Intraoral stops 1mm length | Pkt | √
6.2.78 | Lingual cleats | Pkt | √
6.2.79 | Lip bumper with loops | Pkt | √
6.2.80 | Lower expander screw | Piece | √
6.2.81 | Mandibular bow screw | Piece | √
6.2.82 | Micro sectional screw | Piece | √
6.2.83 | Mini Twin Vs 022 brackets | Set | √
6.2.84 | NiTi 0.014 lower | Pkt | √
6.2.85 | NiTi 0.014 upper | Pkt | √
6.2.86 | NiTi 0.016 lower | Pkt | √
6.2.87 | NiTi 0.016 upper | Pkt | √
6.2.88 | NiTi 0.017x0.025 lower | Pkt | √
6.2.89 | NiTi 0.017x0.025 upper | Pkt | √
6.2.90 | NiTi 0.018 lower | Pkt | √
6.2.91 | NiTi 0.018 upper | Pkt | √
6.2.92 | NiTi 0.019 x 0.025 lower | Packet | √
6.2.93 | NiTi 0.019 x 0.025 upper | Packet | √
6.2.94 | NiTi closed coil long | Spool | √
6.2.95 | NiTi closed coil medium | Spool | √
6.2.96 | NiTi closed coil small | Spool | √
6.2.97 | NiTi Closed coil spring- heavy force 200g | Packet | √
6.2.98 | NiTi Closed coil spring- light force 100g | Packet | √
6.2.99 | NiTi Closed coil spring- medium force 150g | Packet | √
6.2.100 | NiTi distalizing (stop wound coil) spring | Packet | √
6.2.101 | NiTi open coil spring - heavy | Packet | √
6.2.102 | NiTi open coil spring - light | Packet | √
6.2.103 | NiTi open coil spring - medium | Packet | √
6.2.104 | NOLA dry field system-cheek retractor, Tongue Guards,"Y" Connector, Suction tubes and adapter | Set | √
6.2.105 | O-ring modules Black | Pkt | √
6.2.106 | O-ring modules Blue | Pkt | √
6.2.107 | O-ring modules Clear | Pkt | √
6.2.108 | O-ring modules Green | Pkt | √
6.2.109 | O-ring modules Light Blue | Pkt | √
6.2.110 | O-ring modules Maroon | Pkt | √
6.2.111 | O-ring modules Neon Green | Pkt | √
6.2.112 | O-ring modules Orange | Pkt | √
6.2.113 | O-ring modules Pink | Pkt | √
6.2.114 | O-ring modules Purple | Pkt | √
6.2.115 | O-ring modules Red | Pkt | √
6.2.116 | O-ring modules Silver | Pkt | √
6.2.117 | O-ring modules Teal | Pkt | √
6.2.118 | O-ring modules Tooth color | Pkt | √
6.2.119 | O-ring modules Violet | Pkt | √
6.2.120 | O-ring modules Yellow | Pkt | √
6.2.121 | Patient relief wax | Pkt | √
6.2.122 | Preformed Kobayashi ligature ties | Packet | √
6.2.123 | Preformed Left lower first molar band with cleats kit, lower double tube | Set | √
6.2.124 | Preformed Left upper first molar band with cleats kit, triple tube | Set | √
6.2.125 | Preformed ligature long ties | Packet | √
6.2.126 | Preformed Right lower first molar band with cleats kit,lower double tube | Set | √
6.2.127 | Preformed Right upper first molar band with cleats kit, triple tube | Set | √
6.2.128 | Preformed second molar band kits lower left, single tube | Set | √
6.2.129 | Preformed second molar band kits lower right, single tube | Set | √
6.2.130 | Preformed second molar band kits upper left, single tube | Set | √
6.2.131 | Preformed second molar band kits upper right, single tube | Set | √
6.2.132 | Preformed short ligature ties | Packet | √
6.2.133 | Rapid molar distalizer | Packet | √
6.2.134 | Retraction screw - medium | Piece | √
6.2.135 | Retraction screw - Mini | Piece | √
6.2.136 | Rotation wedges | Pkt | √
6.2.137 | Sectional screw - Medium | Piece | √
6.2.138 | Sectional screw - Mini | Piece | √
6.2.139 | Stainless steel wire 0.014 lower | Packet | √
6.2.140 | Stainless steel wire 0.014 upper | Packet | √
6.2.141 | Stainless steel wire 0.016 lower | Packet | √
6.2.142 | Stainless steel wire 0.016 upper | Packet | √
6.2.143 | Stainless steel wire 0.016 x 0.022 lower | Packet | √
6.2.144 | Stainless steel wire 0.016 x 0.022 upper | Packet | √
6.2.145 | Stainless steel wire 0.017 x 0.022 lower | Packet | √
6.2.146 | Stainless steel wire 0.017 x 0.022 upper | Packet | √
6.2.147 | Stainless steel wire 0.018 lower | Pkt | √
6.2.148 | Stainless steel wire 0.018 upper | Pkt | √
6.2.149 | Stainless steel wire 0.018x0.025 lower | Pkt | √
6.2.150 | Stainless steel wire 0.018x0.025 upper | Pkt | √
6.2.151 | Stainless steel wire 0.019x0.025 lower | Pkt | √
6.2.152 | Stainless steel wire 0.019x0.025 upper | Pkt | √
6.2.153 | Stainless steel wire 0.021x0.025 lower | Packet | √
6.2.154 | Stainless steel wire 0.021x0.025 upper | Packet | √
6.2.155 | Stop/spacer 2mm - 022 | Pkt | √
6.2.156 | Stop/spacer 2mm - 022 x 028. | Packet | √
6.2.157 | Stop/spacer 2mm - 030 | Packet | √
6.2.158 | Stripping Strips, Diamond studded | Packet | √
6.2.159 | Thermoforming materials - Essix A+ 1.5mm | Pkt | √
6.2.160 | Thermoforming materials - Essix A+ 1mm | Pkt | √
6.2.161 | Thermoforming materials - Essix Embrace 1mm | Pkt | √
6.2.162 | Thermoforming materials - Essix Embrace 2mm | Pkt | √
6.2.163 | Thermoforming materials - Essix Night Guard | Pkt | √
6.2.164 | Thermosealer - 10mm, Role | Roll | √
6.2.165 | Thermosealer - 5mm, Role | Roll | √
6.2.166 | Thermosealer - 7.5mm, Role | Roll | √
6.2.167 | Tissue guard (clear), 0.027 inch | Spool | √
6.2.168 | TMA preformed pendulum springs – left and right | Packet | √
6.2.169 | TMA Reverse curve with T-loops lower 0.017x0.025 (broad) | Packet | √
6.2.170 | TMA Reverse curve with T-loops lower 0.016x0.022 (broad) | Packet | √
6.2.171 | TMA Reverse curve with T-loops upper 0.017x0.025 (broad) | Packet | √
6.2.172 | TMA Reverse curve with T-loops upper 0.016x0.022 (broad) | Packet | √
6.2.173 | TMA Upper 0.017x0.025 | Pkt | √
6.2.174 | TMA Lower 0.017x0.025 | Pkt | √
6.2.175 | TMA upper 0.016x0.022 | Packet | √
6.2.176 | TMA lower 0.016x0.022 | Packet | √
6.2.177 | TMA upper large 0.018, U 016 x 022 | Packet | √
6.2.178 | TMA lower large 0.018, L 016 x 022 | Packet | √
6.2.179 | Tooth stripper files | Packet | √
6.2.180 | Traction screw | Piece | √
6.2.181 | Transbond adhesive kit plus primer + Etchant, self cure | Set | √
6.2.182 | Transbond kit (light cure), plus primer and etchant | Piece | √
6.2.183 | Trapezoidal Expansion screw | Piece | √
6.2.184 | Universal expansion screw - Maxi | Piece | √
6.2.185 | Universal expansion screw - medium | Piece | √
6.2.186 | Universal expansion screw - mini | Piece | √
6.2.187 | Weldable buttons | Packet | √
6.2.188 | Wilcock SS round wire, 012 Premium plus | Spool | √
6.2.189 | Wilcock SS round wire, 014 Premium plus | Spool | √
6.2.190 | Wilcock SS round wire, 016 Premium plus | Spool | √
6.2.191 | Wilcock SS round wire, 018 Premium plus | Spool | √
6.2.192 | Wilcock SS round wire, 020 Premium plus | Spool | √
6.2.193 | Orthocryl Powder Ploymer, Green | 1 kg/ pkt | √
6.2.194 | Orthocryl liquid Monomer / Green | 50ml | √
6.2.195 | Orthocryl Powder Ploymer, Neon | 1kg/pkt | √
6.2.196 | Orthocryl liquid Monomer / Neon | 50ml | √
6.2.197 | Orthocryl Powder Ploymer, Purple | 1kg/pkt | √
6.2.198 | Orthocryl liquid Monomer / Purple | 50ml | √
6.2.199 | Orthocryl Powder Ploymer , Yellow | 1kg/pkt | √
6.2.200 | Orthocryl liquid Monomer/ Yellow | 50ml | √
6.2.201 | Low fusing silver soldering flux | spool | √
6.2.202 | Anti-flux | 50g | √
6.2.203 | Self-Igniting Butane Micro-Torch (refillable): gas torch that creates a precise flame emitting from a nozzle head, allowing direct heat application to an object. | pcs | √
6.2.204 | Adhesive removing pliers with diamond inset | Piece | √
6.2.205 | Bird beak pliers with cutter | Piece | √
6.2.206 | R.E.D. Distraction Hallo | Set | √
6.2.207 | Arch hook crimping pliers | Piece | √
6.2.208 | Arch wire turret | Piece | √
6.2.209 | Band contouring pliers (Johnson) | Piece | √
6.2.210 | Band cutting scissors-curved | Piece | √
6.2.211 | Band slitting pliers | Piece | √
6.2.212 | Bite plane thermoplier | Piece | √
6.2.213 | Bite stick (band adapter) autoclavable | Piece | √
6.2.214 | Bonding Organization tray | Piece | √
6.2.215 | Bracket holder with alignment tip | Piece | √
6.2.216 | Cinch back pliers | Piece | √
6.2.217 | Convertible cap removal tool kit with replacement tip | Piece | √
6.2.218 | Direct bracket remover pliers | Piece | √
6.2.219 | Distal end cutter safety hold | Piece | √
6.2.220 | Distal twister | Piece | √
6.2.221 | Double ended band seater | Piece | √
6.2.222 | Elastic hook-forming thermoplier | Piece | √
6.2.223 | Essix IPR Deluxe kit Raintree USA | Piece | √
6.2.224 | Heavy wire bending pliers (Headgear pliers set) | Piece | √
6.2.225 | How plier | Piece | √
6.2.226 | Jarabak plier | Piece | √
6.2.227 | Ligature cutter small | Piece | √
6.2.228 | Ligature cutter standard | Piece | √
6.2.229 | Ligature tucker | Piece | √
6.2.230 | Mallet handle band pusher (Mershon) | Piece | √
6.2.231 | Mandibular tooth moving thermoplier | Piece | √
6.2.232 | Maxillary tooth moving thermoplier | Piece | √
6.2.233 | Mesial-distal thermoplier | Piece | √
6.2.234 | Micro ramp thermoplier | Piece | √
6.2.235 | Mini Mathieu needle holder | Piece | √
6.2.236 | Nance Tie Back Plier | Piece | √
6.2.237 | Optical pliers | Piece | √
6.2.238 | Orthodontic Impression tray, Rim Locked, Paed lower L1 | Piece | √
6.2.239 | Orthodontic Impression tray, Rim Locked, Paed lower L2 | Piece | √
6.2.240 | Orthodontic Impression tray, Rim Locked, Paed lower L3 | Piece | √
6.2.241 | Orthodontic Impression tray, Rim Locked, Paed lower L4 | Piece | √
6.2.242 | Orthodontic Impression tray, Rim Locked, Paed Upper U1 | Piece | √
6.2.243 | Orthodontic Impression tray, Rim Locked, Paed Upper U2 | Piece | √
6.2.244 | Orthodontic Impression tray, Rim Locked, Paed Upper U3 | Piece | √
6.2.245 | Orthodontic Impression tray, Rim Locked, Paed Upper U4 | Piece | √
6.2.246 | Orthodontic Impression trays perforated, rim locked adult lower L1 | Piece | √
6.2.247 | Orthodontic Impression trays perforated, rim locked,adult lower L2 | Piece | √
6.2.248 | Orthodontic Impression trays perforated, rim locked adult lower L3 | Piece | √
6.2.249 | Orthodontic Impression trays perforated, rim locked adult lower L4 | Piece | √
6.2.250 | Orthodontic Impression trays perforated, rim locked adult, upper U1 | Piece | √
6.2.251 | Orthodontic Impression trays perforated, rim locked adult, upper U2 | Piece | √
6.2.252 | Orthodontic Impression trays perforated, rim locked adult, upper U3 | Piece | √
6.2.253 | Orthodontic Impression trays perforated, rim locked adult, upper U4 | Piece | √
6.2.254 | Photo Cheek retractors, large, set of 3 | Piece | √
6.2.255 | Photo Cheek retractors, Small, set of 3 | Piece | √
6.2.256 | Posterior band removing pliers | Piece | √
6.2.257 | Rectangular arch forming pliers (Tweed) | Piece | √
6.2.258 | Slot forming thermoplier | Piece | √
6.2.259 | Stackable Band Organizer | Piece | √
6.2.260 | Surgical hook crimping pliers | Piece | √
6.2.261 | Torquing pliers | Piece | √
6.2.262 | Tweed loop forming pliers | Piece | √
6.2.263 | Undercut Enhancing thermoplier | Piece | √
6.2.264 | Utility arch/loop forming pliers (Rickettts) | Piece | √
6.2.265 | V-bending pliers | Piece | √
6.2.266 | Weingart utility pliers | Piece | √
6.2.267 | Wire contouring pliers | Piece | √
7 | Paedodontic
7.1 | Consumables
7.1.1 | Acidulated Phosphate Fluoride(APF) 1.23% gel(12,300 ppmF) strawberry flavor 17.5 oz | Bottle of 17.5 Oz | √ | √
7.1.2 | Sodium Fluoride rinse 0.2% (920 ppmF) | 1Ltr/ bottle | √ | √
7.1.3 | Sodium Fluoride rinse 0.05%(230 ppmF) | 250ml/ bottle | √ | √
7.1.4 | Fluoride Tablets 0.25mg | 1000 Tabs/ Pkt | √ | √
7.1.5 | Fluoride Tablets 0.50mg | 1000 Tabs/ Pkt | √ | √
7.1.6 | Sodium Fluoride Varnish 5% ( 2.26% F)10 ml tube with applicator brush | 50tubes/ pkt | √ | √
7.1.7 | Precontoured Primary Molar Bands | Set | √ | √
7.1.8 | Precontoured Permanent First Molar bands | Set | √ | √
7.1.9 | Stainless steel crowns for primary dentitions(Pre-contoured) | Set | √ | √
7.1.10 | Crown scissors, curved serrated blade 4" | Piece | √ | √
7.1.11 | Crown Scissors,straight,110 mm | Piece | √ | √
7.1.12 | Transparent Pediatric Strip Crown Forms | Set | √ | √
7.1.13 | Molt Mouth Gag, external with silicon tip, child/infant, 11cm | Piece | √ | √
7.1.14 | Molt Mouth Gag, external with silicon tip, medium,14 cm | Piece | √ | √
7.1.15 | Polycarbonite lip cheek retractor with wings 0118WC (child) | Box of 3 pieces | √ | √
7.1.16 | Polycarbonite lip cheek retractor with wings 01189WM(Medium) | Box of 3 pieces | √ | √
7.1.17 | Papoose board padded, small with chest strap and abdominal flaps (30” x 7”) | Piece | √ | √
7.1.18 | Papoose board padded, regular with chest strap and abdominal flaps (36” x 11”) | Piece | √ | √
7.1.19 | Dual Arch Fluoride plastic tray, Pedo(aqua) | Set | √ | √
7.1.20 | Dual Arch Fluoride plastic tray, Large (blue) |  | √ | √
7.1.21 | Dual Arch Fluoride plastic tray, medium (White) | Set | √ | √
7.1.22 | Dual Arch Fluoride plastic tray, small (yellow) | Set | √ | √
7.1.23 | DSP IOPA X-ray film holder DR540501(Snap-A-Ray) | 1pcs/ pkt | √ | √
7.1.24 | X-ray apron with collar, pediatric 0.3mm (21" x 22") | Piece | √ | √
7.1.25 | Rubber dam 5”x5”, thin green, Mint (Box) | 52pcs/ box | √ | √
7.1.26 | Rubber dam clamp, SS White,No.206 | Set | √ | √
7.1.27 | Rubber dam clamp, SS White,No.207 | Set | √ | √
7.1.28 | Rubber dam clamp, SS White No. 208 | Set | √ | √
7.1.29 | Rubber dam clamp No.12 | DOPC | √ | √
7.1.30 | Rubber dam clamp No. 13 | Set | √ | √
7.1.31 | Rubber dam clamp, Ivory,No.14 | Set | √ | √
7.1.32 | Rubber Dam Clamp No.0 | Set | √ | √
7.1.33 | Rubber Dam Clamp No.209 | Set | √ | √
7.1.34 | Rubber dam clamp, Ivory, No.8-A | Set | √ | √
7.1.35 | Rubber dam clamp, Ivory, No. 14-A | Set | √ | √
7.1.36 | Rubber dam punch, Ainsworth 160mm | Piece | √ | √
7.1.37 | Rubber Dam Clamp Forceps, Ivory 170mm | Piece | √ | √
7.1.38 | Rubber dam frame(Young’s) 5”x5” | Piece | √ | √
7.1.39 | Abell Contouring pliers N0.112 | Piece | √ | √
7.1.40 | Reynolds contouring pliers 125mm/5” | Piece | √ | √
7.1.41 | Gordon Contouring Pliers 5¼ inch. | Piece | √ | √
7.1.42 | Johnson contouring pliers 135mm/5¼ “. | Piece | √ | √
7.1.43 | Contouring pliers ultra thin No.007-118 | Piece | √ | √
7.1.44 | Crimping Pliers,No.417 | Piece | √ | √
7.1.45 | How plier,straight,135mm | Piece | √ | √
7.1.46 | Transparent Polyester Matrix, 0.05 mm/0.02” thick x 59 ft | Roll | √ | √
7.1.47 | soft stainless matrix, 0.025 mm/0.001” thick x20 foot | Roll | √ | √
7.1.48 | Ferric sulphate 15.5% dropper bottle of 59ml(2oz) | 59ml/ bottle | √ | √
7.1.49 | Reinforced ZOE cement (IRM) Powder 40g,15ml liquid | Set | √ | √
7.1.50 | Needle, 1" 30G (0.40x21mm) short 25 mm | 100pcs/ pkt | √ | √
7.1.51 | Benzocaine topical anesthetic gel 20% 1 oz/29.6ml jar ,raspberry flavour | 25pcs/ pkt | √ | √
8 | Prostodontic
8.1 | Consumables
8.1.1 | Temporary ST. steel 3M , readymade crowns, standard set/pkt, 3M/Bosworth | Set | √ | √ | √
8.1.2 | Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-597-00 | Piece | √ | √ | √
8.1.3 | Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-598-00 | Piece | √ | √ | √
8.1.4 | Cloth polishing buff SHP mounted,Denta, 181-100-01 | Piece | √ | √ | √
8.1.5 | Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-582-00 | Piece | √ | √ | √
8.1.6 | Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-583-00 | Piece | √ | √ | √
8.1.7 | Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-584-00 | Piece | √ | √ | √
8.1.8 | Tungsten carbide, Precision-cut Bur SHP, Dentaurum no. 123-583-00 | Piece | √ | √ | √
8.1.9 | Optra-Dam clamp-free rubber dam. Blue, Box of 5o, Ivoclar | 50sheets/ Pkt | √ | √ | √
8.1.10 | Optra-Gate medium, Box of 5o, Ivoclar | 50/kt | √ | √ | √
8.1.11 | Bur Silicone acrylic small, flame shaped. Shofu- HP/0423-6/Pcs, Grit-Course-Gray | Packet/set | √ | √ | √
8.1.12 | Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0427- Grit-Medium Brown | 6pcs/ pkt | √ | √ | √
8.1.13 | Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0428- Grit-Fine Light-Gray | 6pcs/ pkt | √ | √ | √
8.1.14 | Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0426-6/Pcs, Grit-Course-Gray | Packet/set | √ | √ | √
8.1.15 | Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0427-6/Pcs, Grit-Medium Brown | Packet/set | √ | √ | √
8.1.16 | Bur Silicone acrylic, Large, flame shaped, Shofu- HP/0428-6/Pcs, Grit-Fine Light-Gray | Packet/set | √ | √ | √
8.1.17 | Silicone polishing Cone, Green medium grit, Small | Packet/set | √ | √
8.1.18 | Silicone polishing Cone, Green, fine grit, Small | Packet/set | √ | √
8.2 | Instrunment
8.2.1 | Hanau Alcohol hand Torch plastic for prostho-lab waxing | Piece | √ | √ | √
8.2.2 | Hoopler duplicator, upper/lower set,for CD | Set | √ | √ | √
8.2.3 | Willis gauze for CD for OVD measurement | Piece | √ | √ | √
8.2.4 | Silicone-Rubber Model Base former U/L set (Large) | Set | √
8.2.5 | Silicone-Rubber Model Base former U/L set (Medium) | set | √
8.2.6 | Silicone-Rubber Model Base former U/L set (Small) | set | √
9 | List for Porcelain-Ceramic Work (consumables)
9.1 | Porcelain Vita dentin B1 | 50gm/ pkt | √
9.2 | Porcelain Vita dentin B2 | 50gm/ pkt | √
9.3 | Porcelain Vita dentin B3 | 50gm/ pkt | √
9.4 | Porcelain Vita dentin B 3.5 | 50gm/ pkt | √
9.5 | Porcelain Vita dentin B 4 | 50gm/ pkt | √
9.6 | Porcelain Vita dentin C1 | 50gm/ pkt | √
9.7 | Porcelain Vita dentin C2 | 50gm/ pkt | √
9.8 | Porcelain Vita dentin C3 | 50gm/ pkt | √
9.9 | Porcelain Vita dentin C3.5 | 50gm/ pkt | √
9.1 | Porcelain Vita dentin C4 | 50gm/ pkt | √
9.11 | Porcelain Vita Opaque A1 | 50gm/ pkt | √
9.12 | Porcelain Vita Opaque A2 | 50gm/ pkt | √
9.13 | Porcelain Vita Opaque A3 | 50gm/ pkt | √
9.14 | Porcelain Vita Opaque A3.5 | 50gm/ pkt | √
9.15 | Porcelain Vita Opaque A4 | 50gm/ pkt | √
9.16 | Porcelain Vita Opaque B1 | 50gm/ pkt | √
9.17 | Porcelain Vita Opaque B2 | 50gm/ pkt | √
9.18 | Porcelain Vita Opaque B3 | 50gm/ pkt | √
9.19 | Porcelain Vita Opaque B3.5 | 50gm/ pkt | √
9.2 | Porcelain Vita Opaque B4 | 50gm/ pkt | √
9.21 | Porcelain Vita Opaque C1 | 50gm/ pkt | √
9.22 | Porcelain Vita Opaque C2 | 50gm/ pkt | √
9.23 | Porcelain Vita Opaque C3 | 50gm/ pkt | √
9.24 | Porcelain Vita Opaque C3.5 | 50gm/ pkt | √
9.25 | Porcelain Vita Opaque C4 | 50gm/ pkt | √
9.26 | Porcelain Vita Opaque dentine | 50gm/ pkt | √
9.27 | Porcelain Vita Wash Opaque | 50gm/ pkt | √
9.28 | Gutta –cut, rechargeable unit; 220-230 voltage, complete sets | set | √
9.29 | Ever-stick fabre, GC/Crown &Bridge,1.5mm/pkt | pkt | √
9.3 | Ever-stick fabre, GC/Perio, 1.5mm/pkt | pkt | √
9.31 | Ever-stick fabre, GC/post & core systems, 1.5mm/pkt | pkt | √
9.32 | Compomer TM Brilliant Veneer systems complete sets, Wheledent, pkts | pkt | √
9.33 | Crown cutter carbide burs, 8G, Striaght fissure, Wheledent | pcs | √
9.34 | Magic foam cord- Gingival retraction complete sets with Dispenser and syringe /pkts, | 10/ pkts | √

## Dermatology
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Dermatology Department
1 | Equipments
1.1 | Carbon dioxide (CO2) Laser and accessories | Set | √
1.2 | Fibro Optic (spotlight) for enhancing Leisons in genitals. | Set | √
1.3 | RF-300 (Radio Frequency) and accessories | Set | √
1.4 | Ultra Violet Ray-A (UVA)-Chamber and accessories | Set | √
1.5 | Fluorescent Pulsed Light (FPL-Plasmalite) and accessories | Set | √
1.6 | Dermatoscope- Dermlite DL3 or Welsh Allyn | Set | √
1.7 | Patch test kit - International standard series, with finn chambers | Set | √
2 | Consumables
2.1 | Biopsy Punch (which is sterile, disposable skin biopsy punches ranging from 5mm to 8mm) | Pkt | √
2.2 | Biopsy punch - disposable, size 2mm | Pkt
2.3 | Biopsy punch - disposable, size 3mm | Pkt
2.4 | Biopsy punch - disposable,size 4mm | Pkt
2.5 | TCA (Trichlor Acitic Acid) with applicator | Jar | √
3 | Instrunment
3.1 | Instrument automatic washer/disinfectors | Set | √
3.2 | Ultrasonic cleaning basin, integrated in a work bench | Piece | √
3.3 | Ultrasonic cleaner with lifting system for trays, with lid | Set | √
3.4 | Ultravoilet Woods lamp set for | Set | √
3.5 | Cryo surgery instruments with accessories and liquid nitrogen (very small liquid nitrogen barrel) Power supply 220-240v, 50Hz. Suitable stabilizer. Power cord with round 3 pin top | Set | √
3.6 | Skin hook | Piece | √
3.7 | Comedo extracter | Piece | √
3.8 | Skin curette | Piece | √
3.9 | Iris scissor | Piece | √
4 | Gradle scissor | Piece | √
4.1 | Baby Metzenbaum undermining scissor | Piece | √
4.2 | UV diagnostic Wood's lamp | set | √

## Emergency
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Emergency Department
Equipments and consumables added with the OT and aneathesia section

## ENT
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
ENT Department
1 | Equipments
1.1 | Xenon Cold Light Fountain Compact unit with built in 300 Watt XENON lamp with manually variable light intensity from 0 - 100 %, Provides a light intensity which meets all requirements even for use with fiberscope or narrow-lumen rigid endoscopes, Color Temperature of 6000 Kelvin for daylight illumination, Highest Color Rendering Index (CRI) for naturally color recognition, Maximum light intensity immediately after power up, Fan for silent air-ventilation Novel type of light source adapter for improved heat dissipation at the proximal ends of fiber optic light cables.Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √ | √
1.2 | Head light With ball joint and vertical height adjustment, front section movable in all directions, fiber optic light cable included, with fully adjustable comfortable head band. Power cord with round 3 pin top | Set | √ | √
1.3 | Otoscope With rubber bulb and spare light bulb with inbuilt rechargeable battery | Set | √ | √ | √ | √
1.4 | Rhino-Laryngo-Fibre scope Instrument channel 1.2 mm direction of view 0 degree, angle of view 110 degree, working length 34 cm, outer diameter 3.5 mm.Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √
1.5 | Diagnostic Audiometer, Available Signals: Stimulus: Pure tone, warble tone, 2 external inputs and USB audimetry. MIC input for live speech audiometry.Masking: NBN, WN, SN SIGNALS SPECIFICATION Attenuator step: 1 and 5 dB, Presentation: Continuous, Pulsed (0.5, 1 and 2 Hz) Warble: 5 Hz sin wave modulating signal AVAILABLE OUTPUTS AND TRANSDUCERS AC: Headphones (TDH-39 / DD45) HF headphones (HDA-300) Insert earphones (ER-3A / ER-5A) BC: B-71 bone vibrator Free field Insert masking earphone: IME-100 AVAILABLE TESTS • Pure Tone audiometry (HL and UCL)• Auto threshold (modified Hughsone-Westlake) • HF audiometry: from 8 to 20 kHz (optional) • Speech audiometry (SRT, WRS, MCL, and UCL) • Speech audiometry (SRT, WRS, MCL, and UCL) • SISI test • 2 independent channels Master Hearing Aid • QuickSIN® test (optional) • Visual Reinforcement Audiometry (VRA) and Conditioned Play Audiometry (CPA) test (optional) with video reinforces INCLUDED PARTS • TDH-39 supra-aural headphones • B71 bone vibrator • Patient response switch • Monitor headset with boom microphone • Clip-on microphone for patient-to-operator communication • Medical grade power supply • USB connection cable • User manual • Inventis Software Suite CD • Noah Maestro module • Daisy software and Daisy Maestro module • Four magnetic covers OPTIONAL ACCESSORIES (with order code) | Set | √ | √
1.6 | OPTIONAL ACCESSORIES • ER-3A insert earphones • ER-5A insert earphones • IME-100 insert masking earphone • DD45 complete headset • Amplivox Audiocups noise excluding enclosures for TDH- 39 and DD45 headphones • Desktop microphone for live speech tests • Cable set for sound booths with jack sockets • Cable set for sound booths with RCA sockets • One active speaker FBT J 5A (230V model) • One active speaker FBT J 5A (120V model) |  | √
1.7 | Play audiometer | Set | √
1.8 | OTOREAD SCREENER TEOAE Technical Specifications • Measurement type: OtoAcoustic Emissions • Frequency range DPOAE: 1.5, 2, 2.5, 3, 3.5, 4, 5, 6, 8, 10, 12 kHz • Frequency range TEOAE: 0.7, 1, 1.4, 1.5, 2, 2.5, 2.8, 3.5, 4 kHz• Stimulus intensity: 40 to 70 dB SPL (DPOAE). 83 dB SPL (TEOAE) • Maximum output (Protection): 90 dB (This level is well within OSHA permissible limits of 90 dBA for 8 hours) • Microphone system noise:-20 dB SPL @ 2 KHz(1 Hz bandwidth)-13 dB SPL @ 1 KHz (1 Hz bandwidth) • Battery life: Approximately 300 tests • Display: LCD-display 4 line x 10 character • Weight: 300 g/ 10.6 oz. including batteries • Internal software languages: English, German, French, Spanish, Russian • Screener 6 bands (1 fixed protocol) 1.5-4 kHz Standard accessories • Handheld OtoRead™ unit including probe cord and cradle • Printer including power supply and power cable • Printer Cable • Carrying Bag • Probe cord for extension (100 cm/39 inches) • 2 Thermal rolls paper • Box of 146 eartips (10 size; 3, 4, 6, 7, 8, 9, 10, 11, 12, 13 mm) • 4 probe tips • 4 AA/UW3/R6 alkaline batteries | Set | √ | √ | √
1.9 | Tympanometer (Ipsilateral & contralateral) | Set | √ | √
1.1 | Brainstem evoked response audiometer(ABR), complete set with software | Set | √
1.11 | portable Audiometer: ●Audiometric Headset ● Bone Conductor ● Patient Response Button●External Power Supply ●Audiogram Charts ●Operation Manual CD ●Multilingual CE Instructions for Use ● Amplivox Audiocups, noise reducing headset ●The lightweight carrying case will hold the as well as a noise reducing headset and audiogram charts etc. ●Pure, Warble and Pulsed tones ●Automatic Threshold test ● SISI and Bekesy tests ● Stenger and ABLB tests ●Tone in Noise test (Langenbeck) ● 125Hz - 8000Hz ……….(125-250hz to 8000hz) ●-10dB to 120dBHL output ● Air and Bone Conduction ● WN and NB masking ●5dB and 1dB attenuators ● Insert Phones (optional) | Set | √ | √
1.12 | Hearing Aid analyzer/programmer specifications: ●Frequency Range: 125 to 12500 Hz ● Stimuli: Pink, whilte, speech noise, ISTS, tone sweep ●Stimulus Level: 50-90 dB SPL, 1 or 5 dB steps ● Frequencies per Octave (sweep): 24, 12, 6, 3 ●Frequency Resolution: 1/24, 1/12, 1/6, 1/3 ● Dynamic Range:> 100 dB ● Coupler Mic Intensity:20-150 dB ●Reference Mic Intensity:20-100 dB ●Power:15V / 2 Amp DC ●Size: (W x D x H) 13.7 x 13.7 x 5.5 in. (348 x 348 x 140 mm.) ●Weight:13 lbs. (5.9 kg.) PC SYSTEM REQUIREMENTS: ●Pentium-class or equivalent PC with 1.6 GHz or faster processor ●2 Gb RAM (more is recommended) ●CD-ROM reader ●Screen resolution of 1024x768 or higher ●Available USB port ●150 Mb of available hard disk space ●NOAH 3.7 or later ●Microsoft Windows XP (SP3), Windows 7 or 8 | Set | √
1.13 | HOPKINS Straight Forward Telescope: 0 degree diameter 2.9 mm, length 6 cm | Piece | √
1.14 | HOPKINS Straight Forward Telescope: 0 degree diameter 4 mm, length 18 cm | Piece | √ | √
1.15 | HOPKINS Straight Forward Telescope: 30 degree diameter 2.9 mm, length 6 cm | Piece | √
1.16 | HOPKINS Straight Forward Telescope: 30 degree diameter 4 mm length 18 cm | Piece | √ | √
1.17 | HOPKINS Straight Forward Telescope: 70 degree diameter 4 mm length 18 cm | Piece | √ | √
1.18 | Dr. Speech (software version 4) : With all standard accessories | Set | √
1.19 | Naso- meter II - Model 6400 | Set | √
1.2 | Triveni Speech Trainer: TST- 100 | Set | √
1.21 | Auditory Brainstems Response-complete set | Piece | √
1.22 | Telecam SL II Camera Control unit: Color systems PAL/NTSC, with integrated digital Image Processing Module, Power Supply 100-240 VAC 50/60 Hz. Consisting of : 20213020 Telecam SL II Camera Control Unit 400A Mains cord 2x20221070 Connecting cable, for controlling peripheral devices, Lengh,180cm, 536MK BNC/BNC Video Cable, Length 180cm, 20040082, DV Connecting Cable, 6pin to 4 pin length 450cm.Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √ | √
1.23 | Unidrive multifunction unit for ENT surgery | Piece | √ | √
1.24 | Cordless rechargeable LED head light with charger | Set | √ | √
1.25 | Mastoid drill micromotor - It should be compatible with Unidrive multifunction unit for ENT surgery | Set | √ | √
1.26 | Mastoid drill hand piece - It should be compatible with Unidrive multifunction unit for ENT surgery | Set | √ | √
1.27 | Shaver system for paranasal sinuses micro-shaver handpiece, shaver blade straight, shaver blade curved: (It should be compatible with Unidrive multifunction unit for ENT surgery) Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin top | Set | √ | √
1.28 | Binocular surgical loupe, magnification of 3.5 X, HRP optics, working distance of 420 mm, with extended field of view , adjustable optics right and left to select exact inter-pupillary distance, with spectacle frame and LED lighting plus other accessories. | set | √
1.29 | Mobile suction unit with wide mouthed 2 X 2 Ltr jars,mtr. capacity of - 700 mm Hg +/- 10 regulable , flutter free vaccum control knob, tubing of 8mm ID X 2 | set | √ | √
1.3 | Foot suction: with capacity of -600mmHg t 10 at 200m1 Per Stroke, Heavy Piston Pump, Non - Collapsible Suction Tubing, 1500mml capacity jar wilh over flow system & ABS Lid | set | √ | √
1.31 | Operating ENT Microscope | set | √ | √
2 | Consumables
2.1 | SHEEHY, Ventilation tube, Teflon, sterile, ID 1.25 mm. | Piece | √
2.2 | Absorbable Gelatin Sponge | Piece | √
2.3 | Portex cuffed tracheostomy tubes, Size: 8.0mm | Piece | √
2.4 | Portex cuffed tracheostomy tubes, Size: 7.5mm | Piece | √
2.5 | Portex cuffed tracheostomy tubes: Size: 7.0mm | Piece | √
2.6 | Portex cuffed tracheostomy tubes: Size: 6.0mm | Piece | √
2.7 | Portex cuffed tracheostomy tubes: Size: 4.0mm | Piece | √
2.8 | Hearing aids: Behind the ear (Power- mild) | Piece | √
2.9 | Hearing aids, Behind the ear (Power - moderate | Piece | √
2.1 | Hearing aids, Behind the ear (Power - high) | Piece | √
2.11 | Hearing aids, Body worn (Power - mild) | Piece | √
2.12 | Hearing aids, Body worn (Power -moderate) | Piece | √
2.13 | Hearing aids- Body worn (Power - high) | Piece | √
2.14 | Portex cuffled tracheostomy tubes 5mm OD | Piece | √
2.15 | Portex cuffed tracheostomy tubes,3.0 mm OD | Piece | √
2.16 | Portex tracheostomy tubes. 2.5 mm OD | Piece | √
2.17 | Shiley Tracheostomy Tube Cuffless with Disposable Inner Cannula Fenestrated, size 4, ID 5mm, length 62mm | Piece | √
2.18 | Shiley Tracheostomy Tube Cuffless with Disposable Inner Cannula Fenestrated, size 4, ID 5mm, length 62mm | Piece | √
2.19 | Shiley™ Tracheostomy Tube Cuffless with Disposable Inner Cannula Fenestrated, size 8, ID 7.6mm, length 79mm | Piece | √
2.2 | Shiley™ Tracheostomy Tube XLT Cuffless Distal with Disposable Inner Cannula, ID 5mm, Total length 90mm, Distal extension | Piece | √
2.21 | Shiley™ Tracheostomy Tube XLT Cuffless Distal with Disposable Inner Cannula, ID 7mm, Total length 100mm, Distal extention | Piece | √
2.22 | Shiley™ Tracheostomy Tube XLT Cuffless Distal with Disposable Inner Cannula, ID 7mm, Total length 100mm, Proximal extention | Piece | √
3 | Instrunment
3.1 | Adaptor for suction cannula: With cut off hole, LUER cone length 10cm | Piece | √
3.2 | Barany Noise Apparatus, with ear olive, SS | Piece | √
3.3 | Beales elevator : Curved, width 1.5 mm, length 16 cm, SS | Piece | √
3.4 | Belluci scissors blade: Length 8 mm, working length 8 cm, SS | Piece | √
3.5 | Belluci scissors blade: Length 8 mm, working length 8 cm, curved to left, SS | Piece | √
3.6 | adaptor for suction cannula: Length 8 mm, working length 8 cm, curved to right, SS | Piece | √
3.7 | Billeau ear loop: Length 16 cm - size 1, SS | Piece | √ | √ | √
3.8 | Billeau ear loop: Length 16 cm - size 2, SS | Piece | √ | √ | √
3.9 | Ear Hook: 90 degree, blunt, SS | Piece | √ | √ | √
3.1 | Fisch articulated retroauricular retractor, 3x3 prongs, sharp, length 17cm, SS | Piece | √
3.11 | FISCH Ear Forcep, Delicate, pointed, SS | Piece | √
3.12 | FISCH Scissor, Delicate, curved, sharp/blunt, length 10.5 cm, SS | Piece | √
3.13 | Hartmann ear dressing forcep, SS | Piece | √ | √ | √
3.14 | Hartmann ear forcep: Very fine, serrated, working length 10 cm, SS | Piece | √ | √ | √
3.15 | HARTMANN Ear Speculum, Size O.D 3 mm, SS | Piece | √ | √ | √ | √
3.16 | HARTMANN Ear Speculum, Size O.D 4 mm, SS | Piece | √ | √ | √ | √
3.17 | HARTMANN Ear Speculum, Size O.D 5 mm, SS | Piece | √ | √ | √ | √
3.18 | HARTMANN Ear Speculum, Size O.D 6 mm, SS | Piece | √ | √ | √ | √
3.19 | HARTMANN Ear Speculum, Size O.D 7.5 mm, SS | Piece | √ | √ | √ | √
3.2 | HOLMGREEN self retaining endaural speculum, Outer diameter 5 mm , SS | Piece | √
3.21 | House Dieter malleus nipper, Upbiting, working length 8 cm, SS | Piece | √
3.22 | Instrument rack, 22cm x 12.5 cm, with holder for 5 micro instruments and 5 drills, SS | Piece | √
3.23 | Jobson-Horne probe, With serrated ring and end, length 18 cm, SS | Piece | √ | √ | √ | √
3.24 | LEMPERT Elevator, Length 19 cm, 3 mm wide, SS | Piece | √
3.25 | PLESTER Ear retractor 2x2 prongs, SS | Piece | √
3.26 | PLESTER Elevator, Length 18 cm, 8 mm wide, SS | Piece | √
3.27 | PLESTER Knife, Round, vertical, length 16 cm, standard size 3.5x2.5mm, SS | Piece | √
3.28 | PLESTER retractor with biprong blade on left and solid blade on right, length 11cm, SS | Piece | √
3.29 | PLESTER retractor with biprong blade on right and solid blade on left, length 11cm, SS | Piece | √
3.3 | PLESTER sickle knife, Double edge, length 16cm, standard, medium curved, SS | Piece | √
3.31 | PLESTER Sickle knife, Double edge, length 16 cm, standard, slightly curved, SS | Piece | √
3.32 | PLESTERS sickle knife, Standard, medium curved, length 16 cm, SS | Piece | √
3.33 | POLITZER Ear Knife, Working length 16 cm, curved, SS | Piece | √
3.34 | Seeker with ball end, Angled 45 degree, ball end diameter 1mm | Piece | √
3.35 | SHEA Vein Press, Length 13 cm | Piece | √
3.36 | SPRATT Mastoid Curette, Length 17 - size 1, SS | Piece | √
3.37 | SPRATT Mastoid Curette, Length 17 - size 2, SS | Piece | √
3.38 | SPRATT Mastoid Curette, Length 17 - size 4, SS | Piece | √
3.39 | Suction cannula: Angular LUER lock outer diameter 0.5 mm, working length 6 cm | Piece | √
3.4 | Suction cannula: Angular LUER lock outer diameter 0.7 mm, working length 6 cm | Piece | √
3.41 | Suction cannula: Angular LUER lock outer diameter 1.0 mm, working length 6 cm | Piece | √
3.42 | Suction tube: Conical, LUER, outer diameter 3.5 mm, working length 13 cm | Piece | √
3.43 | Suction tube: Conical, LUER, outer diameter 4.5 mm, working length 13 cm | Piece | √
3.44 | Suction tube: Cylindrical LUER, outer diameter 0.5 mm, working length 9 cm | Piece | √
3.45 | Suction tube: Cylindrical LUER, outer diameter 0.7 mm, working length 9 cm | Piece | √
3.46 | Suction tube: Cylindrical LUER, outer diameter 1.0mm, working length 9 cm | Piece | √
3.47 | TRAUTMANN Ear Knife, Working length 19 cm, Delicate Lance, SS | Piece | √
3.48 | Tuning Fork C-512, SS | Piece | √ | √ | √
3.49 | WAGENER Ear Hook: 90 degree, blunt, SS | Piece | √ | √ | √
3.5 | WULLSTEIN ear forcep: Very fine, oval cupped jaws, working length 12.5 cm, diameter 1mm, SS | Piece | √
3.51 | WULLSTEIN ear forcep, Very fine, oval cupped jaws, working length 8 cm, diameter 0.9 mm, curved to left, SS | Piece | √
3.52 | WULLSTEIN ear forceps, Very fine, oval cupped jaws, working length 8 cm, diameter 0.9 mm, curved to right, SS | Piece | √
3.53 | WULLSTEIN ear forceps, Very fine, oval cupped jaws, working length 8 cm, diameter 0.9 mm, curved upwards, SS | Piece | √
3.54 | WULLSTEIN Needle, Slight curve, length 16.5 cm, SS | Piece | √
3.55 | WULLSTEIN Retractor, 3x3 prongs, length 13 cm, SS | Piece | √
3.56 | WULLSTEIN Scissor, Curved, sharp/blunt, length 14 cm, SS | Piece | √
3.57 | Anterior Pillar Retractor and tonsil dissector with tooth | Piece | √
3.58 | Bronchoscope Tube Universal, Length 43cm - size 6.5 | Piece | √
3.59 | Bronchoscope Tube Universal, Length 43cm - size 7.5 | Piece | √
3.6 | Bronchoscope Tube Universal, Length 43cm - size 8.5 | Piece | √
3.61 | Cotton Carrier, Length 35 cm | Piece | √
3.62 | Cotton Carrier , Length 55 cm | Piece | √
3.63 | Cotton Carrier , with spring handle length 35 cm | Piece | √
3.64 | Cotton Carrier, with spring handle length 55 cm | Piece | √
3.65 | DAVIS - MEYER Mouth Gag, SS | Piece | √
3.66 | Draffin Suspension bipod for use with tongue blade and mouth gags | Piece | √
3.67 | Esophagoscope Alligator, Grasping, double action jaws: Diam 2.5 mm working length 55 cm, SS, for use with oesophagoscope | Piece | √
3.68 | Esophagoscope circular biopsy : Diam 2.5 mm working length 55 cm, SS, for use with oesophagoscope | Piece | √
3.69 | Esophagoscope denture cutting forceps: Diameter 2.5 mm working length 55 cm, for use with oesophagoscope SS | Piece | √
3.7 | Esophagoscope peanut grasping forceps: Diam 2.5 mm working length 55 cm, for use with oesophagoscope | Piece | √
3.71 | Esophagoscope Tube, Length 20 cm - size 3 | Piece | √
3.72 | Esophagoscope Tube, Length 30 cm - size 4 | Piece | √
3.73 | Esophagoscope Tube, Length 30 cm - size 6 | Piece | √
3.74 | Esophagoscope Universal biopsy grasping forcep: Diam 2.0 mm Working length 55 cm, for use with oesophagoscope | Piece | √
3.75 | Esophagoscope Universal biopsy grasping forcep: Diam 2.0 mm Working length 35 cm, for use with laryngoscope | Piece | √
3.76 | Eve's tonsillar snare, Length 27 cm, SS | Piece | √
3.77 | Forceps circular cup, biopsy, working length 35 cm, for use with laryngoscope | Piece | √
3.78 | Forceps Alligator Grasping double action jaws, Diam 1.5 mm, working length35 cm, SS for use with laryngoscope | Piece | √
3.79 | Forceps Alligator Grasping double action jaws, Diam 2.0 mm, working length 55 cm, SS, for use with laryngoscope | Piece | √
3.8 | Forceps Peanut Grasping double action jaws, Diam 2.0 mm working length 35 cm, for use with laryngoscope | Piece | √
3.81 | Fuller's Trachea Tube with fenestration, Sterling silver - size 28 | Piece | √
3.82 | Fuller's Trachea Tube with fenestration, Sterling silver - size 30 | Piece | √
3.83 | Fuller's Trachea Tube with fenestration, Sterling silver - size 36 | Piece | √
3.84 | Fuller's Trachea Tube with fenestration, Sterling silver - size 32 | Piece | √
3.85 | Fuller's Trachea Tube with fenestration, Sterling silver - size 34 | Piece | √
3.86 | Halle Birkett Tonsillar artery forceps, curved, 20 cm, SS | Piece | √
3.87 | Insulated cannula for suction/coagulation, with ergonomic handle, insulated and fitting for unipolar coagulation, working length 23 cm, diameter 3.5 cm | Piece | √
3.88 | KLEINSASSER Cutting Forcep, curved left with 2 mm cupped jaws, working length 23 cm, SS | Piece | √
3.89 | KLEINSASSER Cutting Forcep, curved right with 2 mm cupped jaws, working length 23 cm, SS | Piece | √
3.9 | KLEINSASSER Cutting Forcep, straight, with 2 mm cupped jaws, working length 23 cm, SS | Piece | √
3.91 | Kleinsasser Laryngeal knife, oval straight, SS | Piece | √
3.92 | Kleinsasser laryngeal sickle knife, curved pointed, SS | Piece | √
3.93 | KLEINSASSER Operating Laryngoscope, ength 13 cm, child size, small | Piece | √
3.94 | KLEINSASSER Operating Laryngoscope, length 15 cm, child size | Piece | √
3.95 | KLEINSASSER Operating Laryngoscope, length 18 cm, adult size, medium | Piece | √
3.96 | KLEINSASSER Scissor, curved left working length 23 cm, SS | Piece | √
3.97 | KLEINSASSER Scissor, curved right working length 23 cm, SS | Piece | √
3.98 | KLEINSASSER Scissor , straight working length 23 cm, SS | Piece | √
3.99 | KLEINSASSER Scissor, Angled 45 degree, Working length 23 cm, SS | Piece | √
3.4 | Laryngeal Mirror with handle, size: 5 | Piece | √ | √ | √
3.41 | Laryngeal Mirror with handle, size: 6 | Piece | √ | √ | √
3.42 | Laryngeal Mirror with handle, size: 7 | Piece | √ | √ | √
3.43 | Laryngeal Mirror with handle, size: 8 | Piece | √ | √ | √
3.44 | Laryngeal Mirror with handle, size: 4 | Piece | √ | √ | √
3.45 | Laryngoforce II scissor, curved right, cutting blade 10 mm, working length 23 cm, SS | Piece | √
3.46 | Laryngoforce II scissor, left, cutting blade 10 mm, working length 23 cm, SS | Piece | √
3.47 | Laryngoforce II scissor, straight, cutting blade 10 mm, working length 23 cm, SS | Piece | √
3.48 | Laryngoscope holder and chest support | Piece | √
3.49 | Metzenbaum scissor, curved, length 18 cm, SS | Piece | √
3.5 | Negus Adenoid Curette, curved, length 21cm, with cage - size 2, SS | Piece | √
3.51 | Negus Adenoid Curette, curved, length 21cm, with cage - size 3, SS | Piece | √
3.52 | Negus Adenoid Curette, curved, length 21cm, with cage - size 4, SS | Piece | √
3.53 | Negus forcep, large curved, 20 cm, SS | Piece | √
3.54 | Negus ligature carrier | Piece | √
3.55 | Rigid suction tube with rubber tip, straight, working length 35 cm | Piece | √
3.56 | Rigid Suction Tubes with rubber tip, straight working length 55 cm | Piece | √
3.57 | ROBERTS - JESBERG Hypopharyngoscope, length 20 cm | Piece | √
3.58 | ROBERTS - JESBERG Oval Operating Esophagoscope, length 30 cm | Piece | √
3.59 | ROBERTS - JESBERG Oval Operating Esophagoscope, length 50 cm | Piece | √
3.6 | RUSSELL - DAVIS Tongue Blade with central groove, size: 1 | Piece | √
3.61 | RUSSELL - DAVIS Tongue Blade with central groove, size: 2, SS | Piece | √
3.62 | RUSSELL - DAVIS Tongue Blade with central groove, size: 3, SS | Piece | √
3.63 | RUSSELL - DAVIS Tongue Blade with central groove, size: 4, SS | Piece | √
3.64 | Tele-Laryngo-Pharyngoscope with HOPKINS, Lateral telescope 90 degree | Piece | √
3.65 | Tracheal dilator, two bladed, length 15 cm, SS | Piece | √
3.66 | Universal Biopsy and Grasping forcep, working length 35 cm, diameter 2 mm, for use with laryngoscope | Piece | √
3.67 | Universal Biopsy and Grasping forcep, working length 55 cm, diameter 2.5 mm, for use with oesophagoscope | Piece | √
3.68 | YANKAUER Suction Tube, length 29 cm | Piece | √
3.69 | COTTLE Metal Mallet, SS | Piece | √
3.7 | COTTLE scissor, 10.5 cm, curved, SS | Piece | √
3.71 | CRAIG Nasal Bone Forcep, straight, SS | Piece | √
3.72 | Frazier suction tube, with cut off hole and stylet, angled, working length 10 cm, outer diameter 2 mm | Piece | √
3.73 | Frazier suction tube, with cut off hole and stylet, angled, working length 10 cm, outer diameter 3 mm | Piece | √
3.74 | Freer Elevator sharp and blunt, length 20 cm, SS | Piece | √
3.75 | Freer's suction elevator, with stylet, length 19 cm | Piece | √
3.76 | Heyman nasal scissor, medium, standard model, length 9.5 cm, SS | Piece | √
3.77 | Joseph Elevator, slightly curved, sharp, length 17.5 cm, SS | Piece | √
3.78 | Joseph nasal saw, bayonet type, left, SS | Piece | √
3.79 | Joseph nasal saw, bayonet type, right, SS | Piece | √
3.8 | Joseph retractor, length 15.5 cm, SS | Piece | √
3.81 | JOSEPH scissor, 14 cm, curved, SS | Piece | √
3.82 | Kerrison ronguer, size: 2, SS | Piece | √
3.83 | Killian- Claus septal chisel, V - shaped cutting edge, length 16.5 cm, SS | Piece | √
3.84 | Killian- osteotome | Piece | √
3.85 | KILLIAN STRUYCKEN Nasal Speculum with set screw, length 15 cm, blade length - 65 mm, SS | Piece | √
3.86 | KILLIAN STRUYCKEN Nasal Speculum with set screw, length 15 cm, blade length 40 mm, SS | Piece | √
3.87 | KILLIAN STRUYCKEN Nasal Speculum with set screw, length 15 cm, blade length 90mm, SS | Piece | √
3.88 | Thudicum Nasal Speculum, SS | Piece | √ | √ | √
3.89 | Bipolar Coagulation Forceps, with suction canal, nasal, blunt, insulated, bayonet shaped, length 16 cm, SS | Piece | √
3.9 | Blakesley nasal forceps, 45degrees upturned, through cutting, width 3.5 mm, length 11 cm, SS | Piece | √
3.91 | Blakesley Rhinoforce II Ethmoid forceps, straight, size 1, length 17 cm, SS | Piece | √
3.92 | Blakesley Wilde nasal forcep, 90 degrees upturned, width 3.5 mm, length 11cm, SS | Piece | √
3.93 | Blakesley Wilde nasal forcep, straight, width 2.5 mm, length 11 cm, SS | Piece | √
3.94 | Blakesley Wilde Rhinoforce II, Ethmoid forceps 45 degree upwards curved, length 17 cm, SS | Piece | √
3.95 | Blakesley Wildenasal forceps, straight, width 4.8 mm, length 11 cm, SS | Piece | √
3.96 | Citelli antrum punch through cuttung, size: 1.5 mm, working length 17 cm, SS | Piece | √
3.97 | Eicken Antrum cannula, Luer lock, long curved, working length 11 cm, outer diameter 3 mm, SS | Piece | √
3.98 | Grunwald-Henke Nasal Cutting forceps, length 11 cm , SS | Piece | √
3.99 | Heuwieser antrum grasping forcep, fixed jaw curved 90 degrees, curved downwards, movable jaw backward opening 1200, jaw curved downwards, length 10 cm, SS | Piece | √
3.1 | Lichtwitz antrum trochar needle Luer lock, outer diameter 1.5 mm, working length 10 cm | Piece | √
3.101 | Nasal Sickle knife, pointed, length 19.5cm | Piece | √
3.102 | Probe, double ended, maxillary sinus ostium seeker, ball shaped, length 19cm, SS | Piece | √
3.103 | Rhinoforce Nasal scissors, curved to left, length 13 cm, SS | Piece | √
3.104 | Rhinoforce Nasal scissors, curved to right, length 13 cm, SS | Piece | √
3.105 | Rhinoforce Nasal scissors, straight, length 13 cm, SS | Piece | √
3.106 | Stammberger Antrum punch, backward cutting, sheath 360 degree rotatable with fixing screw, working length 10 cm with cleaning adaptor, SS | Piece | √
3.107 | Takahashi Nasal Forceps, straight, length 11cms, SS | Piece | √
3.108 | Trocar and cannula for Sinoscopy, O.D 3.3 mm, length of the cannula 7.5 cm, oblique beak, SS | Piece | √
3.109 | Trocar and cannula for Sinoscopy, O.D 4 mm, length of the cannula 7.5 cm, oblique beak, SS | Piece | √
3.11 | Trocar and cannula for Sinoscopy, O.D 5 mm, length of the cannula 8.5 cm, fenestrated beak, SS | Piece | √
3.111 | Bronchoscope tube universal, 4.5 mm | Piece | √
3.112 | Lack's tongue depressor | Piece | √ | √ | √ | √ | √
3.113 | Tilley's nasal dressing forceps | Piece | √ | √ | √ | √
3.114 | Noise generator device (tinnitus masker) | Piece | √
3.115 | Vancouver Optical Forceps for pediatric bronchoscopy, diameter 2.9mm, length 36cm, with spring action handle for controlled removal of peanuts and soft foreign bodies | Piece | √
3.116 | Bipolar Coagulation Forcep, 16 cm, pointed, SS | Piece | √ | √
3.117 | Stammberger Telescope handle, flat, standard model, length 11 cm for use with HOPKINS telescope 4 mm, length 18 cm | Piece | √
3.118 | Tumarkin metal aural speculum slotted, set of 8, Nos 0 to 7, black finish. | Piece | √
3.119 | Halle Birkett Tonsillar artery forceps, straight, 20 cm, SS | Piece | √
3.12 | Dennis Brown tonsil holding forceps | Piece | √
3.121 | Tonsillar scissors, curved | Piece | √
3.122 | Lacrimal punctum dilator | Piece | √
3.123 | Tilley Lichtwitz trocar and cannula for antral wash out | Piece | √
4 | Accessories
4.1 | Fiber Optic light cable, size 3.5cm, | Piece | √
4.2 | Fiber Optic light cable, size 3.5cm, Length 230 cm | Piece | √
4.3 | Diamond straight shaft burrs with smooth shaft, length 7 cm, size 006 - 070, set of 15, SS | Piece | √
4.4 | Straight shaft burrs, Tungsten Carbide, length 7 cm, sizes 006 - 070, set of 14, SS | Piece | √
4.5 | Adaptors from Bronchoscope to any respiration equipment | Piece | √
4.6 | BENJAMIN -HAVAS Clip for proximal illumination | Piece | √
4.7 | Flexible Brush for Cytology Handle, Length 35 cm, SS | Piece | √
4.8 | Kleinsasser handle for use with above two instruments, SS | Piece | √
4.9 | Prismatic Light Deflector for use with laryngoscope | Piece | √
4.1 | Specimen collector for bronchoscopy | Piece | √
4.11 | Protecting Tube for HOPKINS telescopes | Piece | √
4.12 | Relevant cables and tubes, For Unidrive multifunction unit | Set | √
4.13 | Higginson syringe for antral wash out | Pcs | √
4.14 | Bulbs for operating microscope, Compatible to existing operating microscop | Piece | √
4.15 | Bulbs for Otoscope, Compatible to existing otoscope | Piece | √ | √ | √ | √
4.16 | Bulbs for Cold light source, Compatible to existing | Piece | √
4.17 | Polysomnography. Complete set. Sleep study |  | √
4.18 | Hartmann ear forcep Alligator type (Crocodile forcep), Very fine, serrated, working length 12 cm, SS
4.19 | Hartmann ear forcep Alligator type (Crocodile forcep), Very fine, serrated, working length 10 cm, SS
4.2 | Audiometer accessories head phone, TDH 39 |  | √
4.21 | Audiometer accessories BC vibrator, B71 |  | √
4.22 | Sterilizer dental Autoclave, Table top, Power supply 220-240v, 50Hz. Operating environmental condition: Temperature -10 to 40 ºC, Humidity > 80% , Suitable Voltage stabilizer having high/low cut off with delay timer, Power cord with round 3 pin top. |  | √ | √ | √
4.23 | Back biter for FESS surgey |  | √ | √
4.24 | Crocodile forcep 7cm long |  | √ | √ | √ | √
4.25 | Wax hook |  | √ | √ | √ | √
4.26 | 90 degree hook |  | √ | √ | √ | √
4.27 | Tillis forceps |  | √ | √ | √ | √

## Gyne obst
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Gyne & Obst Department
1 | Equipments
1.1 | Hysteroscope: Telescope 0 & 30 degree, autoclaveable, Telescope sheaths, Operating accessories (stainless steel); (Biopsy, forceps, scissors, graspers, coagulation electrode, colen's knief, roller ball - 6 Piece, Endovision display system - colour monitor (14"). Medical grade monitor with three chip camera, Xenon light source with six spare bulbs, Hysteromat: automatic infusion, suction device for infusing fluids in uterus from 0.1 to 3.0 litre/mintue. air bubble detection facility. Containers for collecting outflowing fluids, Recording facility for operative procedure. Servo voltage stabilizer should be supplied. Power Supply: AC 220-230V . Power cord with round 3 pin top. | Set | √ | √
1.2 | Bed for Obstetric and Labour : with Epoxy coated. I.V Rod (72x28x32"), with Raxine mattress, Adjustable height by hydraulic operated foot pump, kick bucket with ss bowl, leggings 2 nos, handrest 2 nos and with railing with adjustable head end and detachable lower part. | Set |  |  | √ | √
1.3 | Portable video colposcope: Class II, Camera for medical use. Pixel : 4 Million Pixel. CCD : 1/1.8" High Dnsity CDD. Image size : 2272 x 1704, 1024 x 68. Lens : 4 Times Zoom. Focus Distance : F=7.85~32.0mm (35mm Format calculation 38~155. Iris : F5.1 - F10.3. Lens composition : 8 Clusters 10 Pieces: 1. Digital Zoom: Maximum 4 Times. Autofocus : TTL AF. Shot Distance : 30 Cm~oo. Memory Tool : 16M Compact Flash Card. Sidelight Form : TTL. Shutter speed : 1 - 1/2300 Second. Interface : USB. Requester Program 1 copy. Power supply: AC 220V- 230V, Power cord with round 3 pin top and suitable voltage stabilizer. To be supplied with LCD monitor. | Set | √ | √ | √ |  | To be supplied in district hospitals where there are Gynecologists.
1.4 | CTG (cardiotocograph) machine twin monitor: with facility for fetal heart rate, uterine contraction recording in single and twin pregnancies. FHR to be displayed from 0-240 beats/minute, uterine contraction to be displayed from 0 - 100 mmHg '- Facility for maternal BP, ECG and Pulse oxymetry (optional) '- Facility for fetal pulse oximetry (optional)'- Ultrasonic transducers (four) with Piece machine '- Tocotransducers (two) with Piece machine '- Remote event marker '- Facility for alarm in case of very low and high fetal heart rate '- Facility for hoOK up to central nursing station (optional)'- Facility for hoOK up to central comptuer facility to store data on hard disk '- Thermal paper to last one year supply. CTG table.Power supply: AC 220V- 230V, to be supplied with two cardio probes. Power cord with round 3 pin top and suitable voltage stabilizer ( delete) | Set | √ | √ | √ |  | supplied to where there are gynecologist
1.5 | Doppler fetal heart monitor: with numerical display-2.5MHz, speaker 0.6W output, 50-280bpm, AC230V, Single phase, with rechargeable battery, DC9.6V-12V | Set | √ | √ | √ | √
1.6 | Laparoscopic set : with all accessories for laparoscopic surgery including family planning | Set | √ | √
1.7 | LEEP Machine: RSF - 4000K (220 - 240 volts). Mode - Cut, Cut & amp; Coag, Coag, Fulguration, Bipolar. Output voltage - Variable from 500 to 600 +_ 20%. Output power - 100 W. Output Frequency - Monopolar mode > 1 MHz. Bipolar mode > 1MHz. Mode & amp; output Waveform. Cut (Fully Filtered). Cut & amp; Coag (Fully Rectified). Coag (Partially Rectified). Fulguration (Spark - Gap). Blend. Class I type BF. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √ |  | Supplied to where there are gynecologist at DH
1.8 | Cryotherapy set with different sizes probes connection pipe to carbon dioxide cylinder with adaptor with accessories (cryoprobes for ectocervix) | Set | √ | √ | √ |  | Supplied to where there are gynecologist at DH
1.9 | Ultrasound Unit: Perinatal Medicine colour Doppler, 3.5 MHz probe 5-10 MHz probe, Transvaginal probe, Printer, Cart for the Unit, CINE Loop Facility, Clinically analysis package, Vascular package, OB/GYN package, General imaging package, In built 3D package (part of the OB/GYN package), DIACOM comparable for networking communication, 3D & Panoramic image display technology, Independent monitor tilt and swived adjustment. TRANSDUCER: 2 to 5 MHz broadbank curved array transducer for abdomen, OB.GYN application. 3 to 12 MHz broadband Linear transducer for vascular & Small parts application. 4 to 8 MHz broadband curved endovaginal application. Colour printer for inbuilt image management system. CD ROM recorder, USB, DVD. DIACOM media with MOD. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS. | Set | √
1.1 | Vaccum extractor for instrunmental delivery. with siliconized cups (40 mm, 50 mm, 60 mm) with electric motar and twin bottle. Complete standard set. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √
1.11 | Manual vacuum aspiration (MVA) kit, Disposable: Vaccum aspirator syringe of 50 ml volume with adaptors and suction cannulas of different sizes (2mm to 12mm size), Made of plastic | Set | √ | √ | √ |  | supplied to where there are trained doctors in BHU I level
1.12 | Delivery bed: Electrohydraulic operation. Lithotomy position with all standard accessories such a stirrup, clamp rods etc. Detacheable leg section, leg section should be split in two parts. Arm supports - one pair. Separate head section. Lateral tilt 15 - 30 degrees. Trendlenberg 30 degres. antitrendlenberg 20 degree - 30 degree. Protected against blood, liquor or waste sepage into machine. Hand clutches - one pair, to hold while bearing down during childbirth, Remote control, Detachable fluid collector ( according to bed size) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √
1.13 | Ultrasound machine (Portable) with only abdominal Probe : (Small ultrasound machine with gray scale monitor with abdominal probe only) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ |  |  | Only for impatient department
1.14 | Bed for Gyne Examination and procedures with leg stirrups for lithotomy position and motorised table with the capacity to hold upto 250kg of weight | Set | √ | √ | √ |  | supplied to where there are gynecologist at DH level
1.15 | Breast pumps (Electric) with variable pressure regulating device, Siliastic cups, on/off switch, reusable containers, silent fan and battery back-up.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Piece | √ | √ | √
1.16 | Infusion Pump: Pump type: Peristaltic finger type, Control method: Drip rate control, Required infusion set: Standard PVC IV set with 3.4 - 4.5 mm OD tubing, Flow rate range: 1 - 450 ml/h, Purge rate: more than 450 ml/h, Alarm: Door open, Occlusion, Low battery, Internal malfunction, Air in line, Infusion complete, Empty container, Reminder, abnormal drips. Mounting: Pole or Bed frame. Power supply: 220 - 240V =/- 10% 50 Hz.power cord with round 3 pin top | Set | √ | √
1.17 | Syringe Pump: Flow rage (ml/h), 0.1 - 200 ml/h 10 ml syringe, 0.1 - 300 ml/h, 5 ml to 50 ml syringe, (0.1 ml/h increments), Rapid flow rate 200 - 900 ml/h depending on size of syringe, Accuracy: Within +/-1 to 2%, Alarm: occlusion, near empty, end, system malfunction, syringe disengage, reminder, low battery, Power source: 240 V, 50Hz and power cord with round 3 pin top. Built in battery, Mounting: Pole or bed frame. | Set | √ | √ | √
1.18 | Patient warming system with drape warming blankets : Fast warm up, Temperature range Low 32', medium 38', and high 40' - 44'C, Multiple mounting options: cart, bedrail, IV pole or floor, Drape warming blankets: for upper and lower part of body (adult) & pediatric warming blanket ( set of 10 blankets each). AC power 220 - 240V, 50 Hz, Hose 5 ft. with nozzle clip. Power supply: 220-240V. Power cord with round 3 pin top | Set | √ | √ | √
1.19 | Sequential Compression Device for management of DVT prophalyais The Kendall SCD™ 700 series controller is the latest compression technology from Medtronic, delivering clinically proven sequential, gradient, circumferential compression (to the leg, foot or both simultaneously) to help prevent deep vein thrombosis (DVT) and pulmonary embolism (PE). The controller allows patients to receive VTE prevention in the comfort of their own home during the critical DVT risk period, offering the same standard of care that doctors currently prescribe and patients utilize over time in the hospital environment. With the Kendall SCD™ Express Compression System, fully transportable vascular therapy is ready to go. Move patients from room to room, floor to floor, wherever you need to take them – no plugs, no hassles, and no interruption to their protection. That’s DVT Prevention. On Time. Anywhere. Kendall SCD™ sequential compression systems deliver all-in-one leg or foot compression to help prevent DVT and PE. Patented Vascular Refill Detection technology allows the system to detect each patient's individual vascular refill time to deliver customized therapy. | set | √ | √
1.2 | Baby cot cradle with mattress | Set | √ | √ | √ |  | Only where there are cesarean facility
2 | Consumables
2.1 | CTG paper compatible with Philips Machine. | Piece | √ | √ | √ |  | where there is CTG machine and Gynae
2.2 | CTG paper compatible with Mastro Mediline Machine. | Piece | √ | √ | √ |  | where there is CTG machine and Gynae
2.3 | CTG paper compatible with CTG Huntleigh Machine. | Piece | √ | √ | √ |  | where there is CTG machine and Gynae
2.4 | Glossy printing paper compatible with ultrasound camera | Piece | √ | √
2.5 | Mattress for Obstetric & labour bed: compatible with the Obstetric & labour bed (72x28x32"). Removable 2 parts mattress with perineal cut out seat section. Polyurethane foam with reduction cover. (Measurement: Upper part with 136 cm in length, 87 cm in breadth and 18 cm thick. Lower Part with 87 cm length, 60 cm breadth and 18 cm thick.) | Piece | √ | √ | √ | √
2.6 | Mattress for Motorized Birthing Bed compatible with Janak Healthcare : Model No. Arrive 1512, Size : 2155mm L x 1010 mm W x 515mm to 880mm H. comprises of 2 pieces. (i) Body Section with concave in middle (ii) Foot Section with convex in the middle | Piece | √ | √
2.7 | Spatula for Pap smear: Sterile, Made of plastic, Shaped like allys bury spatula | Pkt of 100 | √ | √ | √ | √ | discuss with DoMSHI and RH for local production
2.8 | Cytobrush for Pap smear : | Piece | √ | √ | √ | √
2.9 | Silicone Ring pessaries of different sizes from 5 cm to 10 cm. For use in Uterine prolapse | Set | √ | √ | √ |  | where there is Gynae (No. of pieces in the set to be defined)
2.1 | Umblical cord clamp: ETO sterile pyrogen free security lock to prevent accident opening made from non-toxic medical grade polymer with solid grip for easy cord excision | Piece | √ | √ | √ | √
2.11 | Endometrial suction device/curette (Pipette): plastic and disposable | Pkt of 25 | √ | √
2.12 | Intrautrine Insemination (IUI) Cannulas / Catheter (For use in infertility treatment, to deposit enriched sperms into uterine cavity), length-23cm, external diameter 2.6mm, internal diameter 2.2mm, internal cannula diameter 1.1mm, catheter volume 0.35ml (cannula should be flexible) | Piece | √ | √
2.13 | Siliconized vacuum cup: Set of 40mm, 50mm, 60mm with tubings | Set | √ | √ | √
2.14 | Laminaria Tent (Seaweed, Hygroscopic material for Osmotic dialation of cervics) | Piece | √
2.15 | Lugol's Iodine (Potassium Iodide Solution for Colposcopy & LEEP) | Jar | √ | √ | √ |  | Where there are gynecologist
2.16 | Fischie's Clip for use in laproscopic tubal ligation | Packet | √ | √
3 | Instrunments
3.1 | Amniotome (hooked), ss | Piece | √ | √ | √
3.2 | Cervical biopsy forceps (minibite), ss | Piece | √ | √ | √ |  | where there is colposcopy
3.3 | Cervical punch biopsy forceps, ss | Piece | √ | √ | √
3.4 | Dilator, Hegars, set of 8 ss | Set | √ | √ | √
3.5 | Doyen's myoma screw, ss | Piece | √ | √ | √ |  | where there is OT
3.6 | Doyen's retractors, size: Medium, ss | Piece | √ | √ | √ |  | where there is OT
3.7 | Endocervical biopsy curette,ss | Piece | √ | √ | √
3.8 | Forceps, Green-Armitage, straight,ss | Piece | √ | √ | √ |  | where there is OT
3.9 | Forceps, obstetric, Wrigley's,ss | Pair | √ | √ | √ |  | where there is Gynecologist.
3.1 | Forceps, ovum, size: Large , ss | Piece | √ | √ | √
3.11 | Forceps, ovum , size : Small, ss | Piece | √ | √ | √
3.12 | Forceps, tenaculum - single toothed,,ss | Piece | √ | √ | √ | √
3.13 | Forceps, vulsellum - multiple toothed ,ss | Piece | √ | √ | √
3.14 | Leech Wilkinson's cannula for HSG, size: Large , ss | Piece | √ | √ |  |  | where there is Radiologist
3.15 | Leech Wilkinson's cannula for HSG , size: Medium, ss | Piece | √ | √ |  |  | where there is Radiologist
3.16 | Leech Wilkinson's cannula for HSG , size: Small, ss | Piece | √ | √ |  |  | where there is Radiologist
3.17 | Scissor, episiotomy, Blunt tip, size: 6", ss | Piece | √ | √ | √ | √
3.18 | Speculum, cusco's, size: Large ss | Piece | √ | √ | √ | √
3.19 | Speculum, cusco's, size: Medium, ss | Piece | √ | √ | √ | √
3.2 | Speculum, cusco's, size: Small, ss | Piece | √ | √ | √ | √
3.21 | Speculum, sim's, size: Large, ss | Piece | √ | √ | √ | √
3.22 | Speculum, sim's, size : Medium, ss | Piece | √ | √ | √ | √
3.23 | Speculum, sim's, size: Small, ss | Piece | √ | √ | √ | √
3.24 | Uterine curette double ended, sharp/blunt with 14.5mm head size - Large, ss | Piece | √ | √ | √
3.25 | Uterine curette double ended, sharp/blunt with 11.5mm head size - Medium, ss | Piece | √ | √ | √
3.26 | Uterine curette double ended, sharp/blunt with 6.8mm head size - Small , ss | Piece | √ | √ | √
3.27 | Uterine sound , ss | Piece | √ | √ | √ | √
3.28 | Wound exploring probes, ss | Piece | √ | √ | √ |  | where there is OT
3.29 | Heaney hysterectomy forceps, ss | Piece | √ | √ | √ |  | where there is OT
3.3 | Kogan's endocervical speculum, ss | Piece | √ | √ | √ |  | where there is Colposcopy
3.31 | Richardson's Retractor, ss | Piece | √ | √
3.32 | Wartheim's clamps, size: Large , ss | Piece | √ | √
3.33 | Wartheim's clamps, size: Medium, ss | Piece | √ | √
3.34 | Scissor, Cord cutting, size: 6 inch, ss | Piece | √ | √ | √ | √
3.35 | CuT ( Intra Utrine Device: IUD) removable, hook | Piece | √ | √ | √ |  | where there are gynecologists
3.36 | Intra Utrine Device : (Copper "T"380) | Piece | √ | √ | √ | √
3.37 | Auvard's speculum for vaginal hysterectomy, ss | Piece | √ | √
3.38 | Shirodkar's uterine clamp, ss | Piece | √ | √
4 | Leep Accessories
4.1 | LOOP Electrode for LEEP: size: W1 (15 mm), W2 (20 mm), size W3 (25 mm), Endocervical electrodes F2-8mm & Ball electrode 4 inch long W18 - 5mm | Set | √ | √ | √
4.2 | LEEP Patient cable (Hand held piece) | Piece | √ | √ | √
4.3 | LEEP Cautery (Patient) plate for Machine | Piece | √ | √ | √
4.4 | LEEP Smoke evacuation vaginal speculum : Large 835-5355 stamless | Piece | √ | √ | √
4.5 | LEEP Smoke evacuation vaginal speculum : standard 835-5355 stamless | Piece | √ | √ | √
4.6 | Ferric Subsulfate solution for hemostasis (Monsel), 500ml | Jar | √ | √ | √ |  | Required where there are LEEP machine
5 | Accessories of Laparoscope
5.1 | Uterine Mannipulator | Piece | √ | √
6 | Gyne Operating Laproscopy
6.1 | Equipments
6.1.1 | Laproscopic camera 3 chips high definition: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
6.1.2 | Laproscopic camera unit: flat screen, 18 inches | Piece | √ | √
6.1.3 | Laproscopic Medical Monitor: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
6.1.4 | Laproscopic CO2 insufflator: Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √
6.1.5 | Laproscopic Light source: compatible to the above light source | Piece | √ | √
6.1.6 | Power Morcellator with hand set | set
6.1.7 | Laproscopic Light cable: 0 degree,10 mm | Piece | √ | √
6.1.8 | Laproscope lens: 30 degree,10 mm | Piece | √ | √
6.1.9 | Laproscope lens : 0 degree 10mm | Piece | √ | √
6.2 | Consumables
6.2.1 | Cleaning brush: Size: 16 mm | Piece | √ | √
6.2.2 | Disposable 12 mm abdominal trocar | Piece | √ | √
6.2.3 | Disposable 10 mm abdominal trocar | Piece | √ | √
6.2.4 | Reducer: (from 10 mm to 5 mm ) compatible to Trocar | Piece | √ | √
6.2.5 | Reducer: (from 12 mm to 5 mm ) compatible to Trocar | Piece | √ | √
6.2.6 | Disposable 5mm abdominal trocar | Piece | √ | √
6.2.7 | Laparoscopic Bipolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
6.2.8 | Laparoscopic Monopolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
6.2.9 | Liga clip: LT 300 | Piece | √ | √
6.2.10 | Oil for the instruments maintenance | Piece | √ | √
6.2.11 | Single clip applicator: Titanium clip mediun for LT 300 (Ethicon) | Piece | √ | √
6.2.12 | Claw extractor with monopolar with Ratiched handle 33.2 cm, size 5mm | Piece | √ | √
6.2.13 | Curved dissector with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
6.2.14 | Grasper with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
6.2.15 | Laparoscopic Aspiration needle: size bigger than 20 Fr. | Piece | √ | √
6.2.16 | Curved dissector with Bipolar cautery 33.2 cm: size 5mm | Piece | √ | √
6.2.17 | Laproscopic Babcock (10 mm), size 10mm | Piece | √ | √
6.2.18 | Curved scissor with monopolar cautry | Piece | √ | √
6.2.19 | Irrigation and suction cannula | Piece | √ | √
6.2.20 | Lap Myoma screw | Piece
6.2.21 | Uterine elevator hand set | Piece
6.2.22 | Laproscopic Hook dissector : 5mm diameter shaft length 34 cm | Piece | √ | √
6.2.23 | Merry lamp forcep: 32.2 cm, 5 mm | Piece | √ | √
6.2.24 | Metal trocar: size 5mm | Piece | √ | √
6.2.25 | Metal trocar: size 10mm | Piece | √ | √
6.3 | Accessories
6.3.1 | CO2 regulator for D-type cylinder | Piece | √ | √ |  |  | for D-type cylinder
6.3.2 | CO2 regulator for B-type cylinder | Piece | √ | √ |  |  | for B-type cylinder
6.3.3 | CO2 cylinder D type | Piece | √ | √ |  |  | D-type
6.3.4 | CO2 cylinder B type | Piece | √ | √ |  |  | B-type
6.3.5 | High pressure Hose (CO2), connection length 10m | Piece | √ | √ |  |  | connection length10 m.
7 | Gyne Laproscopy for TL
7.1 | Equipments
7.1.1 | Laproscopic camera 3 chips high definition: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
7.1.2 | Laproscopic camera unit: flat screen, 18 inches | Piece | √ | √
7.1.3 | Laproscopic Medical Monitor: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
7.1.4 | Laproscopic CO2 insufflator: Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √
7.1.5 | Laproscopic Light source: compatible to the above light source | Piece | √ | √
7.1.6 | Laproscopic Light cable: 0 degree,10 mm | Piece | √ | √
7.1.7 | Laproscope lens: 30 degree,10 mm | Piece | √ | √
7.2 | Consumables
7.2.1 | Cleaning brush: Size: 16 mm | Piece | √ | √
7.2.2 | Disposable 12 mm abdominal trocar | Piece | √ | √
7.2.3 | Disposable 10 mm abdominal trocar | Piece | √ | √
7.2.4 | Reducer: (from 10 mm to 5 mm ) compatible to Trocar | Piece | √ | √
7.2.5 | Reducer: (from 12 mm to 5 mm ) compatible to Trocar | Piece | √ | √
7.2.6 | Disposable 5mm abdominal trocar | Piece | √ | √
7.2.7 | Laparoscopic Bipolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
7.2.8 | Oil for the instruments maintenance | ml | √ | √
7.2.9 | Filshie clip applicator | Piece | √ | √
7.2.10 | Grasper with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
7.2.11 | Curved dissector with Bipolar cautery 33.2 cm: size 5mm | Piece | √ | √
7.2.12 | Laproscopic Babcock (10 mm), size 10mm | Piece | √ | √
7.2.13 | Metal trocar: size 5mm | Piece | √ | √
7.2.14 | Metal trocar: size 10mm | Piece | √ | √
7.3 | Accessories
7.3.1 | CO2 regulator for B-type cylinder | Piece | √ | √ |  |  | for B-type cylinder
7.3.2 | CO2 cylinder B type | Piece | √ | √ |  |  | B-type
7.3.3 | High pressure Hose (CO2), connection length 10m | Piece | √ | √ |  |  | connection length10 m.
8 | Accessories of Hysteroscope
8.1 | Ob/Gyn Operating Hysteroscopic parts ( Richard Wolf)
8.1.1 | A: 5.5mm Optiflow Hysteroscope sheath with 5 Fr working channel includes: 1 : Optiflow sheath-5.5mm 219mm 5 Fr. channel, Part # 8974.3511 2: Operative outer sheath-5mm, 205mm WL, Part # 8974.152 3: Operative inner sheath-205mm, WL, 7 Fr channel, Part # 8974.251 4: Operative sheath, 4mm, 245mm WL, 3 Fr channel, Part # 8974.121 5: Hysteroscope-30 ° 2.7mm, Part # 8974.402 | Set | √ | √
8.1.2 | B : Hand Equipments (Richard Wolf) |  | √ | √ | √ | √
8.1.3 | Biopsy Forcep: 3 Fr.410mm WL, Part # 829.131 | Piece | √ | √
8.1.4 | Micro Scissor: 5Fr.355mm WL, Part # 8642.681 | Piece | √ | √
8.1.5 | Grasping Forcep: 5 Fr.355mm WL, Part # 828.131 | Piece | √ | √
8.1.6 | Punch Biopsy Forcep : 5 Fr.340mm WL , Part # 8642.631 | Piece | √ | √
8.1.7 | C: Bipolar Electrode: (Richard Wolf) 1: BipoTrode electrode 5Fr. 36mm WL, Part # 823.005 2: ERBE Bipolar cable, 3M, Part # 8108.135 3: ConMed, Valley lab Bipolar cable, 3M 4: ERBE VIO 300D Bipolar generator | Set | √ | √
8.1.8 | D: Fluid Manager System: (Richard Wolf) -2222.223-complete system, includes Hystero pump II (2222.011), Fluid monitoring module (2223.011) - Inflow Tubing, disposable, sterile boc of 10x2 (4170.223) | Set | √ | √
8.1.9 | E: Princess 21Fr. Bipolar resectoscope: ( Richard Wolf) - 8653.018-19Fr. Continous flow inner sheath - 8653.021-21 Fr. Continuous flow outer sheath - 8653.18- Obturator | Set | √ | √
9 | Accessories
9.1 | Probe for "Huntleigh "CTG machine: Model No. BD400XS-2, Serial No. 714X0201563-08 | Piece | √ | √ | √ |  | where there are CTG Machine
9.2 | Probe for "Mastro Mediline "CTG machine: Model No. NEWCGT12, Serial No. 1201006 | Piece | √ | √ | √ |  | where there are CTG Machine
9.3 | Transvaginal Probes compatible with Philips | Piece | √
9.4 | Transabominal Probes compatible with Philips | Piece | √
9.5 | Camera for Ultrasound Machine | Piece | √
9.6 | Camera for Colposcopy Machine compatible with Pro-Mis | Piece | √ | √
9.7 | Delivery Light : Ceiling mounted, Dome shadowless- Two light source | Piece | √ | √ |  |  | for laparoscopic tubal ligation
9.8 | Toco probe for CTG machine (Hiuntleigh) | Piece | √ | √ | √ |  | where there is CTG
9.9 | TVS probe with groove for biopsy needle compatable with existing machine. | piece | √ |  |  |  | for TVS guided pelvic biopsy

## Medical
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Internal Medicine Department (Medical Department)
1 | Cardiology
1.1 | Echo cardiography (Echo Lab) machine; with Adult and Paediatric Probes, TEE Probe with printer (Black and white). Power supply 220-240v, 50Hz. Suitable UPS. Power cord with round 3 pin top | Set | √ | √
1.2 | Tread Mill Test (TMT) machine; Power supply 220-240v, 50Hz. Suitable Stabilizer, Power cord with round 3 pin top | Set | √
2 | Consumables
2.1 | Pigtail catheter, Size: 4F | Piece | √
2.2 | Pigtail catheter, Size: 5F | Piece | √
2.3 | Pigtail catheter, Size: 6F | Piece | √
2.4 | Printing paper for Echo machine; Black and white | Roll | √ | √
2.5 | Short guide wire, J-tipped | Piece | √
2.6 | TMT printing paper | Roll | √
2.7 | Vascular Access sheath, with dilator, size: 4F | Piece | √
2.8 | Vascular Access sheath, with dilator, size: 5F | Piece | √
2.9 | Vascular Access sheath, with dilator, size: 6F | Piece | √
2.1 | Pericardiocentesis kit | Kit | √

## OT
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Operation Thearter
1 | Equipments
1.1 | Instrument Trolley: stainless steel (length -75 cm, breadth - 55 cm. Wheel size: 3 inch and railings on three sides) | Piece | √ | √ |  |  | where there is OT
1.2 | Mayo Table: Stainless steel structure, height adjustable from 85 cm to 120 cm, manually operated, stainless steel top tray detachable, 650x530x20mm. | Piece | √ | √ |  |  | where there is OT
1.3 | Mayo table (stain less steel): Dimension of the tray: 73-75 cm X47-50 cm X 6 cm Adjustable Height: 75-140 cm Weight: 10-15 Kgs Wheel: total three castors, two of them should have brake. | Piece | √ | √ |  |  | where there is ortho surgeon
1.4 | Operation table Radiolucent table top, power adjustable height, back, trendelenburg and lateral tilt. Mobile base with manually operated locks remote control facility for table movement operation. Manual removable split leg and head rest sections, pendent hand control, arm board with ball and socket joint adjustment, boay strap polyster fabric. Internal battery power and all accessories for general surgery, kidney bridge, leggings (1pair.) One extra remote control, list of essential spare parts and accessories with rates . 3 years warrenty. Power supply: 220V-230V, Power cord of round 3 pin round type plug top , operating condition:-10-+40 degree centrigrade, RH>80%, altitude>2000m | Set | √ | √ |  |  | where there is ortho surgeon
1.5 | LED Operation Theatre light Shadow less, 2 or more than 2 domes. Light handles, (In built cctv camera-optional) Can be used for hours without disruption, heat resistant. Should be fixed to ceiling.'- Back up UPS for 4 Hrs. 210-240 V, 50Hz. with extra 8 set of OT light handle.operating condition:-10-+40 degree centrigrade,RH>80%,altitude>2000m.3 pin round type plug top | Set | √ | √ |  |  | where there is ortho surgeon
1.6 | Surgical Diathermy machine with vessel sealing device: should contain 1: Monapolar- Foot switch, connecting cable, Hand piece. 2. Bipolar-Food switch, 10 numbers of bipolar connecting cables, 4 numbers of curved bipolar forceps and two numbers of straight bipolar forceps. 3. Vessel sealing handsets and accessories. | Set | √ | √ | √ |  | where there is ortho surgeon
1.7 | Operation Theatre Light handles: compatible to the existing OT light/machine | Piece | √ | √ |  |  | Where there is OT
1.8 | OT fumigation machine: Power supply: 220V-230V, Power cord of round 3 pin round type plug top | piece | √ | √ |  |  | Where there is functional OT
1.9 | Collin Rectractor (Adult): (24 Cm/10”) (38 X 60 mm) | piece | √ | √
1.1 | Collin Rectractor (Paed): (20 Cm/8”) (18 X40 mm) | piece | √ | √
1.11 | Cleaning brush, (210 X95 mm) | piece | √ | √
1.12 | Cleaning brush, (330 X 130 X50 mm) | piece | √ | √
1.13 | Cleaning brush, (400 X 190 X 8/12 mm) | piece | √ | √
2 | Consumables
2.1 | Abdominal Sponges (Mops): 12 x 12 inch, thick layer, X-ray detectable | Piece | √ | √ |  |  | where there is functional OT
2.2 | Abdominal sponges (mops): 15 x 15 inch, x-ray detectable. | Piece | √ | √ |  |  | where there is functional OT
2.3 | Disposable surgical cap (OT caps) | Piece | √ | √ |  |  | where there is functional OT
2.4 | Silicon “T” tube: size FG-6 | Pkt. | √ | √ |  |  | where there is surgeon
2.5 | Silicon “T” tube: size FG-8 | Pkt. | √ | √ |  |  | where there is surgeon
2.6 | Silicon “T” tube: size FG-10 | Piece | √ | √ |  |  | where there is surgeon
2.7 | Silicon “T” tube: size FG-12 | Piece | √ | √ |  |  | where there is surgeon
2.8 | Medical grade suction tubings wide bore silicon: 4 meter long, with ties in the end for its connection with any type of cannula (Autoclavable) | Piece | √ | √ | √
2.9 | Prolene mesh: 15 cm x 15 cm | Piece | √ | √ |  |  | Where there is functional OT
2.1 | Prolene mesh: 6 cm x 11 cm | Piece | √ | √ |  |  | Where there is functional OT
2.11 | Prolene mesh: 7 cm x 12 cm | Piece | √ | √ |  |  | Where there is functional OT
2.12 | Close wound suction unit: size 8 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ |  |  |  | Added for ENT sugery
2.13 | Close wound suction unit: Size 10, (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ |  |  |  | Added for ENT sugery
2.14 | Close wound suction unit: size 12 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ | √
2.15 | Close wound suction unit: size 14 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ |  |  | Where there is surgeon
2.16 | Close wound suction unit: size 16 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ | √
2.17 | Close wound suction unit: size 18 (Containing Bellow unit-1, connecting tube 1, Redon 2, curved needle 1) | Set | √ | √ | √
2.18 | Scalpel blades: No 10, packet of 100, SS | Pkt. | √ | √ | √ | √
2.19 | Scalpel blades: No 11, packet of 100, SS | Pkt. | √ | √ | √ | √
2.2 | Scalpel blades: No 15, packet of 100, SS | Pkt. | √ | √ | √
2.21 | Scalpel blades: No 20, packet of 100, SS | Pkt. | √ | √ | √
2.22 | Haemostatic Gel foam: (standard) 7 x 5 x 1 cm | Piece | √ | √ |  |  | Where there is functional OT
2.23 | Linear GI Stapler: (55 mm) | Piece | √ | √
2.24 | Linear GI Stapler Cartridge: (55 mm) | Piece | √ | √
2.25 | Linear GI Stapler: (75 mm) | Piece | √ | √
2.26 | Linear GI Stapler Cartridge: (75 mm) | Piece | √ | √
2.27 | Angular GI Stapler | Piece | √ | √
2.28 | Curved Intraluminal Stapler : with Adjustable Height Staples 33mm | Piece | √ | √
2.29 | Curved Intraluminal Stapler : with Adjustable Height Staples 25mm | Piece | √ | √
2.3 | Curved Intraluminal Stapler: with Adjustable Height Staples 21mm | Piece | √ | √
2.31 | Re- loadable Linear stapler: 30 mm | Piece | √
2.32 | Reload for Linear stapler: 30 mm | Piece | √
2.33 | Re- loadable Linear stapler: 60 mm | Piece | √
2.34 | Reload for Linear stapler: 60 mm | Piece | √
2.35 | Re- loadable Linear cutter: 75 mm | Piece | √
2.36 | Reload for Linear cutter: 75 mm | Piece | √
2.37 | Re- loadable Linear Cutter: 55 mm | Piece | √
2.38 | Reload for Linear cutter: 55 mm | Piece | √
2.39 | Round Stapler | Piece | √
2.4 | Stapler for skin: disposable, 35 mm | Piece | √ | √ |  |  | Where there is functional OT
2.41 | Skin stapler: 35 mm clip remover | Piece | √ | √ | √
2.42 | Suction nozzle Yangkuar plastic | Piece | √
2.43 | Oxidized Regenerated Cellulose: (Absorbable haemostat) 5 cm x 7.5 cm | Pkt. | √ | √ |  |  | Where there is surgeon
2.44 | Tube drain: size - 8, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.45 | Tube drain: size - 10, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.46 | Tube drain: size - 12, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.47 | Tube drain: size - 14, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.48 | Tube drain: size - 16, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.49 | Tube drain: size - 18, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.5 | Tube drain: size - 20, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.51 | Tube drain: size - 22, PVC (Polyvinyl Chloride) | Piece | √ | √ |  |  | do
2.52 | Theatre Incise Drape: 56 cm x 84 cm | Piece | √ | √ |  |  | do
2.53 | Theatre Incise Drape: 45 cm x 28 cm | Piece | √ | √ |  |  | do
2.54 | Theatre Incise Drape: 40 cm x 42 cm | Piece | √ | √ |  |  | do
2.55 | Theatre Incise Drape: 30 cm x 28 cm | Piece | √ | √ | √
2.56 | Disposable Suprapubic catheter set: 16 Fr. | Set | √ | √ |  |  | General
2.57 | Disposable Suprapubic catheter set: 14 Fr. | Set | √ | √ | √ |  | General
2.58 | Disposable Suprapubic catheter set: 8 Fr. | Set | √ | √ | √ |  | General
2.59 | Carbon cautry tip cleaner: with x-ray detectable | Packet | √ | √ | √
2.6 | Disposable surgical gowns, small | Piece | √ | √ |  |  | Where there is functional OT
2.61 | Disposable surgical gowns, medium | Piece | √ | √ |  |  | do
2.62 | Disposable surgical gowns, large | Piece | √ | √ |  |  | do
2.63 | Bard Parker Blade handle No. 3, SS | Piece | √ | √ | √ | √
2.64 | Bard Parker Blade handle No. 4, SS | Piece | √ | √ | √
2.65 | Bard Parker Blade handle No. 7, SS | Piece | √ | √ | √
2.66 | Gloves sterile, Non-powdered, Non latex Size 6.5 | Pair | √ | √
2.67 | Gloves sterile, Non-powdered, Non latex Size 7 | Pair | √ | √
2.68 | Goves sterile, Non-powdered, Non latex Size 7.5 | Pair | √ | √
2.69 | Gloves sterile, Non-powdered, Non latex Size 8 | Pair | √ | √
2.7 | Patient Positioning Pads (Visco elastic polymer): a. Donut head pad with centre dish-adult, size: 20cmx7.5cmx4.5cm, pediatric size: 14cmx5.7cmx3.2 cm. b. Horseshoe head pad-adult, size: 20.5x7.5x7.5 cm, pediatric size: 14x5.7x3.2 cm. c. Dome shaped positioned semi circular design with flat bottom configuration 36x15x7.6 cm. d. Contoured chest roll, small: 30.5x16.13 cm, medium: 40.6x16x13 cm, large: 50.8x16x13 cm. | set | √
2.71 | Patient Positioning Pads (Visco elastic polymer): b. Horseshoe head pad-adult, size: 20.5x7.5x7.5 cm, pediatric size: 14x5.7x3.2 cm. | set | √
2.72 | Patient Positioning Pads (Visco elastic polymer): c. Dome shaped positioned semi circular design with flat bottom configuration 36x15x7.6 cm. | set | √
2.73 | Patient Positioning Pads (Visco elastic polymer): d. Contoured chest roll, small: 30.5x16.13 cm, medium: 40.6x16x13 cm, large: 50.8x16x13 cm. | set | √
3 | Instrument
3.1 | Cooley vascular clamp: Atraumatic grip, 17 cm, SS | Piece | √ | √
3.2 | Dissecting right angle forcep baby adson: size - 16 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
3.3 | Dissecting right angle forcep baby adson: size - 18 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
3.4 | Dissecting right angle forcep baby adson: size - 20 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
3.5 | Fistula Probe: 12.5 cm, SS | Piece | √ | √ |  |  | Where there is surgeon
3.6 | Forceps, artery: curved, size - 12.5 cm, SS | Piece | √ | √ | √ | √
3.7 | Forceps, artery: curved, size - 15 cm, SS | Piece | √ | √ | √ | √
3.8 | Forceps, artery: curved, size - 20 cm, SS | Piece | √ | √ | √
3.9 | Forceps, artery: mosquito curved, size - 10 cm , SS | Piece | √ | √ | √
3.1 | Forceps, artery: mosquito curved, size - 12.5 cm , SS | Piece | √ | √ | √
3.11 | Forceps, artery: mosquito curved, size - 10 cm , SS | Piece | √ | √ | √
3.12 | Forceps, artery: mosquito curved, size - 12.5 cm , SS | Piece | √ | √ | √
3.13 | Forceps, artery: straight, size - 15 cm, SS | Piece | √ | √ | √
3.14 | Forceps, artery: straight, size - 20 cm, SS | Piece | √ | √ | √
3.15 | Forceps, artery: straight, size - 25 cm, SS | Piece | √ | √
3.16 | Forceps, dissecting: non-tooth, size - 10 cm, SS | Piece | √ | √ | √ | √
3.17 | Forceps, dissecting: non-tooth, size - 12.5 cm, SS | Piece | √ | √ | √ | √
3.18 | Forceps, dissecting: non-tooth, size - 15 cm, SS | Piece | √ | √ | √
3.19 | Forceps, dissecting: non-tooth, size - 20 cm, SS | Piece | √ | √ | √
3.2 | Forceps, dissecting: toothed, size - 10 cm SS | Piece | √ | √ | √
3.21 | Forceps, dissecting: toothed, size - 12.5 cm SS | Piece | √ | √ | √
3.22 | Forceps, dissecting: toothed, size - 15 cm SS | Piece | √ | √ | √
3.23 | Forceps, dissecting: toothed, size - 20 cm SS | Piece | √ | √ | √
3.24 | Forceps, dissecting: toothed, size - 25 cm SS | Piece | √ | √ | √
3.25 | Forceps, Kocher's: curved, size - 15 cm, SS | Piece | √ | √ | √
3.26 | Forceps, Kocher's: curved, size - 20 cm, SS | Piece | √ | √ | √
3.27 | Forceps, Kocher's: straight, size - 15 cm, SS | Piece | √ | √ | √
3.28 | Forceps, Kocher's: straight, size - 20 cm, SS | Piece | √ | √ | √
3.29 | Forceps, sinus, size - 15 cm, SS | Piece | √ | √ | √
3.3 | Forceps, sinus, size - 20 cm, SS | Piece | √ | √ | √
3.31 | Forceps, sponge holding, size - 15 cm, SS | Piece | √ | √ | √ | √
3.32 | Forceps, sponge holding, size - 20 cm, SS | Piece | √ | √ | √
3.33 | Forceps, sponge holding, size - 25cm, SS | Piece | √ | √ | √
3.34 | Forceps, tissue, Allis, size - 5 inch, SS | Piece | √ | √ | √
3.35 | Forceps, tissue, Allis, size -6 inch, SS | Piece | √ | √ | √
3.36 | Forceps, tissue, Allis, size - 7 inch, SS | Piece | √ | √ | √
3.37 | Forceps, tissue, Babcock, size - 5 inch, SS | Piece | √ | √ |  |  | where there is surgeon
3.38 | Forceps, tissue, Babcock, size - 6 inch, SS | Piece | √ | √ |  |  | do
3.39 | Forceps, tissue, Babcock, size - 7 inch, SS | Piece | √ | √ |  |  | do
3.4 | Intestinal clamp, crushing curved (payr), size - 8 inch, SS | Piece | √ | √ |  |  | where there is functional OT
3.41 | Intestinal clamp, crushing straight (payr), size - 8 inch, SS | Piece | √ | √ |  |  | do
3.42 | Intestinal clamp, non-crushing curved with non- traumatic serration, size 21.5 cm/8.5 inch, SS | Piece | √ | √ |  |  | do
3.43 | Intestinal clamp, non-crushing straight with non- traumatic serration, size: 21.5 cm/ 8.5 inch, SS | Piece | √ | √ |  |  | do
3.44 | Kidney stone forceps Randall (Sets of 4) | Set | √ | √
3.45 | Micro-adson (delicate dissecting forceps tooth), size: 12 cm, 4.75 inch, 1 x 2 teeth, SS | Piece | √ | √
3.46 | Micro-adson (non-tooth), size: 12 cm, 4 .75 inch, SS | Piece | √ | √
3.47 | Needle holder, size - 10 inch , SS | Piece | √ | √ | √
3.48 | Needle holder, size - 10 inch, Tungsten carbide | Piece | √ | √
3.49 | Needle holder, size - 4 inch, SS | Piece | √ | √ | √
3.5 | Needle holder, size - 4 inch, Tungsten carbide | Piece | √ | √
3.51 | Needle holder, size - 5 inch, SS | Piece | √ | √ | √
3.52 | Needle holder, size - 5 inch, Tungsten carbide | Piece | √ | √
3.53 | Needle holder, size - 6 inch, SS | Piece | √ | √ | √
3.54 | Needle holder, size - 6 inch, Tungsten carbide | Piece | √ | √
3.55 | Needle holder, size - 8 inch, SS | Piece | √ | √ | √ | √
3.56 | Needle holder, size - 8 inch, Tungsten carbide | Piece | √ | √
3.57 | Non traumatic bull dog clamp DeBakey, 3 sm, SS | Piece | √ | √
3.58 | Non traumatic bull dog clamp DeBakey, 8 cm, 3.125 inch, SS | Piece | √ | √
3.59 | Non traumatic bull dog clamps (curved), 6 cm, SS | Piece | √ | √
3.6 | Non traumatic bull dog clamps (curved), 6.7 cm, SS | Piece | √ | √
3.61 | Non traumatic bull dog clamps (curved), 8.5 cm, SS | Piece | √ | √
3.62 | Non traumatic bull dog clamps (straight), 6 cm, SS | Piece | √ | √
3.63 | Non traumatic bull dog clamps (straight), 6.7 cm, SS | Piece | √ | √
3.64 | Non traumatic bull dog clamps (straight), 8.5 sm, SS | Piece | √ | √
3.65 | Pelvis Retractor St. Marks, 29cm, 193 x 60 mm, SS | Piece | √ | √
3.66 | Retractor Deaver, 100 x 25mm (AxB), 30cm, 12 inch, SS | Piece | √ | √ |  |  | where there is functional OT
3.67 | Retractor Deaver, 100 x 50mm (AxB), 30cm, 12 inch, SS | Piece | √ | √ | √
3.68 | Retractor, langenback, 14cm, SS | Piece | √ | √ | √
3.69 | Retractor, langenback, 16cm, SS | Piece | √ | √ | √
3.7 | Retractor, langenback, 21cm, SS | Piece | √ | √ | √
3.71 | Retractor, Liver, SS | Piece | √ | √ | √
3.72 | Retractor, Morris , SS | Piece | √ | √ | √
3.73 | Safety pin (Instrument holder), SS, 13.5 cm for holding ring handle instruments | Piece | √ | √ | √
3.74 | Ring retractor set, Kirschner, size - 340 x 240 mm, SS | Piece | √ | √
3.75 | Scissors curved metzenbaum, 4 inch, SS | Piece | √ | √ | √
3.76 | Scissors curved metzenbaum, 5 inch carbide coated | Piece | √ | √ | √
3.77 | Scissors curved metzenbaum - 6 inch, SS | Piece | √ | √ | √ | √
3.78 | Scissors curved metzenbaum - 8 inch carbide coated | Piece | √ | √ | √
3.79 | Scissors curved metzenbaum - 10 inch carbide coated | Piece | √ | √ | √
3.8 | Scissors Mayo, Curved - 10 inch, SS | Piece | √ | √ | √
3.81 | Scissors Mayo, Curved - 6 inch, SS | Piece | √ | √ | √
3.82 | Scissors Straight metzenbaum -4 inch, SS | Piece | √ | √ | √
3.83 | Scissors Straight metzenbaum - 5 inch, SS | Piece | √ | √ | √
3.84 | Scissors Straight metzenbaum -6 inch, SS | Piece | √ | √ | √ | √
3.85 | Scissors Straight metzenbaum - 8 inch, SS | Piece | √ | √ |  |  | where there is surgeon
3.86 | Suture cutting scissors, 23 cm | Piece | √
3.87 | Dissecting scissors curved, 23 cm | Piece | √
3.88 | Self-retaining abdominal Retractor, Balfour - Adult, SS (100 x 35 mm) | Piece | √ | √ | √
3.89 | Self-retaining abdominal Retractor, Balfour - Child, SS (24 x 20 mm) | Piece | √ | √ | √
3.9 | Bookwalter retractor with instrument case (Consisting of accessories that can be used for advanced stomach, thoractomy, liver, gallbladder, spleen, Gynecology and other abdominal organs surgery) Horizontal Bar Post Coupling Medium Oval Ring 10.5 x 14.5" I.D. (26.7 x 36.8cm) Malleable Retractor Blade (2ea) 2'' x 6'' (51mm x 152mm) Malleable Retractor Blade 1.5'' x 6'' (38mm x 152mm) Harrington Retractor Blade Malleable Retractor Blade 3'' x 6'' (76mm x 152mm) Kelly Retractor Blade (2ea) 2'' x 3'' , 51mm x 76mm Kelly Retractor Blade 2'' x 4'' 51mm x 101mm Kelly Retractor Blade 2'' x 5'' 51mm x 127mm Kelly Retractor Blade 2'' x 6'' 51mm x 152mm Liver Retractor (Detail list of accessories should be clearly mentioned by the supplier during quotation) | set | √ |  |  |  | Added for onco sugery
3.91 | Self retaining laminectomy retractors, Beckman-Eaton, 32 cm long | Piece | √ | √ |  |  | do
3.92 | Self retaining retractor, GELPI , 17.5 cm long | Piece | √ | √ |  |  | do
3.93 | Self retaining retractor, GELPI, 18.5 cm long | Piece | √ | √ |  |  | do
3.94 | Skin HooKs - Large, SS | Piece | √ | √ | √
3.95 | Skin HooKs - Small, SS | Piece | √ | √ | √
3.96 | Thoracotomy Lung Rectractor - Adult, SS | Piece | √ | √
3.97 | Thoracotomy Lung Rectractor - Child, SS | Piece | √ | √
3.98 | Thoracotomy Rib Rectractor - Adult, SS | Piece | √ | √
3.99 | Thoracotomy Rib Rectractor - Child, SS | Piece | √ | √
3.1 | Towel Clamps, Backhaus with locking device, size - 11.5 cm, SS | Piece | √ | √ | √
3.101 | Towel Clamps, Backhaus, size – 11 cm | Piece | √ | √ | √
3.102 | Towel Clamps, Jones/Jones-Martin, 8.5 cm | Piece | √ | √ | √
3.103 | Varicose vein stripper with probes, SS | Set | √ | √ |  |  | where there is surgeon
3.104 | Humby's skin grafting handle | Piece | √ | √ | √
3.105 | Humby's skin grafting blade | Pkt. of 10 | √ | √ | √
3.106 | Self-retaining abdominal Retractor, Balfour - Adult, SS | Piece | √ | √ |  |  | where there is surgeon
3.107 | Self-retaining abdominal Retractor, Balfour - Child, SS | Piece | √ | √ |  |  | do
3.108 | Self-retaining abdominal Retractor, Balfour - Adult, SS | Piece | √ | √ |  |  | do
3.109 | Self-retaining abdominal Retractor, Balfour - Child, SS | Piece | √ | √ |  |  | do
3.11 | Self-retaining abdominal Retractor, Balfour - Adult, SS | Piece | √ | √ |  |  | do
3.111 | Self-retaining abdominal Retractor, Balfour - Child, SS | Piece | √ | √ |  |  | do
4 | Accessories
4.1 | Disposable Bipolar hand piece: with connecting cable compatible with Cautery machine (Martin) | Piece | √ | √
4.2 | Disposable Bipolar hand piece: with connecting cable compatible with Cautery machine (ERRBE) | Piece | √ | √
4.3 | Disposable Monopolar hand piece: with connecting cable compatible with Cautery machine (Martin) | Piece | √ | √
4.4 | Disposable Diathermy plate: with connecting cable for children (compatible with Martin machine) | Piece | √ | √ | √
4.5 | Disposable Diathermy plate: with connecting cable for children (compatible with ERRBE machine) | Piece | √ | √ | √
4.6 | Disposable Diathermy plate: with connecting cable for adult (compatible with Martin machine) | Piece | √ | √ | √
4.7 | Disposable Diathermy plate: with connecting cable for adult (compatible with ERRBE machine) | Piece | √ | √ | √
4.8 | Long tip for the monopolar cautry hand piece: (compatible with ERBEE Cautry machine) | Piece | √ | √ | √
4.9 | Long tip for the monpolar cautry hand piece: (compatible with Martin Cautry machine) | Piece | √ | √ | √
5 | New Items
5.1 | Povidone –Iodine 10% w/v (500 ml), for surgical skin preparation | 500mls | √
5.2 | Surface and equipment disinfectant spray. Aldehyde free (250 ml-500 ml with dispenser)For Electronic device, Laparoscopic screens, C-Arms, Monitors, emergency trolley’s | 250-500mls | √
5.3 | Disposable diathermy plate: For OR/ICU/Wards (H2O2 10% and silver nitrate 0.01%) | 500mls | √
5.4 | Surgical hand scrub solution: Chlorhexidine 5% V/V, Ethyl alcohol 61% W/V, with emollient and moisturizer) with dispenser | 500mls | √
5.5 | Surgical Hand disinfectant gel with dispenser: Bactericidal, tuberculocidal, Virucidal | 500mls | √
5.6 | Microfiber cloth (green): For cleaning operating equipments | piece | √
5.7 | Microfiber clot(red): For cleaning operating Table | piece | √
5.8 | Operation theatre cleaning cart: Single bag capacity, above 8 inch ball bearing wheel, sufficient storage space, interchangeable handles, lightweight, rubber tool grips and hooks, secure cabinet doors, water and chemical resistant), length 48-49 inch, width 21-22 inch, height 43-45 inch | piece | √
5.9 | Surgical gloves powdered: Sterile, size-6 | pair | √
5.1 | Surgical gloves powdered: Sterile, size-6.5 | pair | √
5.11 | Surgical gloves powdered: Sterile, size-7 | pair | √
5.12 | Surgical gloves powdered: Sterile, size 7.5 | pair | √
5.13 | Surgical gloves powdered: Sterile, size-8 | pair | √
5.14 | Disposable diathermy plate: For Force-Triad (Covidone) with connecting cable | piece | √
5.15 | Disposable diathermy plate: For Schiller diathermy Machine with connecting cables | piece | √
5.16 | Para formaldehyde tablets: For disinfection of endoscopic cables and accessories | piece | √
5.17 | Glutaraldehyde 2% with OPA: For high level disinfection | 5 ltrs | √
5.18 | Automatic Liquid Dispenser ( For Soap/disinfectant solution): Stainless steel or plastic | 250-500mls | √
Note on B61: dell: add near gelfoam

## opthalmology
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Opthalmology
1 | Equipments
1.1 | SnellenTrail Set - Non Illuminated, Power supply 220-240v, 50Hz.Power cord with round 3 pin top. | Piece | √ | √ | √
1.2 | Cabinet Roller Test Type Drum for 3 Meter Distance. Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √ | √
1.3 | Internally Illuminated Rotating Drum Distance Vision | Piece | √ | √ | √
1.4 | Near Vision Eye Testing Drum, Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √ | √
1.5 | Binomag Loupe With Adjustable Headband | Piece | √ | √ | √
1.6 | Slit Lamp - Haag Streight Slit Lamp, Power supply 220-240v, 50Hz. | Piece | √ | √
1.7 | Slit Lamp - Topcon SL D2, Voltage 100-120 / 200-240, Power 160VA, | Piece | √ | √ | √
1.8 | Lensometer | Piece | √ | √ | √
1.9 | Bipolar Cautery, Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √
1.1 | Ophthalmoscope Direct - 3.5V Coaxial 11720 | Piece | √ | √ | √
1.11 | Retinoscope - Streak - 18240 | Piece | √ | √ | √
1.12 | Indirect Ophthalmoscope - 12500 | Piece | √ | √
1.13 | Tonometer - Schiotz | Piece | √ | √ | √
1.14 | Tonopen Reichert | Piece | √ | √
1.15 | Perkins Tonometer | Piece | √ | √
1.16 | Plantech Refraction Unit with LCD screen and remote control | Set | √ | √
1.17 | AB Scan, Power supply 220-240v, 50Hz. | Set | √ | √
1.18 | Operating Microscope - portable - OPMI - 1 -FR, Power supply 220-240v, 50Hz. | Set | √ | √
1.19 | Operating Microscope ( CL-995D), The main scope adopts the technology of 5-step,3.5,5.4,8.5,13.5,21X, optical routeway and motorized zooming. The Image is clear, good stereoscopic sense and wide view field. The working distance is 200mm. Bi-oculars microscope has the angel between the main and bi-oculars is 90°. coaxial Cold light source illumination, and brilliance of the light spot:≥50000LX. Optional: Light distributor, cameral interface, single-ocular showing scope. | Set | √ | √
1.2 | Automated Humphrey Perimeter | Set | √ | √
1.21 | Digital Retinal Camera | Set | √ | √
1.22 | Clinical Specular Microscopy Non Robo - NSD 9900 | Set | √
1.23 | Eye Bank Kerato analyser Model No. EKA - 04 | Set | √
1.24 | YAG LASER | Set | √ | √
1.25 | Iridex ARGON Red LASER | Set | √ | √
1.26 | LASER Eye Safety Filter | Piece | √ | √
1.27 | Vitrectomy Machine, Power supply 220-240v, 50Hz. | set | √
1.28 | OCT Machine, Power supply 220-240v, 50Hz. | set | √
1.29 | Green Diode LASER For Photocoagulation, Power supply 220-240v, 50Hz. | set | √
1.3 | Lameller Flow Hood | Piece | √
1.31 | Pedriatric Trial Frame | Piece | √ | √ | √
1.32 | Keeler Trial Frame (Pediatric) | Piece | √ | √
1.33 | Paediatric opthalmoscope with alkaline battery | Piece | √
1.34 | Autoclave machine | Piece | √ | √ | √
1.35 | WelchAllyn wall mount Retinoscope set 3.5V | set | √ | √ | √
1.36 | LASER Saftety Glasses | Piece | √ | √
1.37 | Mainster Focal / Grid Retina Argon / Diode Laser Lens | Piece | √
1.38 | Oaiy Abraham Iridectomy Yag Leser Lens | Piece | √
1.39 | Oaiy Abraham Capsulotomy Yag Leser Lens | Piece | √
1.4 | VOLK Super Quad 160 Fundus Laser Lens | Piece | √
1.41 | VOLK Pan Retinal 2.2 Laser Lens (AE031590) | Piece | √
1.42 | Hertal Exophthalmometer | Piece | √ | √
1.43 | Vitrectomy irrigation lens, plano convex | Piece | √
1.44 | Vitrectomy irrigation lens, plano concave | Piece | √
1.45 | Ellman Cautery Radiofrency with Electroepilation Tips | Piece | √
1.46 | Eye Model for Teaching | Piece | √ | √ | √
1.47 | I Care Tonometer | Piece | √ | √
1.48 | Haag Streit Applanation Tonometry | Piece | √ | √
1.49 | Recgargeable battery 3.5Volts Catalog No.72200 made in Japan WELCH ALLYN | Piece | √ | √ | √
1.5 | Recgargeable battery 3.5Volts Catalog No.72300 made in Japan WELCH ALLYN | Piece | √ | √ | √
2 | Vitreo-retinal Sub-speciality Clinic - Equipments, consumables & accessories
2.1 | Vitrectomy - Constellation Vision System 1.Online UPS, 2.Vitrectomy cassette, 3.Tubing accessories, 3.Tubing accessories, 3.Tubing accessories | Piece | √
2.2 | Optical Coherence Tomography - Cirrus HD 1.Online UPS, 2.Coulur Printer, 3.Printing Papers | Piece | √
2.3 | Fundus Camera 1.Colour Printer, 2.Vitrectomy cassette | Piece | √
2.4 | YAG 532 nm Laser Machine YAG 532 nm Laser Machine, 2.Squad lens - 150D, 3.Coupling media, 4.Eye safety spectacles, 5.Fibre optic pipe | Piece | √
2.5 | Argon Laser Machine 1.Smart key, 2.Eye safety spectacles | Piece | √
2.6 | Phaco-machine, Universal II 1.Tubing accessories, 2.Foldable IOLs, 3.IOL injector, 3.IOL injector | Piece | √
2.7 | Vitrectomy Machine | Piece | √
2.8 | Laser Machine. For Endolaser in operation theatre | Piece | √
2.9 | Laser Machine: Retinal laser therapies in OPD | Piece | √
2.1 | Laser safety glasses: Use for safety purpose during laser procedures | Piece | √
2.11 | OCT Machine: High defination scanning of retinal structures | Piece | √
2.12 | Digital Auto Lensmeter 7"Big touch screen Ophthal Display 7" TFT LCD 758x480 pixels PD UV with printer | Piece | √ | √
2.13 | Hand Held Keratometer Macro KM 500 with infra-red tansmission to printer and with charger | Set | √ | √
2.14 | Non-Mydriatic Auto Fundus Cemera AF-330 with color printer | Set | √ | √
2.15 | Digital Vision Chart, LCD 18.5 inch, Wall mount,Remote controlled, 3 meters, Display Resolution 250+-30cd/m2 | Piece | √ | √
2.16 | IOL Master 700 / Optical Biometry with printer | set | √ | √
2.17 | Slit Lamp Digital System SL-7D with color printer Slit Lamp Digital System Viewlight SL-7 Background Illumination adapter Beam Splitter Digital Adapter Camera Canon Rebel XSi (SLR 12.2 mpx) Software for Photo Management Canon Power adapter for Canon Rebel Memory Card 8 GB Electric Table Camera parameters are setup, so you don't have to figure out the correct settings. The electric table is shipped with the setup for the Slit Lamp SL-7. | set | √ | √
3 | PHACO MACHINE
3.1 | Phaco machine (Alcon) Universal II | Piece | √
3.2 | Floridable Intra Ocular Lens | Piece | √
3.3 | Phaco accessories | Piece | √
3.4 | Nugent forcep, Scleral bucking | Piece | √
3.5 | Visulas Yag III Combi | Piece | √
3.6 | Volk Lens 90 Diopter | Piece | √ | √
3.7 | Volk Lens 20 Diopter | Piece | √ | √
3.8 | 3 Angle Mirrors | Piece | √ | √
4 | INSTRUMENTS
4.1 | Graefe Muscle Hook | Piece | √ | √
4.2 | Alfonso Eye Speculum | Piece | √ | √
4.3 | Clark Eye Speclum - Adult size | Piece | √ | √ | √
4.4 | Clark Eye Speclum - Child size | Piece | √ | √ | √
4.5 | Sauer Eye Speculum Solid Blades - Infant size | Piece | √ | √
4.6 | Cottle elevator | Piece | √ | √
4.7 | Kay's scissors | Piece | √ | √
4.8 | Kimura Spatula | Piece | √ | √ | √
4.9 | Sauer Eye Speculum Solid Blades - Child size | Piece | √ | √
4.1 | Maumenee-Park eye speculum | Piece | √ | √
4.11 | Thudicum Nasal Speculum - Infant, | Piece | √ | √
4.12 | Thudicum Nasal Speculum - Child | Piece | √ | √
4.13 | Thudicum Nasal Speculum - Adult | Piece | √ | √
4.14 | Barraquer Eye Speculum (heavy) - Adult | Piece | √ | √ | √
4.15 | Barraquer Eye Speculum (heavy) - Child | Piece | √ | √ | √
4.16 | Desmarres Lid Retractor Infant - 7.5mm Wide | Piece | √ | √ | √
4.17 | Desmarres Lid Retractor - Adult - 13mm Wide | Piece | √ | √ | √
4.18 | Dastoor Iris Retractor | Piece | √ | √
4.19 | Knapp Lacrimal Sac Retractor | Piece | √ | √
4.2 | Killner Retractor for Occuloplasty | Piece | √ | √
4.21 | Jaeger - Lid Plate | Piece | √ | √
4.22 | Stallard Ptosis Plate | Piece | √ | √
4.23 | Tooks Corneal Knife | Piece | √
4.24 | Jameson Strabismus Hook | Piece | √
4.25 | Gass Retinal Detachment Hook - 13mm Flattened with 1.5mm Oval Hole at Tip | Piece | √
4.26 | Indo Gem Retinal Detachment Hook with 2mm Bulbos Tip with Groove Suture | Piece | √
4.27 | Sinskey Lens Manipulator | Piece | √ | √
4.28 | Dastoor Spatula for Pupil & Iris Repositor | Piece | √ | √
4.29 | Lang Lacrimal Sac Dissector | Piece | √ | √
4.3 | Kirby Lens Loop & Expressor | Piece | √ | √
4.31 | Wells Enucleation Spoon | Piece | √ | √
4.32 | Mule Evisceration Spoon | Piece | √ | √
4.33 | Meyerhoefer Chalazion Curette 2.0mm Cup | Piece | √ | √ | √
4.34 | Flieringa Scleral Fixation Ring - 12mm, | Piece | √
4.35 | Flieringa Scleral Fixation Ring - 14mm, | Piece | √
4.36 | Flieringa Scleral Fixation Ring - 16mm, | Piece | √
4.37 | Flieringa Scleral Fixation Ring - 18mm | Piece | √
4.38 | Hesshurg Barron Vacuum corneal trephine - 6mm | Piece | √
4.39 | Hesshurg Barron Vacuum corneal trephine - 6.50mm | Piece | √
4.4 | Hesshurg Barron Vacuum corneal trephine - 7.00mm | Piece | √
4.41 | Hesshurg Barron Vacuum corneal trephine - 7.50mm | Piece | √
4.42 | Hesshurg Barron Vacuum corneal trephine - 8.00mm | Piece | √
4.43 | Hesshurg Barron Vacuum corneal trephine - 8.50mm | Piece | √
4.44 | Hesshurg Barron Vacuum corneal trephine - 9.00mm | Piece | √
4.45 | Hesshurg Barron Vacuum corneal Punch - 6.00mm | Piece | √
4.46 | Hesshurg Barron Vacuum corneal Punch - 6.50mm | Piece | √
4.47 | Hesshurg Barron Vacuum corneal Punch - 7.00mm | Piece | √
4.48 | Hesshurg Barron Vacuum corneal Punch - 7.50mm | Piece | √
4.49 | Hesshurg Barron Vacuum corneal Punch - 8.00mm | Piece | √
4.5 | Hesshurg Barron Vacuum corneal Punch - 8.50mm | Piece | √
4.51 | Hesshurg Barron Vacuum corneal Punch - 9.00mm | Piece | √
4.52 | Hesshurg Barron Vacuum corneal Punch - 9.5 mm | Piece | √
4.53 | Corneal Transplant Trephine | Piece | √
4.54 | Flexible Ruler with Inches & mm Marking, Stainless Steel | Piece | √
4.55 | Castroveijo Caliper Straight | Piece | √ | √
4.56 | Bowman Lacrimal Probe Cylinder Stainless Steel, 00 & 0 | Piece | √ | √ | √
4.57 | Bowman Lacrimal Probe Cylinder Stainless Steel, 1 & 2, | Piece | √ | √ | √
4.58 | Bowman Lacrimal Probe Cylinder Stainless Steel, 3 & 4 | Piece | √ | √ | √
4.59 | Catalano Lacrimal Intubation Set, Stainless Steel Malleable Probes With Silicone Tubing | Piece | √ | √
4.6 | Nettle Ship Punctum Dilator | Piece | √ | √ | √
4.61 | West Bone Chisel | Piece | √ | √
4.62 | West Bone Gauge | Piece | √ | √
4.63 | Freer Periosteal Twin Elevator Double Ended | Piece | √ | √
4.64 | Kerrison Bone Nibbling Ronguer | Piece | √ | √
4.65 | Graefee Fixation Forceps-0.12mm | Piece | √ | √
4.66 | Elschnig Fixation Forceps | Piece | √ | √
4.67 | Dastoor Superior Rectus Forceps | Piece | √ | √
4.68 | Jaffe Tying Forceps - Straight, | Piece | √ | √
4.69 | Mcpherson Tying Forceps - Curved | Piece | √ | √
4.7 | Mcpherson Corneal Forceps | Piece | √ | √
4.71 | Utrta Capsulorexhis forceps | Piece | √ | √
4.72 | Lims Corneal Forceps | Piece | √ | √
4.73 | Moorfield Conjunctival Suture Forceps | Piece | √ | √
4.74 | Kuglen Hook [Iris Hook & Lens Manipulator - Push & Pull Model] | Piece | √ | √
4.75 | Barraquer Cilia Forceps | Piece | √ | √ | √
4.76 | Berke Ptosis Clamp | Piece | √ | √
4.77 | Snellen Entropion Clamp - Right | Piece | √ | √
4.78 | Snellen Entropion Clamp - Left | Piece | √ | √
4.79 | Lambert Chalazion Forcep -10mm, | Piece | √ | √ | √
4.8 | Lambert Chalazion Forcep -15mm | Piece | √ | √ | √
4.81 | Watzke Sleeve Spreading Forceps | Piece | √ | √
4.82 | Scleral Depressor | Piece | √ | √
4.83 | Metallic Retractor | Piece | √ | √
4.84 | Hartman Haemostatic Mosquito Forceps - Straight | Piece | √ | √
4.85 | Hartman Haemostatic Mosquito Forceps - Curved | Piece | √ | √
4.86 | Schaedel - Towel Clips | Piece | √ | √
4.87 | Diefenbach Serrefine Clamps | Piece | √ | √
4.88 | Allis Tissue Forceps | Piece | √ | √ | √
4.89 | Bonn Iris Scissor | Piece | √ | √
4.9 | Knapp Strabismus Scissor | Piece | √ | √
4.91 | Vannas Scissor Straight | Piece | √ | √
4.92 | Gills Welsh Vannas 11mm Angled Blade | Piece | √ | √
4.93 | Westcott Tenotomy Scissor Blunt Tips | Piece | √ | √
4.94 | Kalt Needle Holder Straight | Piece | √ | √
4.95 | Castroviejo Needle Holder Without Lock | Piece | √ | √
4.96 | Mcpherson Needle Holder Micro Straight 11mm Jaws | Piece | √ | √
4.97 | Barrquer Blade Braker & Holder | Piece | √ | √
4.98 | Lacrimal Cannula 25 Guage 1/2 Curved, SS | Piece | √ | √
4.99 | Simcoe Irrigating & Aspirating Cannula, SS | Piece | √ | √
4.1 | Drews Irrigating Vectis | Piece | √ | √
4.11 | Stainless Steel Sterilizing Case - 210x125x20mm | Piece | √ | √
4.12 | Stainless Steel Sterilizing Case - 330x220x20mm | Piece | √ | √
4.13 | Titanium Forceps Straight Tying | Piece | √ | √
4.14 | Titanium Forceps Straight Tooth | Piece | √ | √
4.15 | Titanium Mcpherson Forceps Angled | Piece | √ | √
4.16 | Titanium Needle Holder Barquer Straight | Piece | √ | √
4.17 | Castroviejo CorneoScleral Punch 1.5mm | Piece | √ | √
4.18 | Skin hooks - double prong | Piece | √ | √
4.19 | Paufique forceps | Piece | √ | √
4.2 | Stevens scissors - Curved | Piece | √ | √
4.21 | Jewellers forceps | Piece | √ | √
4.22 | Castro viejo needle holder with Lock | Piece | √ | √
4.23 | Westcott scissors - Sharp Tip | Piece | √ | √
4.24 | Westcott scissors - Blunt Tips | Piece | √ | √
4.25 | Malleable retractors sizes - 6X200mm, | Piece | √ | √
4.26 | Malleable retractors sizes - 8X200mm, | Piece | √ | √
4.27 | Malleable retractors sizes - 10X200mm, | Piece | √ | √
4.28 | Malleable retractors sizes - 12X200mm | Piece | √ | √
4.29 | Sewell's retractors | Piece | √ | √
4.3 | Sayer retractors | Piece | √ | √
4.31 | Stainless Stell Corneal Schields | Piece | √ | √
4.32 | Ball Implant Sizer | Piece | √ | √
4.33 | Ball Implant Introducer | Piece | √ | √
4.34 | Colorado tips for the cautery machine | Piece | √ | √
4.35 | Skin hooks- single and double prong | Piece | √ | √
4.36 | Cage & Jar (Moist Chamber For Enucleated Eye Storage) | Piece | √ | √
4.37 | Corneal Burrs Battery Operated | Piece | √ | √
4.38 | Groove Director For Lacrimal Intubation Tubes/Mickey Mouse | Piece | √ | √
4.39 | Implant Introducer - 20mm | Piece | √ | √
4.4 | Implant Introducer - 18mm | Piece | √ | √
4.41 | Implant Introducer - 16mm | Piece | √ | √
4.42 | Prism Bar- Horizontal + Vertical | Piece | √ | √
4.43 | Loose Prisms | Piece | √ | √
4.44 | Jackson Cross Cylinder 0.50D & 1.00D | Piece | √ | √ | √
4.45 | Lens Bar | Piece | √ | √
4.46 | Stand Magnifiers and Reading Magnifiers | Piece | √
4.47 | Illuminated Stand Magnifiers - 2.8X, | Piece | √
4.48 | Illuminated Stand Magnifiers - 3.0X, | Piece | √
4.49 | Illuminated Stand Magnifiers - 4.0X, | Piece | √
4.5 | Illuminated Stand Magnifiers - 5.0, | Piece | √
4.51 | Illuminated Stand Magnifiers - 6.0X, | Piece | √
4.52 | Illuminated Stand Magnifiers - 10.0X | Piece | √
4.53 | Illuminated Stand Magnifiers - 12.5X | Piece | √
4.54 | Hgher Additions (Prismatic Bino Comfort) - 1X, | Piece | √
4.55 | Hgher Additions (Prismatic Bino Comfort) - 1.25X, | Piece | √
4.56 | Hgher Additions (Prismatic Bino Comfort) - 1.5X, | Piece | √
4.57 | Hgher Additions (Prismatic Bino Comfort) - 2X, | Piece | √
4.58 | Hgher Additions (Prismatic Bino Comfort) - 2.5X | Piece | √
4.59 | Noves Bino - 1X, | Piece | √
4.6 | Noves Bino - 1.25X, | Piece | √
4.61 | Noves Bino - 1.5X, | Piece | √
4.62 | Noves Bino - 2.0X, | Piece | √
4.63 | Noves Bino - 2.5X | Piece | √
4.64 | Hand Held Magnifyer And Reading Magnifiers | Piece | √
4.65 | Illuminated Pocket magnifiers - 3X, | Piece | √
4.66 | Illuminated Pocket magnifiers - 3.5X, | Piece | √
4.67 | Illuminated Pocket magnifiers - 4X, | Piece | √
4.68 | Illuminated Pocket magnifiers - 5X, | Piece | √
4.69 | Illuminated Pocket magnifiers - 7X | Piece | √
4.7 | Illuminated Pocket magnifiers - 10X, | Piece | √
4.71 | Illuminated Pocket magnifiers - 12.5X | Piece | √
4.72 | Aplanatic Reading Magnifiers - 2.4X, | Piece | √
4.73 | Aplanatic Reading Magnifiers - 2.7X, | Piece | √
4.74 | Aplanatic Reading Magnifiers - 2.9X, | Piece | √
4.75 | Aplanatic Reading Magnifiers - 3.8X, | Piece | √
4.76 | Aplanatic Reading Magnifiers - 4.5X | Piece | √
4.77 | Aspheric Reading Magnifiers - 2.8X, | Piece | √
4.78 | Aspheric Reading Magnifiers - 3.0X, | Piece | √
4.79 | Aspheric Reading Magnifiers - 3.5X, | Piece | √
4.8 | Aspheric Reading Magnifiers - 4.0X, | Piece | √
4.81 | Aspheric Reading Magnifiers - 5.0X, | Piece | √
4.82 | Aspheric Reading Magnifiers - 6.0X | Piece | √
4.83 | End-Gripping Forceps - 1286 - B, 20G | Piece | √
4.84 | Serrated Jaw Forceps - 1286 - C, 20G | Piece | √
5 | Microscissors
5.1 | Vertical Scissors - 1286 - E, 20G | Piece | √
5.2 | Straight Scissors - 1286 - J, 20G | Piece | √
6 | ILM Forceps
6.1 | Eckardt End-Gripping Forceps - 1286 - I | Piece | √
7 | FB Forceps |  | √
7.1 | Avci FB Forceps - 2286 - I, 20G/0 | Piece | √
7.2 | Avci FB Forceps - 2286 - H, 17G/1 | Piece | √
8 | Vitrectomy Lenses
8.1 | Flat Lens - 1284 - DW | Piece | √
8.2 | Biconcave Lens (For FAE) - 1284 - EW | Piece | √
8.3 | Wide Field Lens (For Central PO & Viterous) Sterior - 1284 - GW | Piece | √
8.4 | 20 Degree Prism Lens (For Posterior Periphery) - 1284 - HW | Piece | √
8.5 | 30 Degree Prism Lens (For The Posterior Periphery Beyond The Equator) - 1284 - IW | Piece | √
8.6 | Suturable vitrectomy lens ring (adult size) - 1284 - R | Piece | √
8.7 | Suturable vitrectomy lens ring (Pedriatric size) - 1284 - R | Piece | √
8.8 | Magnifying Lens - 1284 F | Piece | √
8.9 | Prismatic Biconcave Lens - 1284 - L | Piece | √
9 | Reusable Vitrectomy Accessories
9.1 | Scleral Plug Forceps - 1277 | 3.5/4.0mm | √
10 | Reusable Microforceps
10.1 | End-Gripping Forceps - 1286 - B06 | Piece | √
10.2 | Eckardt End-Gripping Forceps - 1286 - W06 | Piece | √
10.3 | Serrated Jaw Forceps - 1286 - C06 | Piece | √
10.4 | Vertical Scissors - 1286 - E06 | Piece | √
10.5 | Straight Scissors - 1286 - J06 | Piece | √
11 | 23G Backflush Instruments
11.1 | Backflush With Blunt Needle With Passive Aspiration - 1281 - AD06 | Piece | √
11.2 | Backflush With Blunt Needle With Active Aspiration - 1281 - BTD06 | Piece | √
12 | ILM Procedures
12.1 | Diamond Dusted Sweeper - 1290 - DSS06, 23G/0 | Piece | √
12.2 | Soft Tipped Canula With Leur Lock Connector - 1272 - DS23, 23G/0 | Piece | √
12.3 | Spaide Soft Tipped Diamond Dusted Canula For Trimcinolone Removal - 1272 - SDT23, 23G/0 | Piece | √
12.4 | Retinal Magnetic Forceps, 23G/0 | Piece | √
12.5 | Kerrison Rongeur 1.50mm | Set | √ | √
12.6 | Kerrison Rongeur 2.00mm | Piece | √ | √
12.7 | Kerrison Rongeur 2.50mm | Piece | √ | √
12.8 | Kerrison Rongeur 3.00mm | Piece | √ | √
13 | Consumables
13.1 | Sterile Microsurgical Blade (Cresent) 2mm Wide, Tip Angled | Piece | √ | √
13.2 | Sterile Microsurgical Blade (Keratome) 3.2mm, Tip Angled | Piece | √ | √
13.3 | Sterile Microsurgical Blade (Keratome) 2.8m,m Tip Angled | Piece | √ | √
13.4 | Sterile 15 Degree Knife Side Port | Piece | √ | √
13.5 | Injection Methyl Cellulose Intra Ocular (Visco Elastic) In Prefilled Syringes 2ml | Piece | √ | √
13.6 | Injection Hyalorinadase 1500iu (vials) | Piece | √ | √
13.7 | Optisol GS (Corneal Storage Media) | Piece | √
13.8 | MK Media [McCarey-Kaufman] | Piece | √
13.9 | Aurosling - L527 | Piece | √
13.1 | Aurolac 0275D Frontalis Sling | Piece | √
13.11 | Trypan Blue 0.06% | Piece | √ | √
13.12 | Carbachol Intra Ocular Solution 0.01% | Piece | √ | √
13.13 | Bandage Contact Lens | Piece | √ | √
13.14 | Punctum Plugs | Piece | √ | √
13.15 | Cyanocrilate Glue | Piece | √ | √
13.16 | Eye shields - Small, Medium, Large | Piece | √ | √
13.17 | Eye Drape - SMS Polypropylene fabric water repellant Sheet size: 100x70 cm‚ Incise area: 7x9 cm‚ Drainage pouch of transparent PE film | Piece | √ | √
13.18 | Incise Drape Small D1020 | Piece | √
13.19 | Silicone Band 240 (2.5X0.6X125.0mm) | Piece | √
13.2 | Silicone Tyre 276 | Piece | √
13.21 | Silicone Sponge (3.0X5.0X100mm) | Piece | √
13.22 | Octane (5ml) Perfluoro-n-octane | Piece | √
13.23 | Polydimethyl Siloxane (10ml) | mL | √
13.24 | Sulphur Hexafluride (SF6) | Piece | √
13.25 | Perfluoropropane (C3F8) | Piece | √
13.26 | 30GX1/2" (0.3X13mm) Needle | Piece | √
13.27 | Ringer Lactate (Bottle With Rubber Cork) | Piece | √
13.28 | Injection Bevacizumab | Piece | √
13.29 | Injection Fluorescein for FFA | Piece | √
13.3 | Injection Triamcinolone Acetonide. | Piece | √ | √
13.31 | Schimer's Test Strip | Piece | √ | √ | √
13.32 | Thermal Paper URK-700 Auto | Piece | √ | √
13.33 | Thermal Paper Ref No. 112 60 FDT, Zeiss Humphrey Systems | Piece | √ | √
13.34 | Thermal paper (AB Scan) 110mmX18m | Piece | √ | √
13.35 | Plastic Universal Eye Schield | Piece | √ | √
13.36 | Snellen Distance Vision E Chart | Piece | √ | √ | √
13.37 | Snellen Distance Vision A Chart | Piece | √ | √ | √
13.38 | Visual Acuity ETDRS Chart | Piece | √ | √
13.39 | Visual Acuity LogMAR Chart | Piece | √ | √
13.4 | Near Vision Chart | Piece | √ | √ | √
13.41 | Isihara Color Vision Chart | Piece | √ | √ | √
13.42 | Universal Conformers - 20mm, | Piece | √
13.43 | Universal Conformers - 22mm, | Piece | √
13.44 | Universal Conformers - 24mm | Piece | √
13.45 | Artificial Eyes Small sizes (Right Eye) | Piece | √ | √ | √
13.46 | Artificial Eyes Small sizes (Left Eye) | Piece | √ | √ | √
13.47 | Artificial Eyes Medium Sizes (Right Eye) | Piece | √ | √ | √
13.48 | Artificial Eyes Medium Sizes (left Eye) | Piece | √ | √ | √
13.49 | Artificial Eyes Big Sizes (Right Eye) | Piece | √ | √ | √
13.5 | Artificial Eyes Big Sizes (Left Eye) | Piece | √ | √ | √
13.51 | Eye Spehere Solid Type - 13mm, | Piece | √
13.52 | Eye Spehere Solid Type - 15mm, | Piece | √
13.53 | Eye Spehere Solid Type - 17mm, | Piece | √
13.54 | Castroveijo Implant Solid Plastic - 13mm, | Piece | √
13.55 | Castroveijo Implant Solid Plastic - 15mm, | Piece | √
13.56 | Castroveijo Implant Solid Plastic - 17mm | Piece | √
13.57 | Medpor imlants- sheet and ball implant | Piece | √
13.58 | Eye Spere Solid Plaque - 19mm, | Piece | √
13.59 | Eye Spere Solid Plaque - 17mm, | Piece | √
13.6 | Eye Spere Solid Plaque - 15mm, | Piece | √
13.61 | Eye Spere Solid Plaque - 13mm | Piece | √
13.62 | Suction Cannula | Piece | √ | √
13.63 | Cambridge Crowding Cards | Piece | √ | √
13.64 | Cambridge Single cards | Piece | √ | √
13.65 | TNO Test For Stereoscopic Vision | Piece | √ | √
13.66 | RAF Ruler | Piece | √ | √
13.67 | Portable Low Vision Chart | Piece | √ | √
13.68 | Armsler grid | Piece | √ | √
13.69 | Cardiff Picture Chart (Pedriatric) | Piece | √ | √
13.7 | Portable Love & Bailey Low Vision Chart | Piece | √ | √
13.71 | Infusion Canula 5.mm | Piece | √
13.72 | Infusion Canula 7.mm | Piece | √
13.73 | Infusion Canula | Piece | √
13.74 | 23G Endolaser Pipes, 23G/0 | Piece | √
13.75 | Trochar Set Disposable | Piece | √
13.76 | Bulbs - Retinoscope Bulb 08200u | Piece | √ | √ | √
13.77 | Bulbs - Ophthalmoscope 11730 - 04900u | Piece | √ | √ | √
13.78 | Bulbs - Indirect Ophthalmoscope Ophthalmoscope - 1200u | Piece | √ | √
13.79 | Bulbs - Pocket Ophthalmoscope - 3900u | Piece | √ | √ | √
13.8 | Bulbs - Penlite - 07600u | Piece | √ | √
13.81 | Bulbs for Topcon Slitlamp - SL-3F | Piece | √ | √
13.82 | Bulbs for Topcon Slitlamp - SL-D2 | Piece | √ | √
13.83 | Bulbs for Slitlamp Haag Streit - BQ-900 | Piece | √ | √
13.84 | Bulbs for Slitlamp Haag Streit - BM-901 | Piece | √ | √
13.85 | Air Cannula (RYCROFT)27Gx22mm,angle 45' ,4mm bend | Piece | √
13.86 | Constellation Vitrotectomy Cassette 25 Gauze | Piece | √
13.87 | Constellation Vitrotectomy Cassette 23 Gauze | Piece | √
13.88 | Constellation Vitrotectomy Cassette 20 Gauze | Piece | √
13.89 | Perfluorocarbon Liquids (PFCL) (heavy liquid) | Piece | √
13.9 | I Care probe for Tonometery | Piece | √ | √
14 | PCIOL
14.1 | PC IOL-Single piece 13mm Length - +13.00D | Piece | √ | √
14.2 | PC IOL-Single piece 13mm Length - +14.00D | Piece | √ | √
14.3 | PC IOL-Single piece 13mm Length - +15.00D | Piece | √ | √
14.4 | PC IOL-Single piece 13mm Length - +16.00D | Piece | √ | √
14.5 | PC IOL-Single piece 13mm Length - +16.50D | Piece | √ | √
14.6 | PC IOL-Single piece 13mm Length - +17.00D | Piece | √ | √
14.7 | PC IOL-Single piece 13mm Length - +17.50D | Piece | √ | √
14.8 | PC IOL-Single piece 13mm Length - +18.00D | Piece | √ | √
14.9 | PC IOL-Single piece 13mm Length - +18.50D | Piece | √ | √
14.1 | PC IOL-Single piece 13mm Length-+19.00D | Piece | √ | √
14.11 | PC IOL-Single piece 13mm Length-+19.50D | Piece | √ | √
14.12 | PC IOL-Single piece 13mm Length-+20.00D | Piece | √ | √
14.13 | PC IOL-Single piece 13mm Length-+20.50D | Piece | √ | √
14.14 | PC IOL-Single piece 13mm Length-+21.00D | Piece | √ | √
14.15 | PC IOL-Single piece 13mm Length-+21.50D | Piece | √ | √
14.16 | PC IOL-Single piece 13mm Length-+22.00D | Piece | √ | √
14.17 | PC IOL-Single piece 13mm Length-+22.50D | Piece | √ | √
14.18 | PC IOL-Single piece 13mm Length-+23.00D | Piece | √ | √
14.19 | PC IOL-Single piece 13mm Length-+23.50D | Piece | √ | √
14.2 | PC IOL-Single piece 13mm Length-+24.00D | Piece | √ | √
14.21 | PC IOL-Single piece 13mm Length-+24.50D | Piece | √ | √
14.22 | PC IOL-Single piece 13mm Length-+25.00D | Piece | √ | √
14.23 | PC IOL-Single piece 13mm Length-+25.50D | Piece | √ | √
14.24 | PC IOL-Single piece 13mm Length-+26.00D | Piece | √ | √
14.25 | PC IOL-Single piece 13mm Length-+27.00D | Piece | √ | √
14.26 | PC IOL-Single piece 13mm Length-+28.00D | Piece | √ | √
14.27 | PC IOL-Single piece 13mm Length-+29.00D | Piece | √ | √
14.28 | PC IOL-Single piece 13mm Length-+30.00D | Piece | √ | √
14.29 | PC IOL-Single piece 13mm Length-+31.00D | Piece | √ | √
14.3 | PC IOL-Single piece 13mm Length-+32.00D | Piece | √ | √
14.31 | PC IOL-Single piece 13mm Length-+33.00D | Piece | √ | √
14.32 | PC IOL-Single piece 13mm Length-+34.00D | Piece | √ | √
14.33 | PC IOL-Single piece 13mm Length-+35.00D | Piece | √ | √
15 | AC IOL
15.1 | AC IOL-Single piece 13mm Length - +13.00D | Piece | √ | √
15.2 | AC IOL-Single piece 13mm Length - +14.00D | Piece | √ | √
15.3 | AC IOL-Single piece 13mm Length - +15.00D | Piece | √ | √
15.4 | AC IOL-Single piece 13mm Length - +16.00D | Piece | √ | √
15.5 | AC IOL-Single piece 13mm Length - +17.00D | Piece | √ | √
15.6 | AC IOL-Single piece 13mm Length - +18.00D | Piece | √ | √
15.7 | AC IOL-Single piece 13mm Length - +19.00D | Piece | √ | √
15.8 | AC IOL-Single piece 13mm Length - +20.00D | Piece | √ | √
16 | PC Foldable Len
16.1 | Foldable IOL dioptre size 19.00D | Piece | √
16.2 | Foldable IOL dioptre size 19.50D | Piece | √
16.3 | Foldable IOL dioptre size 20.00D | Piece | √
16.4 | Foldable IOL dioptre size 20.50D | Piece | √
16.5 | Foldable IOL dioptre size 21.00D | Piece | √
16.6 | Foldable IOL dioptre size 21.50D | Piece | √
16.7 | Foldable IOL dioptre size 22.00D | Piece | √
16.8 | Foldable IOL dioptre size 22.50D | Piece | √
16.9 | Foldable IOL dioptre size 23.00D | Piece | √
17 | Foldable IOL dioptre size 23.50D | Piece | √
17.1 | Foldable IOL dioptre size 24.00D | Piece | √

## orthopedic
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Orthopedic Department
1 | Basic Orthopaedic Supplies:
1.1. | Plaster of Paris
| Paster of Paris: Size 7.5cmx 2.7m | Roll | √ | √ | √
| Paster of Paris: Size 10cm x 2.7m | Roll | √ | √ | √
| Paster of Paris: Size 15cm x 2.7m | Roll | √ | √ | √
| Fiber glass cast: Size- 10cm x 2.7m | Roll | √ | √
| Fiber glass cast: Size- 7.5cm x 2.7m | Roll | √ | √
1.2 | Cotton cast padding (soft rolls) :
| Cotton cast padding (soft rolls) : Size 10cm x 3m | Roll | √ | √ | √
| Cotton cast padding (soft rolls) : Size 15cm x3m | Roll | √ | √ | √
1.3 | Stokinette Rolls:
| Stokinette Rolls: Size 5cmx 10m | Roll | √ | √
| Stokinette Rolls: Size 7.5cm x 10m | Roll | √ | √
| Stokinette Rolls: Size 10cm x 10m | Roll | √ | √
1.4 | Dressing Materials for Advanced Wound Mangt.
| i. Knited elastic fixation bandage (Tensofix): Size: 10 x 2.4m unstretched and 4m stretched | Roll | √ | √
| ii. Adhesive incise drape, eg. OPSITE Incise, 30 x 28cm | 10/Pack | √ | √
| iii. Adhesive incise drape, eg. OPSITE Incise, 45 x 55cm | 10/Pack | √ | √
| iv. Adhesive film dressing with absorbent pad, eg. OPSITE Post-Op, 15.5 x 8.5cm | 20/Box | √ | √
| v. Adhesive film dressing with absorbent pad, eg. OPSITE Post-Op, 25 x 10cm | 20/Box | √ | √
| vi. Absorbent dressing with adhesive border, eg. PRIMAPORE, 15 x 8cm | 20/Box | √ | √
| vii. Absorbent dressing with adhesive border, eg. PRIMAPORE, 25x10cm | 20/Box | √ | √
| viii. Hydrocellularantimicrbial foam dressings, eg. Allevyn Ag Gentle Border, 12.5 x 12.5cm | 10/Pack | √ | √
| ix. Hydrocellularantimicrbial foam dressings, eg. Allevyn Ag Gentle Border, 17.5 x 17.5cm | 10/Pack | √ | √
| x. Hydrofiber Dressing, eg. Aquacel Ag Surgical, 9 x 15cm | 10/Pack
| xi. Hydrofiber Dressing, eg. Aquacel Ag Surgical, 9 x 25cm | 10/Pack
| xii. Hydrofiber Dressing, eg. Aquacel Foam Adhesive, 17.5 x 17.5cm | 10/Pack
| xiii. Hydrofiber Dressing, eg. Aquacel Foam Sacral, 20 x 16.9cm | 5/Pack
1.5 | Negative Pressure Wound Therapy (NPWT) System
| Closed Cyclic NPWT PU Wound Management device | No | √ | √
1.6 | Disposable supplies:
| CCNPWT canister (500ml) with Bacteria filter, Y Connector and Drainage tube- | No | √ | √
| CCNPWT PU Wound Drainage Kit- Small/7.5 x 10 x 3cm | No | √ | √
| CCNPWT PU Wound Drainage Kit- Medium/12.5 x 18 x 3cm | No | √ | √
| CCNPWT PU Wound Drainage Kit- Large/15 x 26 x 3cm | No | √ | √
| Bacteriostatic & Fungistatic Dressings- small/medium/large | No | √ | √
Abdominal Wound Kit
1.7 | Splints, Braces & Accessories
| Universal Buck's skin traction splint | No | √ | √ | √
| Leg Traction Brace with velcro closure: Size: small | No | √ | √ | √ | √
| Leg Traction Brace with velcro closure: Size: medium | No | √ | √ | √ | √
| Leg Traction Brace with velcro closure: Size: large | No | √ | √ | √ | √
| Thomas splint adult size | No |  |  | √ | √
| Thomas splint child size | No |  |  | √ | √
| Traction Splints: Hare / Sagar traction splint: universal size | No | √ | √
| Crammer Wire Splints: Size - 24''x2" | No |  |  | √ | √
| Crammer Wire Splints: Size - 36"x3" | No |  |  | √ | √
| Crammer Wire Splints: Size - 36"x4" | No |  |  | √ | √
| Cervical orthosis (Philadelphia): small | No | √ | √
| Cervical orthosis (Philadelphia): medium | No | √ | √
| Cervical orthosis (Philadelphia): large | No | √ | √
| Spine Board and Strapping Systems-eg BAXSTRAP TM | No | √ | √
| Jewitt's Brace or Ash Brace: medium | No | √ | √
| Jewitt's Brace or Ash Brace: Large | No | √ | √
| Taylor's Brace: small | No | √ | √
| Taylor's Brace: medium | No | √ | √
| Taylor's Brace: Large | No | √ | √
| Knee Brace long- medium | No | √ | √
| Knee Brace long- large | No | √ | √
1.8 | Steinmann Pins:
| Steinmann Pins, with trocar tip, 3.5mm dia., length 150mm | No | √ | √
| Steinmann Pins, with trocar tip, 3.5mm dia., length 200mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size - 4.0mm x 175mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size - 4.0mm x 200mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size - 4.0mm x 225mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size -4.5mmx250mm | No | √ | √
| Steinmann Pins, with trocar tip and middle thread: Size- 5.0mm x 250mm | No | √ | √
1.9 | Bone Hand Drill : Bunnell- Hand drill with three-jaw chuck & key for shanks up to 4 mm diameter,cannulated
1.1 | Spica Cast Table (Paediatric)
2 | Basic Equipment for Orthopaedic Surgery
| i. Oscillating Plaster Saw, 220V/50Hz with the following parts: a. Two Spanners to exchange the saw blades; b. Spare saw blade (dia. 50mm); c. Spare saw blade (dia. 65mm); c. Spare saw blade (dia. 65mm); d. Spare saw blade, special hardness, (dia. 50mm); e. Spare saw blade, special hardness, (dia. 65mm);f. Plaster Spreader, 22. 5cm, 9" | Set | √ | √ | √ |  | For district hospitals wherever there are orthopaedic Technicians
| ii. Plaster Spreader, 22. 5cm, 9" | No | √ | √ | √ |  | Same as above
| iii. Plaster-of- Paris Shears (Plaster cutting scissors): Bruns, 24 cm, 9 1/2" | No | √ | √ | √ |  | Same as above
2.1 | C Arm Machine : C-arm high frequency mobile X-ray Machine Suitable for the operating room and emergency room for real time guidance of orthopaedic operations. | No | √ | √ | √
2.2 | Electric / battery Powered Bone Drills | [5 empty] | Will be needed in Gelephu in 2017
| a. Hand Pieces: 1. Duel Trigger Rotatory Hand Piece 2. Sagittal Saw Hand Piece | No | √ | √
| b. Drilling and Reaming Attachments consisting of: 1. Small Synthes Quick Connect 2. Stryker 1/4' Adjustable Jacob's Chuck ( keyed) 3. Synthes' Reaming Attachment (Large) 4. Synthes' DHS Reaming Attachment 5. Wire Collet (0.7mm -2.8mm) 6. Pin Collect (2.0mm-4.0mm) | No | √ | √
| c. Accessories consisting of: 1. Battery Charger 2. Batteries (3 pairs) 3. Aseptic Battery Kit 4. Sterilization Case for 3 hand Pieces 5.Saw Blades all sizes (3 pairs) | No | √ | √
2.3 | Pneumatic Tourniquet System:
a | i. Zimmer or Stryker equivalent doul channel tourniquet | Set | √ | √
| ii. Tourniquet Cuff :Paediatric size | No | √ | √
| iii. Tourniquet Cuff : size 12 inches | No | √ | √
| iv. Tourniquet Cuff : size 18 inches | No | √ | √
| v. Tourniquet Cuff : size 24 inches | No | √ | √
| vi. Tourniquet Cuff : size 30 inches | No | √ | √
| vii. Tourniquet Cuff : size 34 inches | No | √ | √
b | Eshmarch :
| Eshmarch : 5m long and 7.5cm wide | No | √ | √
| Eshmarch : 5m long 10cm wide | No | √ | √
| Eshmarch : 5m long and 15cm wide | No | √ | √
| Size:LxWxH= 480mmx260mmx5mm | No | √ | √
3 | General Orthopaedic Instruments
3.1 | Bone Cutters:
| Ruskin Rongeurs, Straight, 7.5 inches | No | √ | √
| Ruskin Rongeurs, Curved, 7.5 inches | No | √ | √
| Ruskin Rongeurs, Straight, 9.0 inches | No | √ | √
| Ruskin Rongeurs, Curved, 9.0 inches | No | √ | √
| Ruskin-Liston Bone Cutting Forceps, Straight | No | √ | √
| Ruskin-Liston Bone Cutting Forceps, Angled | No | √ | √
3.2 | Bone Mallets:
| Bone Mallets, 250gms | No | √ | √
| Bone Mallets, 400gms | No | √ | √
3.3 | Osteotomes, Chisels and Gouges:
| Mini Osteotomes, 6 mm | No | √ | √
| Mini Osteotomes, 8 mm | No | √ | √
| Mini Osteotomes, 10 mm | No | √ | √
3.4 | Bone Curettes:
| Endoscopic Curettes, Ring Curette | No | √ | √
| Endoscopic Curettes, Oval Cup Curette | No | √ | √
3.5 | Bone Holding Instruments:
| Lane Bone Holding Forceps, 250mm, with ratchet | No | √ | √
| Lane Bone Holding Forceps, 300mm, with ratchet | No | √ | √
| Pelvic Reduction Forceps, short tips, angular, short pointed-ball tip | No | √ | √
| Pelvic Reduction Forceps, long tips, angular, long pointed-ball tips | No | √ | √
| Bone Reduction Forceps, with rachet, serrated, length 140mm | No | √ | √
| Bone Reduction Forceps, with rachet, serrated, length 170mm | No | √ | √
| Bone Reduction Forceps, with rachet, serrated, length 200mm | No | √ | √
| Bone Reduction Forceps, with rachet, pointed, length 140mm | No | √ | √
| Bone Reduction Forceps, with rachet, pointed, length 170mm | No | √ | √
| Bone Reduction Forceps, with rachet, pointed, length 200mm | No | √ | √
| Plate and Bone holding forceps, one pointed tip and foot plate | No | √ | √
3.6 | Elevators, Awls and Bone Hooks:
| Periosteal Elevators, 6mm straight edge, straight blade | No | √ | √
| Periosteal Elevators, 13mm straight edge, straight blade | No | √ | √
3.7 | Retractors:
| Homann Retractors, 8mm blade width, length 8.5 inches | No | √ | √
| Blount Retractors, double prong retractor | No | √ | √
| Humeral Head Retractor | No | √ | √
| Volkman Retractors, 2 prongs, sharp | No | √ | √
| Volkman Retractors, 3 prongs, sharp | No | √ | √
| Volkman Retractors, 4 prongs, sharp | No | √ | √
| Volkman Retractors, 2 prongs, blungt | No | √ | √
| Volkman Retractors, 3 prongs, blunt | No | √ | √
| Volkman Retractors, 4 prongs, blunt | No | √ | √
| Skin Hooks, single pronged | No | √ | √
| Skin Hooks, two pronged | No | √ | √
3.8 | Operating Scissors: Kelly 16cm, 6 1/4" | No | √ | √
| Operating Scissors: Mayo 14.5cm, 5 3/4" (straight) | No | √ | √
| Operating Scissors: Mayo 14.5cm, 5 3/4" (curved) | No | √ | √
| Operating Scissors: Mayo 17cm, 6 3/4" | No | √ | √
| Operating Scissors: Metzenbaum 11.5cm, 4 1/2" (straight) | No | √ | √
| Operating Scissors: Metzenbaum 11.5cm, 4 1/2" (curved) | No | √ | √
| Operating Scissors: Metzenbaum 14.5cm, 5 3/4" (curved) | No | √ | √
| Operating Scissors: Metzenbaum 18cm, 7" (curved) | No | √ | √
3.9 | Needle Holders : Halsey 13cm, 5 1/8" | No | √ | √
| Needle Holders : Crile-Wood 15cm, 6" | No | √ | √
| Needle Holders : Crile-Wood 23cm, 9" | No | √ | √
| Needle Holders : Mayo Hegar 16cm, 6 1/4" | No | √ | √
| Needle Holders : Mayo-Hegar 18.5cm, 7" | No | √ | √
Needle Holders : Mayo-Hegar 27cm, 10 3/4"
| Needle Holders : Kilner 13cm, 5 1/6" | No | √ | √
3.1 | Thumb Tissue Forceps: Adson, toothed, 12cm | No | √ | √
| Thumb Tissue Forceps: Adson, toothed, 15cm | No | √ | √
| Thumb Tissue Forceps: Adson-Brown, toothed, 12cm | No | √ | √
3.11 | Bone Files : Size- 22cm | No | √ | √
3.12 | Wire Instruments along with a Sterilizing tray for wire instruments consisting of : a. Wire passer, 70mm bending diameter; b. Wire holding forceps for circlage wires, L 170mm, c. Vice Grip, self retaining, L 180mm, d. Paralel pliers, flat nosed, L 160mm, e. Wire bending pliers, L 155mm , f.Wire cutter, small, L 160mm, g. Wire cutter, large, L 220mm, h.Wire Cutters, 48cm for hard wires/pins up to 6mm diam | No | √ | √
3.13 | Circlage Wires in coils of length 10m: Diameter 0.6 mm | No | √ | √
| Circlage Wires in coils of length 10m: Diameter 0.8mm | No | √ | √
| Circlage Wires in coils of length 10m: Diameter 1.0mm | No | √ | √
3.14 | Kirscher Wires with trocar tip: 1.0 mm dia. Lengths 150mm | No | √ | √
| Kirscher Wires with trocar tip: 1.25mm dia., lengths 150mm | No | √ | √
| Kirscher Wires with trocar tip: 1.6mm dia., lengths 150mm | No | √ | √
| Kirscher Wires with trocar tip: 2.0mm dia., lengths 150mm | No | √ | √
Kirscher Wires with trocar tip:2.5 mm dia., lengths 230mm
| Kirscher Wires with trocar tip:2.5 mm dia., lengths 280m | No | √ | √
| Kirscher Wires with threaded tip: 1.25mm dia., length 150mm | No | √ | √
| Kirscher Wires with threaded tip: 2.0mm dia., length 150mm | No | √ | √
| Kirscher Wires with threaded tip: 2.5mm dia., length 230mm | No | √ | √
| Kirscher Wires with threaded tip: 3.0mm dia., length 280mm | No | √ | √
4 | TRAUMATOLOGY & GENERAL OSTEOSYNTHESIS (Synthes/Zimmer or Equivalent quality products)
4.1 | Small Fragment Instrument & Implant Set- LC-DCP | No | √ | √
a | Small Fragment Instrument & Implant Set- LC-DCP Graphic Case | No | √ | √
b | Small Fragment Instrument Set, st. steel (Synthes/Zimmer or equivalent) | Complete Set | √ | √
c | Non-Locking screws, 3.5mm diameters
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 14mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 16mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 18mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 20mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, Ti, length 22mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 14mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 16mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 18mm | No | √ | √
| Cortex Screw, dia 3.5 mm, self - tapping, st. steel, length 20mm | No | √ | √
d | DCP & LC- Dynamic Compression Plate 3.5 mm
| LC- Dynamic Compression Plate 3.5 mm, Ti. 5 Holes | No | √ | √
| LC- Dynamic Compression Plate 3.5 mm, Ti. 6 Holes | No | √ | √
| LC- Dynamic Compression Plate 3.5 mm, Ti. 7 Holes | No | √ | √
| DCP 3.5mm, st. steel, 5 holes | No | √ | √
| DCP 3.5mm, st. steel, 6 holes | No | √ | √
| DCP 3.5mm, st. steel, 7 holes | No | √ | √
e | Locking 1/3 Tubular Plate 3.5mm, 5 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 6 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 7 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 8 holes, Ti: | No | √ | √
| Locking 1/3 Tubular Plate 3.5mm, 9 holes, Ti: | No | √ | √
f | Non-Locking 1/3 Tubular Plate 3.5mm, 6 holes, st. steel | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 7 holes, st. steel | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 8 holes, st. steel | No | √ | √
| Non-Locking 1/3 Tubular Plate 3.5mm, 9 holes, st. steel | No | √ | √
g | Reconstruction plate 3.5mm, 5 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 6 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 7 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 8 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 10 holes, Ti | No | √ | √
| Reconstruction plate 3.5mm, 6 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 7 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 8 holes, st. steel | No | √ | √
| Reconstruction plate 3.5mm, 10 holes, st. steel | No | √ | √
h | Cancellous screws 4.0 mm, fully threaded, self taping:
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 14mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 16mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 18mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 20mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 22mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 24mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 26mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 28mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 30mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 35mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 45mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 50mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 55mm | No | √ | √
| Cancellous screws 4.0 mm, fully threaded, self taping: Screw length 60mm | No | √ | √
i | Cancellous bone screw 4.0mm, partially threaded:
| Cancellous bone screw 4.0mm, partially threaded: Screw length 14mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 16mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 18mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 20mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 22mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 24mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 26mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 28mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 30mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 35mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 40mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 45mm | No | √ | √
| Cancellous bone screw 4.0mm, partially threaded: Screw length 50mm | No | √ | √
j | Washer, 7.0mm for small screws | No | √ | √
4.2 | Large Fragment Basic Instrument Set for LC-DCP and DCP
a | Basic Instrument Set LC-DCP and DCP Graphic Case, perforated, autoclavable, complete with cover | No | √ | √
b | Basic Instruments for LC-DCP and DCP, Synthes/Zimmer or Equivalent | Complete Set | √ | √
c | Large Fragment Instrument Implants
| Large Fragment Screws (4.5 mm): | No | √ | √
| 4.5mm cortex screws, self taping, length 30mm | No | √ | √
| 4.5mm cortex screws, self taping, length 32mm | No | √ | √
| 4.5mm cortex screws, self taping, length 34mm | No | √ | √
| 4.5mm cortex screws, self taping, length 36mm | No | √ | √
| 4.5mm cortex screws, self taping, length 38mm | No | √ | √
| 4.5mm cortex screws, self taping, length 40mm | No | √ | √
4.3 | Cannulated Cancellous Screw 4.0 mm System
a | Instrument Set for Cannulated Screw 4.0 mm, st. steel (Synthes/Zimmer or Equivalent): Contents consisting of:
| Guide wire,1.2mm dia,with threaded tip | 5 | √ | √
| Drill bit,dia.2.7mm,cannulated | 2 | √ | √
| Countersink,cannulated for 4.0mm cannulated screws | 1 | √ | √
| Tap,cannulated for 4.0mm cannulated screws | 2 | √ | √
| Screwdriver,hexagonal,cannulated sw 2.5 with holding sleeve | 1 | √ | √
| Screw holding forceps | 1 | √ | √
| Protection sleeve,8.0/3.0 | 1 | √ | √
| Guide sleeve,3.0/1.2 | 1 | √ | √
| Stylet,dia.1.2mm | 2 | √ | √
| Direct measuring device | 1 | √ | √
| Screw Extraction device | 1 | √ | √
| Screwdriver,sw2.5,solid | 1 | √ | √
| Basic tray for 4.0mm cannulated screw instrumentation | 2 | √ | √
| Top tray for 4.0mm cannulated screw instrumentation | 2 | √ | √
| Lid for 4.0mm cannulated screws instrumentation | 2 | √ | √
b | 4.0mm Cannulated Screw Box | 2 | √ | √
c | Implants (4.0mm Cannulated Screws, Ti):
| 4.0mm cannulated screws, total length/thread length : Screw length 20/7mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw lengths 22/7mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 24/8mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 26/9mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 28/9mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 30/10mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw lengths 32/11mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 34/11mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 36/12mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 38/12mm | No | √ | √
| 4.0mm cannulated screws, total length/thread length : Screw length 40/13mm | No | √ | √
| Washer 7/0/3.6 mm for screws 2.7 to 4.0 mm, st. steel | No | √ | √
4.4 | Instrument Set for Cannulated Screw 6.5 mm, st. steel (Synthes/Zimmer or Equivalent): Contents consisting of: | No | √ | √
| Basic tray for 6.5mm cannulated screws instrumentation | 1 | √ | √
| Top tray for 6.5mm cannulated screws instrumentation | 1 | √ | √
| Lid for 6.5mm cannulated screw instrumentation tray | 1 | √ | √
| Guide wire,2.0mm dia with threaded tip | 5 | √ | √
| Drill bit,dia.4.5mm,cannulated | 2 | √ | √
| Countersink, cannulated for 6.5mm cannulated screws | 1 | √ | √
| Tap ,cannulated for 6.5mm cannulated screws | 1 | √ | √
| Parallel wire guide | 1 | √ | √
| Screwdriver, hexagonal, cannulated for 6.5mm cannulated screws | 1 | √ | √
| Depth gauze - 01no | 1 | √ | √
| Drill sleeve,8.0/5.0 | 1 | √ | √
| Protection sleeve, 11.0/8.0 | 1 | √ | √
| Stylet,dia.2.0mm | 2 | √ | √
| Drill sleeve,5.0/2.0 | 1 | √ | √
| Trocar | 1 | √ | √
| Screwdriver, SW 3.5,solid | 1 | √ | √
| Parallel wire guide, adjustable | 1 | √ | √
| Spanner | 1 | √ | √
5 | Angled Blade Plate Ststem:
a | Angled Blade Plate Instrument Set | Set | √ | √
Contents consisting of:
| Grapic Case for Angled Blade plate instrument set without content | 1 | √ | √
| Kirschner wire with trocar tip, 2.0mm/150mm, SS | 5 | √ | √
| Drill Bit, 2.0mm dia, length 100/75 mm for quick coupling | 3 | √ | √
| Double Drill Sleeve 4.5/3.2 | 1 | √ | √
| Chisel Guide | 1 | √ | √
| Seating chisel, length 320mm | 1 | √ | √
| Inserter/Extractor with adjustable clamps length, 270mm | 1 | √ | √
| Slotted Hammer, length 270mm | 1 | √ | √
| Impactor | 1 | √ | √
| Postioning plate, triangular, 90°/50°/40° | 1 | √ | √
| Postioning plate, triangular, 80°/70°/30° | 1 | √ | √
| Postioning plate, triangular, 100°/60°/20° | 1 | √ | √
| Condylar plate guide | 1 | √ | √
b | Angled Blade Plate Implant:
| Angled Blade Plate 95 degree, 60 mm blade 5 holes | No | √ | √
| Angled Blade Plate 95 degree, 60 mm blade 6 holes | No | √ | √
| Angled Blade Plate 95 degree, 60 mm blade 7 holes | No | √ | √
| Angled Blade Plate 95 degree, 60 mm blade 8 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 5 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 6 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 7 holes | No | √ | √
| Angled Blade Plate 95 degree, 65 mm blade 8 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 5 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 6 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 7 holes | No | √ | √
| Angled Blade Plate 130 degree, 60 mm blade 8 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 5 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 6 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 7 holes | No | √ | √
| Angled Blade Plate 130 degree, 65 mm blade 8 holes | No | √ | √
6 | Proximal Femoral Nail (PFN) System: | No | √ | √
a | PFN Instrument Set (Contents) consisting of:
| Proctecting sleeve | 1 | √
| Guide wire sleeve | 1 | √
| Proximal cannulated drill | 1 | √
| Quick coupling handle,T | 1 | √
| Screwdriver,T | 1 | √
| Tissue proctector | 1 | √
| Insertion handle | 1 | √
| Connecting screw | 2 | √
| Connector | 1 | √
| Adaptor | 1 | √
| Sliding hammer | 1 | √
| Combination wrench | 1 | √
| Proximal device bolt | 1 | √
| Guide sleeve,8.2/2.0 | 1 | √
| Trocar,4.0 | 1 | √
| Guide pin,threaded,3.2/400 | 3 | √
| Guide wire,threaded,2.0/340 | 2 | √
| Guide wire,2.0/340 | 2 | √
| Guide wire,2.5 / 340 | 2 | √
| Guide wire,threaded,2.5 / 320 | 2 | √
| Depth gauze for proximal screw | 1 | √
| Step cannulated reamer,4.5/6.3 | 1 | √
| Step cannulated reamer,6.3 / 9 | 1
| Stoper for step reamer | 2 | √
| Depth gauze | 1 | √
| Drill sleeve, 8.2 / 4.0 | 1 | √
| Drill sleeve, 8.2 / 11 | 2 | √
| Drill sleeve, 9.2 / 11 | 1 | √
| Guide sleeve, 9.2 / 2.5 | 1 | √
| Trocar, 2.5 | 1 | √
| Trocar, 2.0 | 1 | √
| Screwdriver , sw 3.5 | 1 | √
| Wrench for conical bolt | 1 | √
| Impactor-extractor rod | 1 | √
| Proximal locking device Radiolucent | 1 | √
| Cannulated Awl | 1 | √
| Wrench, sw 3 | 1 | √
| Drill bit, Stop, 4.0 | 1 | √
| Drill bit, dia.4.0 | 2 | √
| Screwdriver,cannulated,sw 5.0 | 1 | √
| Locking screw holder | 1 | √
| Screw inserter T type for 9.0mm screw | 1 | √
| Screwdriver sleeve | 1 | √
b | PFN Implants:
| Proximal femoral nail, 130˚, short, Dia. 9mm, Length 240mm, Titanium | No. | √
| Proximal femoral nail, 130˚, short, Dia. 10mm, Length 240mm, Titanium | No. | √
| Proximal femoral nail, 130˚, short, Dia. 11mm, Length 240mm, Titanium | No. | √
| Proximal femoral nail, 130˚, short, Dia. 12mm, Length 240mm, Titanium | No. | √
| End cap for Proximal femoral nail, without extension, Titanium | No. | √
| End cap for Proximal femoral nail, 5mm extension, Titanium | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 80mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 85mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 90mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 95mm | No. | √
| Femoral neck screw for PFN, Dia. 9.0mm, self tapping, cannulated, Titanium, 100mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 70mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 75mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 80mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 85mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 90mm | No. | √
| Locking screw, Dia. 6.3mm, self tapping, cannulated, Titanium, 95mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 30mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 32mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 34mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 36mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 38mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 40mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 42mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 44mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 46mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 48mm | No. | √
| Locking screw, Dia. 4.9mm, Length 24mm, self tapping, Titanium, 50mm | No. | √
7 | Rush Rods
| Rush Rod 3mm diameter of various lengths | No. | √ | √
| Rush Rod 5mm diameter of various lengths | No. | √ | √
| Rush Rod 6mm diameter of various lengths | No. | √ | √
8 | External Fixation System:
A | Small External Fixator System: | No. | √ | √
a | Small External Fixator Graphic Case with graphic outlines, complete with cover | No. | √ | √
b | Small External Fixator Instruments, consisting of:
| Combination Wrench, 7 mm width across flats | 1 | √ | √
| Socket Wrench, 7 mm width across flats | 1 | √ | √
| Small External Fixator Drill Sleeve System | 1 | √ | √
c | Small External Fixator Implants:
| 1.6 mm Kirscher Wire, 150 mm, trocar point | No. | √ | √
| 2.0 mm Kirscher Wire, 150 mm, trocar point | No. | √ | √
| 2.5 mm Kirscher Wire, 150 mm, trocar point | No. | √ | √
d | Fixation components
| Closed clamps, 4.0 mm/2.5 mm | No. | √ | √
| Closed clamps, 4.0 mm/4.0 mm | No. | √ | √
| Open clamps, 4.0 mm/2.5 mm | No. | √ | √
| Open clamps, 4.0 mm/2.5 mm | No. | √ | √
| Spring-Loaded Nut | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 60mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 80mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 100mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 120mm, | No. | √ | √
| Connecting Rods, 4.0mm dia. carbon fiber, length 140mm, | No. | √ | √
| Connecting Bars, 4.0mm dia. carbon fiber, length 160mm, | No. | √ | √
| Connecting Bars, 4.0mm dia. carbon fiber, length 180mm, | No. | √ | √
| Connecting Bars, 4.0mm dia. carbon fiber, length 200mm, | No. | √ | √
e | Schanz Screws for small & large external fixator, thread length 20mm | No. | √ | √
| Schanz Screw 4.0/3.0 self drilling, total length 80mm | No. | √ | √
| Schanz Screw 4.0/3.0 self drilling, total length 100mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 60mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 80mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 100mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 120mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 140mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 160mm | No. | √ | √
| Schanz Screw 4.0mm, self drilling, total length 175mm | No. | √ | √
| Schanz Screw, 4.5mm, self drilling, total length 150mm | No | √ | √
| Schanz Screw, 4.5mm, self drilling, total length 175mm | No | √ | √
| Schanz Screw, 5.0mm, self drilling, total length 150mm | No | √ | √
| Schanz Screw, 5.0mm, self drilling, total length 150mm | No | √ | √
B | Large External Fixator System:
a | Large External Fixator Instruments, consisting of:
| Combination Wrench, 11 mm width across flats | 1 | √ | √
| Socket Wrench, 11 mm width across flats | 1 | √ | √
| Large External Fixator Drill Sleeve System | 1 | √ | √
b | Large External Fixator Components:
| Adjustable clamps | No | √ | √
| Open clamps | No | √ | √
| Universal Joint for two tubes | No | √ | √
| Tube-to-tube clamps | No | √ | √
| Transverse clamp | No | √ | √
| Combination clamp | No | √ | √
| Universal clamp | No | √ | √
c | Connecting Rods, 11.0mm dia. carbon fiber
| Connecting Rods, 11.0mm dia. carbon fiber, length 100mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 150mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 200mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 250mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 300mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 350mm, | No | √ | √
| Connecting Rods, 11.0mm dia. carbon fiber, length 400mm, | No | √ | √
d | Reduction wires (olive wire) for Illizarov Fixator
| Reduction wire (olive), 1.8mm | No | √ | √
| Reduction wire (olive), 1.8mm | No | √ | √
9 | Arthroplasty System
a | Prosthesis for Hemiarthroplasty:
| Austin Moore Endoprosthesis, Head 41.0 mm, Stem 140mm | No | √ | √
| Austin Moore Endoprosthesis, Head 43.0 mm, Stem 140mm | No | √ | √
| Austin Moore Endoprosthesis, Head 45.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 47.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 49.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 51.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 53.0 mm, Stem 160mm | No | √ | √
| Austin Moore Endoprosthesis, Head 55.0 mm, Stem 160mm | No | √ | √
b | Bipolar Prosthesis:
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 37.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 39.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 41.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 43.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 45.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 47.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 49.0 mm, Stem length 155 mm | No | √ | √
| New Bipolar Prosthesis, Non-fenestrated: Head diameter 51.0 mm, Stem length 155 mm | No | √ | √
10 | Spine Implants
a | Pedicle screw ( Polyaxial): Synthes/Zimmer or Equivalent
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 30 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 32 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 35 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 38 mm | No. | √
| Pedicle screw( Polyaxial): Dia 4.2 mm x L 40 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 30 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 35 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 40 mm | No. | √
| Pedicle screw( Polyaxial): Dia 5.0 mm x L 45 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 35 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 40 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 45 mm | No. | √
| Pedicle screw( Polyaxial): Dia 6.0 mm x L 50 mm | No. | √
b | Pedicle screw( Monoaxial): Synthes/Zimmer or Equivalent:
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 30 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 32 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 35 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 38 mm | No. | √
| Pedicle screw( Monoaxial): Dia 4.2 mm x L 40 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 30 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 35 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 40 mm | No. | √
| Pedicle screw( Monoaxial): Dia 5.0 mm x L 45 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 35 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 40 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 45 mm | No. | √
| Pedicle screw( Monoaxial): Dia 6.0 mm x L 50 mm | No. | √
c | Nut for Pedicle screw | No. | √
d | Hard Rods:
| 6.0mm hard rods: Length 250mm | No. | √
| 6.0mm hard rods: Length 350mm | No. | √
| 6.0mm hard rods: Length 450mm | No. | √
11 | Spare Parts for Existing Instruments
a | Synthes DHS/DCS Instrument Spare Parts:
| 338.00, 2.5mm Threaded Guide Wires | No | 10 Nos
| 338.20, DHS/DCS Coupling Screw, Short | No | 3 Nos
| 338.21, DHS/DCS Guide Shaft | No | 3 Nos
| 338.010, Drill Bit | No | 1 No
| 338.47, DCS Reaming Head | No | 1 No
b | Karl Storz Arthroscopy Equipment Spare Parts:
| 495 ND Fiber Optic Cable, 3.5mm, length 400cm | No | 1 No
c | Power Drill Spare Parts:
| ConMed Power Pro, PRO200, Linvatec Cable Ref MC5057 | No | 1 No
| Rodar Power Control Unit, Cable Part number 15.01 | No | 1 No
d | Arthroscopic Surgery Accessory Supplies:
| HOPKINS II Telecope 30 degrees, 4mm, 18cm | Piece | 1
| Take-apart Scissors | Piece | 1

## pathology
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
1 | Pathology Department: Microbiology
1.1 | Culture Media
1.1.1 | Agarose | 100 gm | √
1.1.2 | Alkaline peptone water | 100gm | √ | √ | √
1.1.3 | Bile esculin agar | 100gm | √ | √ | √
1.1.4 | Bismuth sulphate agar | 500gm | √ | √ | √
1.1.5 | Bolton broth | 500gm | √
1.1.6 | Brain heart infusion agar | 500gm | √ | √
1.1.7 | Brain heart infusion broth | 500gm | √ | √ | √
1.1.8 | Cary-blair medium | 500gm | √ | √ | √
1.1.9 | Columbia blood agar base | 500gm | √ | √ | √
1.1.10 | Cooked meat medium | 500gm | √
1.1.11 | Corn meal agar | 500gm | √
1.1.12 | Dnase test agar with toludine blue | 100gm | √ | √ | √
1.1.13 | Thioglycollate Broth | 100gm | √ | √
1.1.14 | Gc agar Base | 500gm | √ | √ | √
1.1.15 | Hippurate hydrolysis broth | 100gm | √ | √ | √
1.1.16 | Luria-Bertani broth | 100gm | √
1.1.17 | Nutrient broth | 500gm | √ | √ | √
1.1.18 | MacConkey agar with CV & Bile salt | 500gm | √ | √ | √
1.1.19 | MacConkey agar without salt | 500gm | √ | √
1.1.20 | Haemophilus test medium | 500gm | √ | √ | √
1.1.21 | Malonate broth | 100gm | √ | √ | √
1.1.22 | Modified CCDA agar | 500gm | √
1.1.23 | MR-VP medium | gm | √ | √ | √
1.1.24 | Mueller hinton agar | 500gm | √ | √ | √
1.1.25 | Nitrate broth | 500gm | √ | √ | √
1.1.26 | Nutrient agar | 500gm | √ | √
1.1.27 | Nutrient gelatin | 100gm | √ | √ | √
1.1.28 | Oxidative Fermentative Basal medium | 100gm | √ | √ | √
1.1.29 | Ornithine decarboxylase broth | 100gm | √ | √ | √
1.1.30 | Lysin decarboxylase broth | 100gm | √ | √ | √
1.1.31 | Arginine decarboxylase broth | 100gm | √ | √ | √
1.1.32 | Rice extract agar | 500gm | √
1.1.33 | Sabrouad dextrose agar | 500gm | √ | √
1.1.34 | Sabrouad dextrose agar with antibiotic supplements | 500gm | √ | √
1.1.35 | Selenite broth | 500gm | √ | √ | √
1.1.36 | SIM medium | 500gm | √ | √ | √
1.1.37 | Simmons citrate agar | 100gm | √ | √ | √
1.1.38 | Soyabean casein digest medium with sps | 500gm | √ | √
1.1.39 | TCBS cholera medium | 100gm | √ | √
1.1.40 | Triple sugar iron agar | 100gm | √ | √ | √
1.1.41 | Urea broth base | 100gm | √ | √ | √
1.1.42 | XLD agar | 500gm | √ | √ | √
1.1.43 | Cystein lactose electrolite deficient agar (CLED) | 500gm | √ | √ | √
1.1.44 | MR ( Methyl Red) broth - peptone - dextrose - potassium- l napthol | 500 gm pack each | √
1.2 | Growth Supplements
1.2.1 | Haemoglobin powder (to be used with GC Agar) | 100gm | √ | √ | √
1.2.2 | Urea 40% solution | 5ml | √ | √ | √
1.2.3 | VCNT supplement | gm | √ | √ | √
1.2.4 | Vitamin growth supplement | gm | √ | √ | √
1.3 | Antibiotics Disc
1.3.1 | Amikacin, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.2 | Amphicilin, 10mcg | 50 Disc/ vial | √ | √ | √
1.3.3 | Aztreonam (30 mcg) | Disc | √ | √ | √
1.3.4 | Carbenicillin 100mcg | Disc | √ | √
1.3.5 | Cefoparazone (30 mcg) | Disc | √ | √
1.3.6 | Cefazolin 30mcg | 50 Disc/ vial | √ | √ | √
1.3.7 | Cefepime (30mcg) | Disc | √ | √ | √
1.3.8 | Cefixime (5mcg) | Disc | √ | √ | √
1.3.9 | Cefotoxime/Calvunalic acid (30/10 mcg) | 50 Disc/ vial | √ | √ | √
1.3.10 | Cefoxitin disc (10mcg) | 50 Disc/ vial | √ | √ | √
1.3.11 | Cefpirome (30 mcg) | Disc | √ | √
1.3.12 | Cefpodoxime - (30mcg) | Disc | √ | √
1.3.13 | Ceftazidime 30mcg | 50 Disc/ vial | √ | √ | √
1.3.14 | Ceftazidime/Calvunalic acid (30/10 mcg) | 50 Disc/ vial | √ | √ | √
1.3.15 | Ceftriaxone, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.16 | Cefuroxime (30mcg) | Disc | √ | √ | √
1.3.17 | Cephalexin, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.18 | Cephotaxime, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.19 | Chloramphenicol, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.20 | Ciprofloxacin, 5mcg | 50 Disc/ vial | √ | √ | √
1.3.21 | Clarithromycin 15mcg | 50 Disc/ vial | √ | √ | √
1.3.22 | Clindamycin 2mcg | 51 Disc/ vial | √ | √ | √
1.3.23 | Cotri-moxazole, 25mcg | 50 Disc/ vial | √ | √ | √
1.3.24 | Doxycycline Hcl, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.25 | Erythromycin, 15mcg | 50 Disc/ vial | √ | √ | √
1.3.26 | Gentamicin, 10mcg | 50 Disc/ vial | √ | √ | √
1.3.27 | Gentamicin, 120mcg | 51 Disc/ vial | √ | √ | √
1.3.28 | Imipenam (10mcg) | 50 Disc/ vial | √ | √ | √
1.3.29 | Meropenem (10mcg) | 51 Disc/ vial | √ | √ | √
1.3.30 | Nalidixic acid, 10mcg | 50 Disc/ vial | √ | √ | √
1.3.31 | Nitrofurantoin, 300mcg | 50 Disc/ vial | √ | √ | √
1.3.32 | Norfloxacin, 10mcg | 50 Disc/ vial | √ | √ | √
1.3.33 | Novobiocin Disc (5mcg ) | 50 Disc/ vial | √ | √ | √
1.3.34 | Oxacillin 1mcg | 50 Disc/ vial | √ | √ | √
1.3.35 | Penicillin G, 10units | 50 Disc/ vial | √ | √ | √
1.3.36 | Pipracillin, 100mcg | 50 Disc/ vial | √ | √ | √
1.3.37 | Polymixin B 300mcg | 50 Disc/ vial | √ | √ | √
1.3.38 | Tobramycin, 10mcg | 50 Disc/ vial | √ | √ | √
1.3.39 | Vancomycin, 30mcg | 50 Disc/ vial | √ | √ | √
1.3.40 | Agumentin | 50 disc/vial | √
1.3.41 | Amphicillin | 50 disc/vial | √
1.3.42 | E-test Cefotoxime (NCCLS) | 100 Strips | √ | √ | √
1.3.43 | E-test Ceftriaxone (NCCLS) | 100 Strips | √ | √ | √
1.3.44 | E-test Penicillin (NCCLS) | 100 Strips | √ | √ | √
1.3.45 | E-test Meropenem (NCCLS) | 100 Strips | √ | √ | √
1.3.46 | E-test Vancomycin (NCCLS) | 100 Strips | √ | √ | √
1.3.47 | E test Oxacillin | 100 strips | √
1.4 | Idenfication disc
1.4.1 | Bacitracin disc | 50 Disc/ vial | √ | √ | √
1.4.2 | Optochin disc | 50 Disc/ vial | √ | √ | √
1.4.3 | Oxidase disc | 50 Disc/ vial | √ | √ | √
1.4.4 | Nitrocefin beta - lactamase | 50 Disc/ vial | √ | √ | √
1.4.5 | V Factor disc | 50 Disc/ vial | √ | √ | √
1.4.6 | X factor disc | 50 Disc/ vial | √ | √ | √
1.4.7 | X+V Factor disc | 50 Disc/ vial | √ | √ | √
1.5 | Stains & Chemicals
1.5.1 | Acetone, ExcelaR | 500 ml | √ | √ | √
1.5.2 | Laboratory disinfectent solution | 5Ltrs | √ | √ | √
1.5.3 | Alberts stain A | 125 ml | √ | √
1.5.4 | Alberts stain B | 125 ml | √ | √
1.5.5 | Alcohol based antiseptic handrub | 500 ml | √ | √ | √
1.5.6 | Amyl alcohol | 500 ml | √ | √
1.5.7 | Anaerobic gas generating kit | pkt | √
1.5.8 | Auramine O | 25 gm | √
1.5.9 | Basic fuchsin | 25 gm | √ | √ | √
1.5.10 | Grams crystal violet | 125 ml | √ | √ | √
1.5.11 | Ethidium Bromide | gm | √
1.5.12 | Glycerol (reagent grade) | 500 ml | √ | √
1.5.13 | Grams iodine | 125ml | √ | √ | √
1.5.14 | Lactophenol cotton blue | 125ml | √ | √
1.5.15 | Methylene blue powder | 25gm | √ | √ | √
1.5.16 | N-Acethyl-l-cysteine | 100gm | √ | √
1.5.17 | N-naphthylethylenediamine dihydrochloride | 25 gm | √ | √
1.5.18 | Phenol crystals | 500 gm | √ | √ | √
1.5.19 | Potassium Iodate | 100 gm | √ | √ | √
1.5.20 | Potassium permanganate | 500 gm | √
1.5.21 | Safranin, 0.5% | 125 ml | √ | √ | √
1.5.22 | Sodium Nitrate | 100gm | √ | √
1.5.23 | Spore strips | Strip | √ | √ | √
1.5.24 | Sulfanilamide | 100gm | √ | √
1.6 | Plastic wares/ Other
1.6.1 | Gas Regulator for Co2 | piece | √
1.6.2 | Gas Regulator for mix Gas cyclinder | piece | √
1.6.3 | Antibiotic Zone Scale TM-C | piece | √ | √ | √
1.6.4 | Test tube caps 12/13mm - blue | piece | √ | √ | √
1.6.5 | Test tube caps 12/13mm - silver | piece | √ | √ | √
1.6.6 | Test tube caps 12/13mm - red | piece | √ | √ | √
1.6.7 | Test tube caps 15/16mm - silver | piece | √ | √ | √
1.6.8 | Bottle w/pouring ring & screw cap (Max. temp. 140˚C) 1000ml | piece | √ | √
1.6.9 | Bottle w/pouring ring & screw cap (Max. temp. 140˚C) 500ml | piece | √ | √
1.6.10 | Conical Tubes (Falcon,polypropyline), 15ml | piece | √ | √
1.6.11 | Conical Tubes (Falcon, polypropyline), 50ml | piece | √ | √
1.6.12 | Cryovials (2ml) | 1000/ pkt | √ | √ | √
1.6.13 | McCartney vial w/aluminium caps w/black rubber seals 14 ml | piece | √ | √
1.6.14 | Micropipette tips: 0.5 - 20 µl | piece | √ | √
1.6.15 | Metal loop TM. CH, changeable | piece | √ | √ | √
1.6.16 | Nichrome loop-D-2 | piece | √ | √ | √
1.6.17 | Nichrome loop-D-1 | piece | √ | √ | √
1.6.18 | Petri plates sterile 90mm diameter. | piece | √ | √ | √
1.6.19 | Sterile disposable pipette, 5ml (Individually wrapped) | 100/pkt | √ | √ | √
1.6.20 | Sterile swab, general purpose with screw cap | 100/ pkt | √ | √ | √
1.6.21 | Sterile swab with wooden stick | 100/ pkt | √ | √ | √
1.6.22 | Sterile screw capped stool container with attached scoop (20ml) | 100/pkt | √ | √ | √
1.6.23 | Sterile, single use filter unit 0.20 µm | piece | √ | √
1.6.24 | Sterile screw caped urine container (50ml) | 500/ pkt | √ | √ | √
1.6.25 | Tips rack, resuable & autoclavable for tips 0.1 - 10 µl, 34 mm | piece | √
1.6.26 | Tips rack, resuable & autoclavable for tips: 0.5 - 20 µl, 46 mm | piece | √
1.6.27 | Tube racks, 24 place capacity for 1.5/2.0 ml | piece | √ | √ | √
1.6.28 | Tube, safe lock for PCR, 1.5 ml | piece | √
1.6.29 | Tube, safe lock for PCR, 2.0 ml | piece | √
1.6.30 | Transparent autoclavable plastic bag (1 ltr) | piece | √ | √
1.7 | Glasswares
1.7.1 | Aspirator bottle with tap, Polythene bottle with tubulation, Capacity: 20litres | piece | √ | √ | √
1.7.2 | Burner, Bunsen (ordinary) for LPG | piece | √ | √ | √
1.7.3 | Gel Electrophoresis Chamber Set with power supply | piece | √
1.7.4 | Glass slide with single cavity | piece | √ | √ | √
1.7.5 | Mix gas (5 % oxygen 15% Co2 85% Nitrogen) with regulator | piece | √
1.7.6 | PCR Gel casting & running tray, with combs | piece | √
1.7.7 | Petri plate Rack | piece | √ | √
1.7.8 | Slide, slope-rack | piece | √ | √
1.7.9 | UV face shield | piece | √
1.7.10 | UV transparent Gel scoops for 10 x 7cm | piece | √
1.8 | Test kit |  | √ | √ | √
1.8.1 | API 20 C | 25 test | √
1.8.2 | API 20 strept | 25 test | √
1.8.3 | API 20A | 25 test | √
1.8.4 | API 20E with complete reagent kit | 25 test | √ | √
1.8.5 | API 20 GN with complete reagent kit | 25 test | √ | √
1.8.6 | API Campy with complete reagent kit | 25 test | √
1.8.7 | API NH with complete reagent kit | 12 test | √ | √
1.8.8 | Rapid 1D 32 E (API) | 25 tests/pkt | √
1.8.9 | Bacterial antigen test kit | 50 test | √ | √
1.8.10 | Helicobacter pylori Elisa test kit(IgM) | 96 test/Kit | √
1.8.11 | Helicobacter pylori Elisa test kit(IgG) | 96 test/Kit | √
1.8.12 | Mineral oil | 100 ml | √ | √
1.8.13 | Rota-Virus, latex test kit | 25 test | √ | √
1.8.14 | Streptococcal grouping test kit | 50 test | √ | √ | √
1.8.15 | DNA extraction kit (bacterial) | test | √
1.8.16 | dNTP | ml | √
1.8.17 | 50 milli mole Magnesium Chloride | gm | √
1.8.18 | TAE buffer (10x) | ml | √
1.8.19 | Taq polymerase with buffer and MgCl2 500unit | piece | √
1.8.20 | TBE buffer (10x) | ml | √
1.8.21 | Kala Azar RK39 Rapid test kit | 25 test/ kit | √ | √
1.8.22 | Dengue Rapid test kit(IgM,IgG) | 25 test/ kit | √ | √ | √
1.8.23 | Dengue Elisa test Kit(IgG) | 96 wells | √ | √
1.8.24 | Scrub typhus Rapid test kit(IgM,IgG) | 30 test/ kit | √ | √ | √
1.9 | Anti Sera
1.9.1 | Escherichia coli Agglutinating Sera poly 2 | 2ml | √
1.9.2 | Escherichia coli Agglutinating Sera poly 2 - 3 | 2ml | √
1.9.3 | Escherichia coli Agglutinating Sera poly 2 - 4 | 2ml | √
1.9.4 | Escherichia coli type 0111:K58(B4) | 2ml | √
1.9.5 | Escherichia coli type 0112:K66(B11) | 2ml | √
1.9.6 | Escherichia coli type 0114:K90(B) | 2ml | √
1.9.7 | Escherichia coli type 0119:K69(B14) | 2ml | √
1.9.8 | Escherichia coli type 0124:K72(B17) | 2ml | √
1.9.9 | Escherichia coli type 0125:K70(B15) | 2ml | √
1.9.10 | Escherichia coli type 0126:K71(B16) | 2ml | √
1.9.11 | Escherichia coli type 0127:K63(B8) | 2ml | √
1.9.12 | Escherichia coli type 0128:K67(B12) | 2ml | √
1.9.13 | Escherichia coli type 0142:K86(B) | 2ml | √
1.9.14 | Escherichia coli type 018C:K77(B21) | 2ml | √
1.9.15 | Escherichia coli type 026:K60(B6) | 2ml | √
1.9.16 | Escherichia coli type 044:K74(L) | 2ml | √
1.9.17 | Escherichia coli type 055:K59(B5) | 2ml | √
1.9.18 | Escherichia coli type 086:K61(B7) | 2ml | √
1.9.19 | Salmonella Agglutinating Sera phase 2 | 2ml | √ | √
1.9.20 | Salmonella Agglutinating Sera polyvalent H, phase 1 & 2 | 2ml | √ | √
1.9.21 | Salmonella Agglutinating Sera (1,2 - H) | 2ml | √
1.9.22 | Salmonella Agglutinating Sera (1,5 - H) | 2ml | √
1.9.23 | Salmonella Agglutinating Sera (1,7 - H) | 2ml | √
1.9.24 | Salmonella Agglutinating Sera (a - H) | 2ml | √ | √
1.9.25 | Salmonella Agglutinating Sera ( b - H) | 2ml | √ | √
1.9.26 | Salmonella Agglutinating Sera (c - H ) | 2ml | √ | √
1.9.27 | Salmonella Agglutinating Sera (d - H) | 2ml | √ | √
1.9.28 | Salmonella Agglutinating Sera (e, n, z 15 - H) | 2ml | √
1.9.29 | Salmonella Agglutinating Sera (eh - H) | 2ml | √ | √
1.9.30 | Salmonella Agglutinating Sera (fg - H) | 2ml | √ | √
1.9.31 | Salmonella Agglutinating Sera (gm - H) | 2ml | √ | √
1.9.32 | Salmonella Agglutinating Sera ( i - H) | 2ml | √ | √
1.9.33 | Salmonella Agglutinating Sera (iv - H) | 2ml | √ | √
1.9.34 | Salmonella Agglutinating Sera (k - H) | 2ml | √ | √
1.9.35 | Salmonella Agglutinating Sera (mt - H | 2ml | √ | √
1.9.36 | Salmonella Agglutinating Sera (r - H) | 2ml | √ | √
1.9.37 | Salmonella Agglutinating Sera polyvalent O, Groups A - G | 2ml | √ | √
1.9.38 | Salmonella Agglutinating Sera (2-0) | 2ml | √ | √
1.9.39 | Salmonella Agglutinating Sera (4-0) | 2ml | √ | √
1.9.40 | Salmonella Agglutinating Sera ( 6, 7-0) | 2ml | √ | √
1.9.41 | Salmonella Agglutinating Sera ( 8-0) | 2ml | √ | √
1.9.42 | Salmonella Agglutinating Sera ( 9-0) | 2ml | √ | √
1.9.43 | Salmonella Agglutinating Sera polyvalent O (Vi) | 2ml | √ | √
1.9.44 | Shigella Flexneri polyvalent 1 - 6, X & Y | 2ml | √ | √
1.9.45 | Shigella Sonnei phases 1 & 2 | 2ml | √ | √
1.9.46 | Shigella Agglutinating Sera dysentry Polyvalent 1-10 | 2ml | √ | √
1.9.47 | Shigella Boydii polyvalent 1 (1 - 6) | 2ml | √ | √
1.9.48 | Shigella Boydii polyvalent 2 (7-11) | 2ml | √ | √
1.9.49 | Shigella Boydii polyvalent 3 (12-15) | 2ml | √ | √
1.9.50 | Vibrio cholera polyvalent O-1 | 2ml | √ | √
1.9.51 | Vibrio cholera polyvalent Inaba | 2ml | √ | √
1.9.52 | Vibrio cholera polyvalent Ogawa | 2ml | √ | √
1.1 | Equipment specific Reagents
1.10.1 | FA (Fan Aerobic) for Bact/Alert 3D | 100 bottles/ box | √
1.10.2 | PF (Pediatric Fan) for Bact/Alert 3D | 100 bottles/ box | √
1.10.3 | FA (Fan Aerobic) for BD 9050 | 50 bottles/ box | √ | √ | √
1.10.4 | PF (Pediatric Fan) for BD 9050 | 50 bottles/ box | √ | √ | √
1.10.5 | Flash Disc Bact Alert Blood culture system | piece | √
1.10.6 | Reflectance Standards for Bact/Alert system (A & B) | piece | √
1.11 | Control strains
1.11.1 | Escherichia coli ATCC 25922 | loop | √ | √ | √
1.11.2 | Pseudomonas aeruginosa ATCC 27853 | loop | √ | √ | √
1.11.3 | Staphylococcus aureus ATCC 25923 | loop | √ | √ | √
1.11.4 | Klebsiella pneumoniae ATCC BAA 1705 | loop | √
1.11.5 | Klebsiella pneumoniae ATCC BAA1706 | loop | √
1.12 | Infectious Disease Serology Reagent
1.12.1 | Anti-HBc for Vitrous ECIQ System | ml | √ | √
1.12.2 | Anti-HBs for Vitrous ECIQ System | ml | √ | √
1.12.3 | Anti-HCV for Vitrous ECIQ System | ml | √ | √
1.12.4 | Anti-HIV for Vitrous ECIQ System | ml | √ | √
1.12.5 | CMV IgG for Vitrous ECIQ System | ml | √ | √
1.12.6 | CMV IgM for Vitrous ECIQ System | ml | √ | √
1.12.7 | HBeAg for Vitrous ECIQ System | ml | √ | √
1.12.8 | HBsAg for Vitrous ECIQ System | ml | √ | √
1.12.9 | Rubella IgG for Vitrous ECIQ System | ml | √ | √
1.12.10 | Rubella IgM for Vitrous ECIQ System | ml | √ | √
1.12.11 | Toxo IgM for Vitrous ECIQ System | ml | √ | √
1.12.12 | Toxo Plasma IgG for Vitrous ECIQ System | ml | √ | √
1.12.13 | Anti HBC | ml | √ | √
1.12.14 | Anti HBs calibrator | ml | √ | √
1.12.15 | Anti HIV calibrator | ml | √ | √
1.12.16 | CMV Calibrator | ml | √ | √
1.12.17 | CMV Control | ml | √ | √
1.12.18 | HBc IgM calibrator | ml | √ | √
1.12.19 | HBe calibrator | ml | √ | √
1.12.20 | HBeAg calibrator | ml | √ | √
1.12.21 | HBsAg calibrator | ml | √ | √
1.12.22 | HCV calibrator | ml | √ | √
1.12.23 | Rubella IgG Calibrator | ml | √ | √
1.12.24 | Rubella IgG control | ml | √ | √
1.12.25 | Toxo IgM calibrator | ml | √ | √
1.12.26 | Toxo Plasma IgM controls | ml | √ | √
1.12.27 | Anti - HBC IgM Controls | ml | √ | √
1.12.28 | Anti HBe Controls | ml | √ | √
1.12.29 | Anti HBs Controls | ml | √ | √
1.12.30 | Anti HIV Controls | ml | √ | √
1.12.31 | HBsAg Controls | ml | √ | √
1.12.32 | HCV Controls | ml | √ | √
1.12.33 | HBc lgM ( 52 well ) for Vitrous ECIQ System | ml | √ | √
1.12.34 | HBe for Vitrous ECIQ System | ml | √ | √
1.12.35 | Diluent A | ml | √ | √
1.12.36 | Diluent B | ml | √ | √
1.12.37 | Signal Reagent | ml | √ | √
1.13 | Rapid/Elisa test Kit
1.13.1 | HBsAg Rapid Test Kit, sensitivity 100%, Specificity >98%, Reference Lab evaluated. Assay type: Rapid Diagnostic Test, Immunochromatography (lateral-flow tests or immunofiltration) Antibody type: monoclonal Solid phase: Nitrocellulose membrane strip Specimen type: Serum/plasma from venous whole blood, Detection type: HBsAg, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
1.13.2 | HCV Rapid Test Kit, sensitivity 100%, Specificity >98%, Reference Lab evaluated. Assay type: Rapid Diagnostic Test, Immunochromatography (lateral-flow testsorimmunofiltration) Antigen type: Recombinant proteins synthetic peptides core and NS3 and NS 5 antigens Solid phase: membrane strip Specimen type: Serum/plasma /blood Detection type: HCV antibodies, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
1.13.3 | HIV Rapid Test Kit: WHO prequalified,Reference Lab evaluated. sensitivity 100%, specificity >98%. Assay type: Rapid Diagnostic Test, Immunochromatography (lateral-flow tests or immunofiltration), Antigen type: Recombinant and synthetic peptides for HIV-1 gp41, HIV-1 gp120, and HIV-2 gp36, Solid phase: Nitrocellulose membrane strip, Specimen type: Serum/plasma / whole blood, Detection type: Combined detection of HIV-1/2 antibodies Subtype detection: M N O sub types, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
1.13.4 | RPR floculation test kits. Assay type: Macroscopic flocculation test Antibody type: Monoclonal antibodies Specimen type: Serum/plasma Detection type: Anti- lipoidal antibodies, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
1.13.5 | TPHA Rapid Test Kit, Assay type: Hemagglutination assay/card test/rapid test Antibody type: Monoclonal antibodies Specimen type: Serum/plasma Detection type: Antibodies against the causative agent of syphilis, Storage conditions: + 2 to +30 ˚ C, Statement of intent: Both for blood donor screening and patient diagnosis, Shelf life : After receiving at central depot should be 12 months | 100 test/ Kit | √ | √ | √
1.13.6 | Anti-HIV 1 &2, ELISA, WHO prequalified, Reference Lab evaluated, sensitivity 100%, Specificity >99.4%, Analytical sensitivity: must be less than 60 picograms/ml and /or less than 0.050 IU/ml as per different standardsAssay type: 3rd generation 'Sandwich' ELISA, Antigen type:Recombinant and synthetic peptides for HIV-1 gp41, HIV-1 0 gp41, and HIV-2 gp36, Solid phase:12 x 8 microwell strips, Specimen type: Serum/plasma, Detection type:Combined detection of HIV-1/2 antibodies, Subtype detection: M N O sub types, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life : After receiving at central depot should be 12 months at 2-8 C. | 96 test/192 tests/ Kit | √ | √
1.13.7 | HBsAg ELISA, Reference Lab evaluated, 100%, Specificity >98%, Analytical sensitivity: Must detect less than 60 picograms/mland /or less than 0.050 IU/ml as per different standards, Assay type: 3rd generation 'Sandwich' ELISA, Antibody type: Monoclonal antibodies, Solid phase: 12 x 8 microwell strips, Specimen type: Serum/plasma Detection type: detection of HBsAg, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life: After receiving at central depot should be 12 months at 2-8 C. | 96 test/ Kit | √ | √
1.13.8 | Anti-HCV ELISA,sensitivity 100%, Specificity >98%, Reference Lab evaluated, Assay type: 3rd generation 'Sandwich' ELISA, Analytical sensitivity: must be less than 60 picograms/ml and /or less than 0.050 IU/ml as per different standards, Antigen type: Encoded with recombinant antigens to Core, NS3, NS4and NS 5 antigens, Solid phase: 12 x 8 microwell strips, Specimen type: Serum/plasma Detection type: Detection of HCV antibodies. Must detect genotypes 1,3 and 6 distributed in Asian population and subtypes of HCV, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life: After receiving at central depot should be 12 months at 2-8 C. | 96 test/ Kit | √ | √
1.13.9 | HIV-1/2 is an in vitro diagnostic test kit for the detection of antibodies to Human Immunodeficiency Viruses type1 and/or type2 (HIV-1/HIV-2) in human serum/plasma. sensitivity 100%, Specificity >98%, Reference Lab evaluated, WHO prequalified, Shelf life : After receiving at central depot should be 12 months at 2-8 C. | 96/ test/Kit
1.13.10 | Infectious mononucleosis test | 20 test/ Kit | √ | √
1.13.11 | Typhidot test for Typhoid | 25 test/ kit | √ | √ | √
1.13.12 | Tuberculin PPD (dil) 5TU/0.1ml | 5ml | √ | √ | √
1.13.13 | TORCH PANEL
1.13.14 | Toxoplasma IgG ELISA | 96 wells | √
1.13.15 | Toxoplasma IgM ELISA | 96 Wells | √
1.13.16 | Rubella IgG ELISA | 96 wells | √
1.13.17 | Rubella IgM ELISA | 96 wells | √
1.13.18 | CMV IgG ELISA | 96 Wells | √
1.13.19 | CMV IgM ELISA | 96 Wells | √
1.13.20 | Herpes IgG ELISA | 96 Wells | √
1.13.21 | Herpes IgM ELISA | 96 Wells | √
1.13.22 | HBeAg ELISA | 96 wells | √
1.13.23 | Ciprofloxacin(1ug) | 50 Disc/ vial | √ | √ | √
1.13.24 | Ceftriaxone(0.5ug) | 50 Disc/ vial | √ | √ | √
1.13.25 | Penicillin G(0.5ug) | 50 Disc/ vial | √ | √ | √
1.13.26 | Tetracycline(10ug) | 50 Disc/ vial | √ | √ | √
1.13.27 | Cefpodoxime(10ug) | 50 Disc/ vial | √ | √ | √
1.13.28 | Azithromycin | 50 Disc/ vial | √ | √ | √
1.14 | Reagents for CD4 counter
1.14.1 | BD FACS Count Control | 25 test/ pkt | √ | √ | √
1.14.2 | BD FACS Count Reagent | 50 test/ pkt | √ | √ | √
1.14.3 | BD FACS Flow Solution | 20Ltrs | √ | √ | √
1.14.4 | Thermal paper Printer roll for BD FACS count | rolls | √ | √ | √
1.14.5 | CD4 count, Pima Bead Std | Pkt | √ | √ | √
1.14.6 | CD4 count, Pima Cartridge | 100test/ kit | √ | √ | √
1.14.7 | Echinococcus ELISA | 96test/ Pkt | √
1.14.8 | Kala azar rapid test kits | 96test/ Pkt | √ | √ | √
1.14.9 | Anaerobic indicator | 50 pcs/ pkt | √ | √
1.14.10 | Catalyst | 10 pcs/ pkt | √
1.14.11 | Anaerobic Mixed gas (5 % H2 5% Co2 90% Nitrogen) with regulator | Pcs | √
1.15 | Clinical Parasitology
1.15.1 | IUI- sperm washing media/ sperm rinse ( Bicarbonate and HEPES buffered medium containing human serum albumin) 21 mM HEPES buffer 4 mM Sodium Bicarbonate 5 mg/ml (0.5%) Human Serum Albumin Storage 2-8°C Shelf Life-2 years from date of manufacture | 12 x12 mL, 100 mL. | √
Any equipments selected, the supplier should be able to install the equipment , give training to the staff and the Biomedical unit in the hospital, only then the equipment will be selected.
1.16 | Equipments
1.16.1 | Balance: 0.0001g Readability. Maximum capacity: 210g, Readability: 0.0001g. Settling time (typical): 10…16s, Adjustment with internal weights: proFACT, fully automatic time- and temperature-controlled adjustment. Adjustment with external weights: custom weight. Linearization: automatic linearization with each internal adjustment. Sensitivity (temperature Sampling pump drift): ± 1.0 ppm/°C. Sensitivity (long-term stability): ± 0.0001%. Free height above weighing pan: 55mm. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
1.16.2 | Media Dispenser: Range: 0.5ml - >5 liters. Accuracy,%: ± at 10 ml. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
1.16.3 | Cellulose acetate sterile syringe filter: Filter pore size 0.45 µm, 25 mm diameter membrane, surfactant free-cellulose acetate (SFCA). sample volume: 0.5 ml to 100 ml. | Set | √ | √
1.16.4 | Micropipette: adjustable, volume range: 0.1 - 2.5 µl | Piece | √
1.16.5 | Micropipette: adjustable, volume range: 2 - 25 µl | Piece | √
1.16.6 | 8 Channel Micropipette: adjustable, volume range: 50 - 200 µl | piece | √ | √
1.16.7 | Micro-centrifuge: Microprocessor controlled. Max. RPM: 15,000 RPM. Max. RCF: 16,853 x g. Max. Capacity: 1.5ml x 24 tubes. 200 µl PCR tube adaptor. Display: back light LCD. Over speed protector, door locking, unbalance detector. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
1.16.8 | Anaerobic Work Station: Pressure controll: Automatic. Temp. Range: Ambient + 5°C to 70°C. Temp. Accuracy: +/-0.2°C at 37°C. Temp. Controller: Digital P.I.D. controller, 4Zone PID. Display: Dual LED display, Circulation fan: 5" convection fan. Standard device: Vacuum pump, Glove, Catalyst Pack. Circulator, Drilite Pack, Vacuum gauge, Pressure gauge, Glass shelf. Power supply: 220-240V, 50/60Hz. It should be complimented with: Filled different gas cylinder 6 Nos each. Gas regulator 6 Nos. Tubings & Voltage stabilizer. | Set | P
1.16.9 | Anaerobic jar: Capacity: 2.5 liters Stainless steel/ polypropylene | Piece | √ | √
1.16.10 | Biological safety cabinet-Class II type A1. Main control: Full digital microprocessor control with LCD digital display. Airflow: should be 70% recirculation 30% exhausted. Filters: Should have 0.3 micron 99.9% efficiency and HEPA should be protected by a perforated metal diffuser covering the top of the work zone. Work access opening should be 8 inches (203mm) high. Cabinet should be easily fumigated employing an established procedure. Sound level shall be no more than 67 dbA measured 15 inches (381mm) above the work tray and 12 inches (305mm) in front of viewing window. Illumination: Mounted Fluorescent lighting with UVG lamp. Utility: Two gas valve/service couplings, three or four (230VAC) internal electrical circuits. Cabinet shall be easily converted to a free-standing console model with the addition of the optional base support stand. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √ | √
1.16.11 | Biological safety cabinet-Class II type II B1: Main control: Full digital microprocessor control with LCD digital display. Airflow: should be 30% recirculation 70% exhausted. Filters: Should have 0.3 micron 99.9% efficiency and HEPA should be protected by a perforated metal diffuser covering the top of the work zone. Work access opening should be 8 inches (203mm) high. Cabinet should be easily fumigated employing an established procedure. Sound level shall be no more than 67 dbA measured 15 inches (381mm) above the work tray and 12 inches (305mm) in front of viewing window. Utility: Two gas valve/service couplings, three or four (230VAC) internal electrical circuits. Cabinet exhaust duct connection shall be 12-inch (305mm) diameter. Cabinet shall be easily converted to a free-standing console model with the addition of the optional base support stand. Power: 220V, 50/60Hz. Accessory: Air vilocity measuring device, Storage Pull-Out Trays. Stainless Steel Arm Rest and elbow rests. | Set | P | P
1.16.12 | Blood culture system (It should have minimum 120 cells): An automated modular system with fluorocesence/coloremetric. Technology for culturing, of bacteria and yeast fungi. The system should be able to manage the bottle handling functions such as loading and unloading bottles and regulates the temperature of the system. It should have minimum 120 cells in separate or one drawer module. The system should be configured independently to test blood culture and body fluid cultures. The system should provide basic data management function. It should have an interface to the Laboratory Information System (LIS) to eliminate manual data entry. Reporting can be customized and scheduled to print at a pre-determined time. Should be supplied with latest computer with OS/ (Windows XP), analysis soft ware standard keyboard, mouse,17" LCD monitor and Laser jet printer Should have latest external computer and a Laser jet printer and bar code reader. It should be supplied with the reflectance standard & UPS. supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √
1.16.13 | Blood culture system (It should have minimum 60 cells): An automated modular system with fluorocesence/coloremetric. Technology for culturing, of bacteria and yeast fungi. The system should be able to manage the bottle handling functions such as loading and unloading bottles and regulates the temperature of the system. It should have minimum 60 cells in separate or one drawer module. The system should be configured independently to test blood culture and body fluid cultures. The system should provide basic data management function. It should have an interface to the Laboratory Information System (LIS) to eliminate manual data entry. Reporting can be customized and scheduled to print at a pre-determined time. Should be supplied with latest computer with OS/ (Windows XP), analysis soft ware standard keyboard, mouse,17" LCD monitor and Laser jet printer Should have latest external computer and a Laser jet printer and bar code reader. It should be supplied with the reflectance standard & UPS. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √
1.16.14 | Bunsen burner: 3 mode system (continuous/sensor/footswitch) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √
1.16.15 | Electrophoresis power supply: For gel electrophoresis 100 hour up/down timer. Voltage adjustable from 40 to 120V. Precise digital control. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √
1.16.16 | Microscope, binocular: phase contrast, darkfield & bright field Illumination: with objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top . | Set | √
1.16.17 | Gel photo documentation: Gel imaging system, PC, includes darkroom, camera, cables, image lab software; samples trays with complete accessories, printers and thermal papers. | set | √
1.16.18 | Auto-steam sterilizer (Vertical). With microprocessor digital control and temperature display. Capacity: 80 - 100 liters. Operating temperature range: 105 - 142°C. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √ | √
1.16.19 | Hot plate & magnetic stirrer: Speed Range: 30 - 1500 RPM. Temperature Range: Ambient to 350°C. Temp. Controller: Hydraulic Thermostat. Heater: Interchangeable plate heater. Plate Area: 180 x 180 mm. Top Plate: Ceramic coated top/plate. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √ | √ | √
1.16.20 | Ice maker: Production: 250 liters/day. Storage Capacity: 125 liters. Material (IN): Stainless steel. (OUT): Steel plate with powder coating. Door: Front sliding door. Safety Device: Power breaker, Water level monitor, Inventory level monitor. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
1.16.21 | Incubator CO2: Main control: VFD Displayed Microprocessor Controller. Temperature range: Ambient + 5ºC to 60ºC. Co2 range: 0 to 20%, +/-0.1% at 5%. Co2 detection: Infra-red sensor. Humidity: Up to 96% RH. Programmability: 10 steps / 10 Memory. Display: VFD Digital, 80 character display. Capacity: 165 liters. Safety device: Over/Under Temp. Protector, Gas high/low monitor, Water level monitor, Power breaker. Power: 220V, 50/60Hz. It should be supplied with the following: Filled Co2 Cylinder 6 Nos. Co2 gas regulator 2 Nos. Tubings, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | set | √
1.16.22 | Thermal cycler: Sample capacity: 96 x 0.2 ml PCR tubes or one 96 PCR plate. Block temperature range: 4 - 99°C. Degree range of gradient, maximum: 1-24°C. Temperature control range of gradient: 30 - 99°C. Temperature of lid: 105°C. Block homogenicity: 20-72 ± 0.3°C, 90°C ± 0.4°C. Control accuracy: ± 0.2°C. Heating speed: approx. 6°C/s. Cooling speed: approx. 4.5°C/s. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. Should be supplied with compatible computer and operating systems. | set | √
1.16.23 | Turbidometer: For measuring 0.5 Macfarland turbidity standard.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
1.16.24 | Microplate strip washer: Microplate: 96 wells plate washer. Wash programs: Up to 40. Wash cycles Up to 9. Fluid flow rates; variable, 150 - 1000uL/well/s. Dispense volumes: 50 to 450µl in 1 µl increments. Dispensing accuracy: <5% at 300µl. Display: LCD display with touch screen. Automatic internal four buffer switch. Washing-manifold type: 1x8 manifold. Soaking time programmable in minutes and up to 60 seconds. Washing speed: 3 aspirate and dispense cycles >300 µl/well of 96 wells 8 tube manifold till final aspirate time: ≤130 seconds. Operating temperature: -10°C to +40°C. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Set | √ | √ | √*
1.16.25 | Microplate reader: Detection method: Absorbance.Read method: End point, kinetic, linear scanning. Microplate type: 96 well plate. Reading speed: 8 single wavelengths, 18 double wavelength. Measurement system: 8-channel optical system. Wavelength range: 300 to 700nm. Filters nm: 340, 405, 450, 492, 630. Filter capacity: 6-filter capacity wheel. Light source: Tungsten halogen. Shaking: Low, Medium, High and variable speeds. Reading range: Absorbance range. Abs; 0.000 - 4.000. Dynamic range: 0 to 4.0 OD. Optical Density (OD) accuracy: <1% at 2.5 OD typical. Reading speed: 8 seconds interval. Display: large LCD display with touch screen. External printer compatible to reader. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √ | √ | √*
1.16.26 | Serological analyzer (Infectious diseases): An automated modular system based on the principle of Chemiluminescence. Reaction technology OR enhanced Electrochemiluminescence. Along with hormone and other parameters the analyzer should be able to do. Hepatitis including HBsAg, HCV serology, HIV serology and TORCH panels. Throughput: 60 - 80 test or more/Hour. Must have continuous sample loading. There should be bar codes reader for reagents, controls & calibrators. Analyzer should be capable of using primary sample tubes itself & be able to detect clots in sample. STAT sample positioning should be available. Facility for on board cooling should be available. It should have external printer, UPS. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √
1.16.27 | Shaker: (Rotary mixer), Programmable Speed Range: 0 - 300 RPM. Speed controller: Electronic controller. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √ | √ | √ | √
1.16.28 | Incubator (Shaking): Microprocessor controlled PID type. Temperature range: 35˚C ±2˚C. Capacity: 45-55 litres. Over temperature cut-out, programmable shaking frequency up to 250 rotations per hour. Power supply: AC 220V- 230V, 50HZ. Power cord with round 3 pin top | Set | √
1.16.29 | Microscope, binocular, fluorescent: Illumination: with objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X Complete with power cord. Power supply: 220-240V, 50/60HZ., | Set | √
1.16.30 | FULLY AUTOMATED URINE SEDIMENTANALYZER: Should be able to identify cells, casts and crystals. Other Test Items should include Urobilinogen, Bilirubin, Ketone, Blood, Protein, Nitrate, glucose, Specific Gravity, Ascorbic acid, pH Wavelength: 557, 620, 656nm Principle: Reflectance photometer Strips: 11 items Throughput: 60 to120 tests per hour, Display: Large LCD display Memory: about 1000 results Printer: Built in thermal printer Interface: Standard RS 232 Working environment: Temperature: -10 to +40°C, Humidity: 0 to 95%; with barcode reader to link with LIS. Altitude: > 2000 meter. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
1.16.31 | High Speed Refrigerated Centrifuge (for TB Lab) varible high speed (35,000 to 50,000 RPM 3000g (rotor speed), brushless induction motor powered, centrifuge with micro-processor controls and alphanumeric display. Maximum capacity: 1 to 2 litres. Self-diagnosis on strat up. Programmable speed regulators(independent RPM sensor) amd break. one touch selection for PRM and RCF. Digital display of all the parameters. Automatic rotor detcetion and program loading (approximately up to 30 program memory for storing protocols) Audio and visual alarm for deviation from set parameter and on program complection. Swinging bucket rotor with holder for 15 ml and 50 ml centrifuge tube with adapter. Hard and soft accelaration and breaking essenrial spears and accessories. Power supply: 220 - 230 V. should be supplied with voltage stablizer. | Pcs | √ | √
1.16.32 | A fully automated ELISA processor ,ELISA Plate Processing.The instrument needs to process the key steps of an enzyme immunoassay (EIA), from sample preparation to interpretation of optical density (OD) results. Components include:Pipettor used for diluting and/or dispensing samples and reagents. Incubation slots preset at 5°C, room temperature, or 37°C. 8-channel manifold washer. Microplate reader. Positive identification of samples, reagents, and controls is performed via the internal barcode reader.4 units 96 well microplates (pipetting,incubator,reading).1 unit independent washer &reader( two 8 nozzle manifolds).Liquid levels are checked and presence of possible clots detected. can be connected to Two-way connectivity to the laboratory information System (LIS). Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Pcs | √ | √
1.16.33 | 1000µl (Filtered pipette tips are used for protection against contamination. Application to PCR, sequencing and other techniques involving the use of vaporous, radioactive, biohazardous or corrosive materials.Virgin polyethylene filters.Ensure that all aerosols and liquids are prevented from entering the pipette.Packaged in rack and pre-sterilized.The filter pipette tip is best price and great quality. DNase/RNase free.Autoclavable. | 96pcs/rack
1.16.34 | 50µl (Filtered pipette tips are used for protection against contamination. Application to PCR, sequencing and other techniques involving the use of vaporous, radioactive, biohazardous or corrosive materials.Virgin polyethylene filters.Ensure that all aerosols and liquids are prevented from entering the pipette.Packaged in rack and pre-sterilized.The filter pipette tip is best price and great quality. DNase/RNase free.Autoclavable.) | 96pcs/rack
1.16.35 | 100µl (Filtered pipette tips are used for protection against contamination. Application to PCR, sequencing and other techniques involving the use of vaporous, radioactive, biohazardous or corrosive materials.Virgin polyethylene filters.Ensure that all aerosols and liquids are prevented from entering the pipette.Packaged in rack and pre-sterilized.The filter pipette tip is best price and great quality. DNase/RNase free.Autoclavable.) | 96pcs/rack
1.16.36 | HIV Antigen-antibody combo ELISA, 4th generation invitro diagnostic kit for the detection of p24 antigen and antibody in human serum or plasma.WHO prequalified, Reference Lab evaluated, sensitivity 100%, Specificity >99.4%, Analytical sensitivity: must be less than 60 picograms/ml and /or less than 0.050 IU/ml as per different standards. Recombinant and synthetic peptides for HIV-1 gp41, HIV-1 0 gp41, and HIV-2 gp36, Solid phase:12 x 8 microwell strips, Specimen type: Serum/plasma, Detection type:Combined detection of HIV-1/2 antibodies and antigen, Subtype detection: M N O sub types, Availability of internal controls both specimen and reagent addition controls, Statement of intent: Both for blood donor screening and patient diagnosis, Color coded reagents to monitor procedural steps. Shelf life : After receiving at central depot should be 12 months at 2-8 C. | 96tests/pack
2 | Pathology Department: Biochemistry and Immunology
2.1 | Reagent & Consumables for ELECSYS 2010
2.1.1 | AFP calset for ELECSYS 2010 | 4x1ml | √
2.1.2 | AFP for ELECSYS 2010 | 100test/ pkt | √
2.1.3 | Assay Cup for ELECSYS 2010 | 60x60 | √
2.1.4 | CEA calset for ELECSYS 2010 | 4xlml | √
2.1.5 | Clean cell for ELECSYS 2010 | 6x380 mL | √
2.1.6 | Cortisol calset for ELECSYS 2010 | 4x1ml | √
2.1.7 | Cortisol for ELECSYS 2010 | 100test/pkt | √
2.1.8 | Estradiol calset for ELECSYS 2010 | 4x1ml | √
2.1.9 | Estradiol for ELECSYS 2010 | 100test/ pkt | √
2.1.10 | Ferritin Calset for ELECSYS 2010 | 4x1ml | √
2.1.11 | Ferritin for ELECSYS 2010 | 100test/pkt | √
2.1.12 | Free PSA II calset for ELECSYS 2010 | 4x1ml | √
2.1.13 | Free PSA II for ELECSYS 2010 | 100test/ pkt | √
2.1.14 | FSH calset for ELECSYS 2010 | 4x1ml | √
2.1.15 | FSH for ELECSYS 2010 | 100test/pkt | √
2.1.16 | HCG -1 Beta calset for ELECSYS 2010 | 4x1ml | √
2.1.17 | HCG -1 Beta for ELECSYS 2010 | 100test/ Pkt | √
2.1.18 | LH calset for ELECSYS 2010 | 4x1ml | √
2.1.19 | LH for ELECSYS 2010 | 100test/Pkt | √
2.1.20 | Maintenance kit for ELECSYS 2010 | Pkt | √
2.1.21 | Preci control - U for ELECSYS 2010 | 4x3 ml | √
2.1.22 | Procell for ELECSYS 2010 | 6x380 mL | √
2.1.23 | Progesterone calset for ELECSYS 2010 | 4x1ml | √
2.1.24 | Progesterone for ELECSYS 2010 | 100test/ pkt | √
2.1.25 | Prolactin calset for ELECSYS 2010 | 4x1ml | √
2.1.26 | Prolactin for ELECSYS 2010 | 100test/ pkt | √
2.1.27 | PSA II calset for ELECSYS 2010 | 4x1ml | √
2.1.28 | Serum CEA for ELECSYS 2010 | 100test/ pkt | √
2.1.29 | SYS clean for ELECSYS 2010 | 5x100ml | √
2.1.30 | SYS Wash for ELECSYS 2010 | 500ml | √
2.1.31 | T3 calset for ELECSYS 2010 | 4x1ml | √
2.1.32 | T3 for ELECSYS 2010 | 200test/ pkt | √
2.1.33 | T4 calset for ELECSYS 2010 | 4x1ml | √
2.1.34 | T4 for ELECSYS 2010 | 200test/ pkt | √
2.1.35 | Testosterone calset for ELECSYS 2010 | 4x1ml | √
2.1.36 | Testosterone for ELECSYS 2010 | 100test/ pkt | √
2.1.37 | Total PSA II for ELECSYS 2010 | 100test/ pkt | √
2.1.38 | TSH calset for ELECSYS 2010 | 4x1ml | √
2.1.39 | TSH for ELECSYS 2010 | 200test/ pkt | √
2.1.40 | Assay tips | 30x120/ tips | √
2.1.41 | CA-125 calset | 4x1ml | √
2.1.42 | CA-125 Reagent | 100test/pkt | √
2.1.43 | CA 19.9 Calset | 4x1ml | √
2.1.44 | CA 19.9 Reagent | 100test/pkt | √
2.1.45 | CA 15.3 Calset | 4x1ml | √
2.1.46 | CA 15.3 Reagent | 100test/pkt | √
2.1.47 | Oncology Control Normal/ Abnormal | 4x3 m l | √
2.1.48 | Diluent Universal | 2x16ml | √
2.1.49 | Vit B12 Calset | pkt | √
2.1.50 | Vit B12 Reagent | pkt | √
2.1.51 | Folate Calset | pkt | √
2.1.52 | Folate Reagent | pkt | √
2.1.53 | Preci control Varia, Vit.B12, Ferritin, Folate | 4x3 mL | √
2.2 | Reagent & Consumables for Electrolyte Analyzer (Easylyte)
2.2.1 | Chloride Electrode for Easylyte | pkt | √ | √ | √
2.2.2 | Daily cleaning solution for Easylyte | 6x15ml | √ | √ | √
2.2.3 | Internal filling solution for Easylyte | 125ml | √ | √ | √
2.2.4 | Membrane Assembly for Easylyte | pkt | √ | √ | √
2.2.5 | Na/K/Cl for Easylyte | Pkt | √ | √ | √
2.2.6 | Potassium Electrode for Easylyte | pkt | √ | √ | √
2.2.7 | Reference Electrode for Easylyte | pkt | √ | √ | √
2.2.8 | Sample detector for Easylyte | pkt | √ | √ | √
2.2.9 | Sample probe for Easylyte | pkt | √ | √ | √
2.2.10 | Sodium Electrode for Easylyte | pkt | √ | √ | √
2.2.11 | Solution Valve for Easylyte | pkt | √ | √ | √
2.2.12 | Tubing Kit for Easylyte | pkt | √ | √ | √
2.2.13 | Wash solution for Easylyte | 50ml | √ | √ | √
2.3 | Reagent & Consumables for ELECTROLYTE Analayzer (DIAMOND)
2.3.1 | Chloride Electrode for DIAMOND | pkt | √ | √ | √
2.3.2 | Daily cleaning solution for DIAMOND | 6x15ml | √ | √ | √
2.3.3 | Internal filling solution for DIAMOND | 125ml | √ | √ | √
2.3.4 | Membrane Assembly for DIAMOND | pkt | √ | √ | √
2.3.5 | Na/K/Cl for DIAMOND | Pkt | √ | √ | √
2.3.6 | Potassium Electrode for DIAMOND | pkt | √ | √ | √
2.3.7 | Reference Electrode for DIAMOND | pkt | √ | √ | √
2.3.8 | Sample detector for DIAMOND | pkt | √ | √ | √
2.3.9 | Sample probe for DIAMOND | pkt | √ | √ | √
2.3.10 | Sodium Electrode for DIAMOND | pkt | √ | √ | √
2.3.11 | Solution Valve for DIAMOND | pkt | √ | √ | √
2.3.12 | Tubing Kit for DIAMOND | pkt | √ | √ | √
2.3.13 | Wash solution for DIAMOND | 50ml | √ | √ | √
2.4 | Reagent & Consumables for Immunology Analyzer (ECIQ)
2.4.1 | Maintenance pack for ECiQ | Pkt |  | √
2.4.2 | Thyroid control for ECiQ | Pkt |  | √
2.4.3 | β-HCG II calibrator for ECiQ | Pkt |  | √
2.4.4 | β-HCG II for ECiQ | Pkt |  | √
2.4.5 | TSH calibrator for ECiQ | Pkt |  | √
2.4.6 | TSH for ECiQ | Pkt |  | √
2.4.7 | TT3 calibrator for ECiQ | Pkt |  | √
2.4.8 | TT3 for ECiQ | Pkt |  | √
2.4.9 | TT4 calibrator for ECiQ | Pkt |  | √
2.4.10 | TT4 for ECiQ | Pkt |  | √
2.4.11 | Universal Wash for ECiQ | Pkt |  | √
2.4.12 | Versa Tips for ECiQ | Pkt |  | √
2.4.13 | Micro sample cup for ECIQ | Pkt |  | √
2.4.14 | Signal Reagent for ECiQ | Pkt |  | √
2.5 | Reagent for Photometer
2.5.1 | Albumin for Photometer | Pkt | √ | √ | √ | √
2.5.2 | Alkaline Phosphatase for Photometer | pkt | √ | √ | √ | √
2.5.3 | ALT (SGPT) for Photometer | pkt | √ | √ | √ | √
2.5.4 | AST (SGOT) for Photometer | pkt | √ | √ | √ | √
2.5.5 | Bilirubin Total & Direct for Photometer | pkt | √ | √ | √ | √
2.5.6 | Cholesterol for Photometer | Pkt | √ | √ | √ | √
2.5.7 | Creatinine for Photometer | Pkt | √ | √ | √ | √
2.5.8 | Gamma GTP for Photometer | pkt | √ | √ | √ | √
2.5.9 | Glucose for Photometer | pkt | √ | √ | √ | √
2.5.10 | Total protein for Photometer | Pkt | √ | √ | √ | √
2.5.11 | Triglycerides for Photometer | pkt | √ | √ | √ | √
2.5.12 | Urea for Photometer | pkt | √ | √ | √ | √
2.5.13 | Uric Acid for Photometer | pkt | √ | √ | √ | √
2.5.14 | Thermal paper roll for Photometer | Roll | √ | √ | √ | √
2.5.15 | Halogen lamp 20W 12V for Photometer | No. | √ | √ | √ | √
2.5.16 | Control normal | Pkt | √ | √ | √ | √
2.5.17 | Control high | Pkt | √ | √ | √ | √
2.5.18 | LE (SLE) slide test | 50test/ pkt | √ | √ | √ | √
2.5.19 | ASO Rapid Test Kit | Pkt | √ | √ | √ | √
2.5.20 | CRP Rapid Test Kit | pkt | √ | √ | √ | √
2.5.21 | RF Rapid Test Kit | pkt | √ | √ | √ | √
2.5.22 | Widal Slide Test | 4x5ml | √ | √ | √ | √
2.6 | Vacuttainer Tubes
2.6.1 | Vacuette tube with glycolytic inhibition - Plastic | 100/pkt | √ | √ | √ | √
2.6.2 | Vacuette tube with gel,10ml capacity - Plastic | 100/pkt | √
2.6.3 | Vacuette tube with serum clot activator 10ml capacity - Plastic | 100/pkt | √
2.6.4 | Vacuette tube with serum clot activator 5ml capacity - Plastic | 100/pkt | √ | √ | √ | √
2.7 | OTHERS
2.7.1 | Bar code label thermal paper Size 2"x1" (500 labels/roll) | Roll | √ | √
2.7.2 | Glucose (powder) | 250gm | √
2.7.3 | Urea (powder) | 250gm | √
2.7.4 | Creatinine (powder) | 250gm | √
2.7.5 | Aspartate Transaminase (lyopholized) | 100gm | √
2.7.6 | Alanine Pyruvate Tramsaminase (lyopholized) | 100gm | √
2.7.7 | Alkaline phosphatase (lyopholyzed) | 100gm | √
2.7.8 | Bilirubin (powder) | 250gm | √
2.7.9 | Albumin (Powder) | 250gm | √
2.7.10 | Lactate Dehydrogenase (lyopholyzed) | 100gm | √
2.7.11 | Micro Haematocrite Capillary Tube, Hepainized (75mm length) | 100 tube/pkt | √ | √ | √
2.7.12 | Capillary sealing compound (Haematrocrit tube): Sealing compound provides perfect seal. Capillary Tubes are put into the sealing compound with a twesting motion. | 10trays/ pkt | √ | √ | √ | √
2.7.13 | Almunium Crimp Neck auto Sampler vial cap | 100/pkt | √
2.7.14 | Amber Glass bottle 10ml with rubber cap | 500/pkt | √
2.8 | Equipments
2.8.1 | ELECTROLYTE ANALYZER: Na, K, Cl, Li+, TCO2 CLIA Classification: moderate complexity Sample: Whole blood, Serum, Plasma, Urine Body fluids Sample Size: Approxmately 100 uL Performance: Na+, K+, Cl+, Li+, TCO2 Accuracy: 0.99 Analysis Time: Less than 60 sec (blood) 90 sec (urine) Data Storage: Patient result and QC results Normal, Low and High Calibration: Automatic or On - Demand Output: Printing of result, RS-232 optional, Autosampler tray optional Ambient Conditions: 15 - 32 °C, humidity: <85% Power: Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS. | Set | √
2.8.2 | ELECTROLYTE ANALYZER: Na, K, Cl, Classification: moderate complexity Sample: Whole blood, Serum, Plasma, Urine Body fluids Sample Size: Approximately 100 uL Performance: Na+, K+, Cl Accuracy: 0.99 Analysis Time: Less than 60 sec (blood) 90 sec (urine) Data Storage: Patient result and QC results Normal, Low and High Calibration: Automatic or On - Demand Output: Printing of result, RS-232 optional, Auto sampler tray optional Ambient Conditions: 15 - 32 °C, humidity: <85% Power: 220 - 240V, 50. power cord with round 3 pin top and appropriate UPS. | Set | √ | √ | √
2.8.3 | Chemistry Analyzer System: Discrete, multi-channel, fully automatic and computerized machine, truly random access, open system with onboard washing of reaction cuvettes. Test modes: stat, batch, random access. Assay modes: endpoint, kinetic, prozone check, rate A, B with blank Channels: on board chemistry and ions. Throughput: 400 – 600 tests per hour with ISE Serum indices: lipemic, Icteric, hemolysed Reagent Volume: minimum (R1- 150 – 450 uL, R2 – 20 – 450 uL ) Sample Volume: 2 – 45 uL, step by 0.1 uL Sample types: serum/plasma/urine/CSF & other fluids like peritoneal/pleural/pericardial etc. Reaction temperature: 37°C Stat positions: about 20 positions Cuvettes: Automatic washing cuvettes preferably non-disposable Printer: inbuilt with use of non thermal paper Interface: mono and bi-directional interface essential Cooling: Cool compartment for reagent 4 – 8 °C Method types: performs routine tests, enzymes, drugs assays, immunoturbidimetric assays, Auto dilution and auto rerun during the run Water Supply: On line water supply plant Wavelength: 340, 380, 412, 450, 505, 546, 570, 605, 660, 700, 740, 800 Absorbance: 0 – 3 Abs. Sample probe: Liquid level detection, clot detection and collision protection Reagent probe: Liquid level detection, collision protection and inventory checking Temperature: 15 – 30 °C Humidity: 35 – 80 % Power supplies: 200 – 240 V, 50/60Hz, 1200 VA with UPS Operation System: Windows XP, Interface RS – 232 Accessories: 1. Adaptors should be provided for different sizes of reagent bottles. 2. Start up reagent kits should be provided free of cost during the installation 3. List of spare parts and consumables should be submitted. 4. Warranty 5. Deionizer should be supplied with the analyzer 6. User training must be provided. | Set | √
2.8.4 | Chemistry Analyzer System: Discrete, multi-channel, fully automatic and computerized machines, truly random access, open system with onboard washing of reaction cuvettes. Test modes: stat, batch, random access. Assay modes: endpoint, kinetic, prozone check, rate A, B with blank Channels: on board chemistry and ions. Throughput: 200 – 400 tests per hour with ISE Serum indices: lipemic, Icteric, hemolysed Reagent Volume: minimum ( R1- 150 – 450 uL, R2 – 20 – 450 uL ) Sample Volume: 2 – 45 uL, step by 0.1 uL Sample types: serum/plasma/urine/CSF & other fluids like peritoneal/pleural/pericardial etc. Reaction temperature: 37°C Stat positions: about 20 positions Cuvettes: Automatic washing cuvettes preferably non-disposable Printer: inbuilt with use of non thermal paperInterface: mono and bi-directional interface essential Cooling: Cool compartment for reagent 4 – 8 °C Method types: performs routine tests, enzymes, drugs assays, immunoturbidimetric assays, Auto dilution and auto rerun during the run Water Supply: On line water supply plant Wavelength: 340, 380, 412, 450, 505, 546, 570, 605, 660, 700, 740, 800 Absorbance: 0 – 3 Abs. Sample probe: Liquid level detection, clot detection and collision protection Reagent probe: Liquid level detection, collision protection and inventory checkingTemperature: 15 – 30 °C Humidity: 35 – 80 % Power supplies: 200 – 240 V, 50/60Hz, 1200 VA with UPS Operation System: Windows XP, Interface RS – 232 Accessories: 1. Adaptors should be provided for different sizes of reagent bottles. 2. Start up kits should be provided free of cost during the installation. 3. List of spare parts and consumables should be submitted. 4. Warranty 5. Deionizer should be supplied with the analyzer. Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS |  |  | √
2.8.5 | PHOTOMETER Light Source: Quartz-halogen lamp 12V/20W Wavelength: 340, 405, 492, 510, 546, 578, 630nm Photometric Range: 0.0000 to 3.0000 ABS, Resolution: 0.0001 ABS Flow cell: Stainless steel body, Measuring volume: 32uL, Aspiration Volume programmable: 200 - 2000uL, linearity: 0-1000 units, Optical path: 10mm Analysis method: End point, kinetic, Fixed Time, Absorbance, Turbidimetry, Linear & non-linear standardization Parameter Setting: Method, Wavelength, Temperature, Reagent blank Y/N, sample blank Y/N, Dalay time, Measuring time, Reaction time, Absorbance limit, Aspiration volume, Standards, Linerity check & Unit for results Temperature Control: 25, 30, 37 °C optional, Precision: +/-0.1°C Display: Large LCD display Printer: Built-in thermal printer Interface: RS 232 serial port Operating Environment: 15 - 30°C, Humidity: 20 - 80% Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS. | Set | √ | √ | √ | √
2.8.6 | FULLY AUTOMATED RANDOM ACESS IMMUNO-ANALYZER: Analyzer should be based on the principle of Chemiluminescence reaction Technology OR enhanced Electrochemiluminescence. 1. Analyzer can be Bench top OR Floor base 2. Reagent Carrousel should have at least 10 positions for keeping reagent 3. Programming by loading 4. On board dilution facility 5. Throughput: 20 – 40 tests or more / hour 6. Assay time should be 9 – 18 minutes maximum 7. Reaction time: 37 °C 8. Temperature: 15 – 30 °C 8. Barcode reader for reagent, control and calibrators 9. STAT sampling position 10. On board cooling facility 11. Interface RS – 232 12. External printing facility 13. Onboard inventory management facility 14. Power: 220 - 240V, 50 Hz. power cord with round 3 pin top and appropriate UPS.S. Accessories: 1. Start up kits should be provided free of cost during the installation. 2. List of spare parts and consumables should be submitted. | Set |  | √
2.8.7 | BILIRUBIN METER: Measurement Parameters: 0 - 30 mg/dL (total bilirubin in serum) Light Source: 6V 1.5A tungsten lamp Wavelength: 461, 551 nm Formula: 2 wavelength differential formula Measurement cell: Capillary tube Sample Volume: 50 - 60uL Display: 3-digit LED Power: 220 - 240V, 50 Hz. power cord with round 3 pin top | Set | √ | √ | √
2.8.8 | Barcode Printer with Thermal paper size: 2'' W x 1'' L | Set | √ | √ | √
2.8.9 | Osmometer: Sample volume: 500 µl, Test time: 9 -13 minutes, Single sample capacity, Units: mOsm/kg H2O, Range: 280 - 350 mOsm/kg H2O, Linearity: <2% from a straight line, Temperature: 18 - 35°C, Humidity: 5 – 80 %, Power: 220 - 240V, 50 Hz. power cord with round 3 pin top. | Set | √ | √
2.8.10 | Centrifuge with Swing Out Rotor. Specification: Maximum speed: 8,500 rpm with 841 rotor; 3,900 rpm with 236 rotor. Electronic 30 minute timer and hold mode. Speed Range: 1,000 to 8,500 rpm in 100 rpm increments. Maximum Volume: 400 mL (8 holes x 50 mL ). Sound Level: 52 dBA. Dimensions: 10 3/4" H x 13 1/4" W x 15 3/4" D. | Pcs | √
2.8.11 | Electronic Magnetic Stirrer, Specification: Item Weight: 4.98 pounds. Voltage: 220-240V/60Hz. Scilogex 81111302 MS-Pro, Stainless Steel Plate, Maximum Stirring Speed: 1500 RPM. Overall Height: 3 inches, Overall Length: 11 inches, Overall Width: 6.25 inches, Plate diameter: 5.3", Digital Temperature Range: 0-340 Degrees Celsius (stirrer without heating system is preferred. If heating is present temperature adjustment is required) # different size of magnetic bar should be provided with magnetic stirrer. # any model magnetic stirrer which fulfill the above specification will be accepted. | Pcs | √
2.8.12 | Floor-Model Freez Dryer. Specification: Dimensions (approx): 32'W x 29.25"D x 33.5"H (outside). Electrical: 208 Volts, 60 Hz, 15A. Capacity: 14 Liters for 24 hours with a maximum low condenser temperature at -85C. This freezer dryer also has digital temperature and vacuum display. The FTS Dura-Dryer includes the TD2COT5000 Dura-Top bulk tray dryer. The tray dryer has non adjustable shelves. Feautures: Corrosion resistant construction results in minimal maintenance, Patended coil design prevents moisture carry-over to the vacuum pump. Clear acrylic door gives condenser. Visibility and automatic defrost reduces turnaround time. | Pcs | √
2.8.13 | Glass Vacuum Flask Kit - Heavy Duty. Capacity: 2 Liter. Filter Funnel with glass Frit base 500 ml, SS Funnel Clamp, SS Funnel Cover, Vac House. | Set | √
2.8.14 | Zippette Digital Bottle Dispenser: Specification: Capacity: 0-30ml. Delivery accruracy ± 0.3%, Precission better than 0.1% CV, adjustable volume by turning the top bottom, Chemical resistent can be used for wide range of equipments, Large digital volume display for easy and precise volume setting, display resolution of 100 ul, Autoclavable at 121°C2 bar. Simple user calibration via front panel buttons for GLP/ISO compliance. Safety right-angled spout design. Borosicate glass barrel protected with a transparent olyprophylene sleeve. Bubble free dispensing. | Set | √
2.8.15 | 13mm Flip off Vial Manual Crimper Hand Tool. Specification: Size arrange: 8mm - 32 mm. Gross weight: 1Kg. Color: Silver. Material: 100% stainless steel. Production Capacity: 20 - 30 bottle/min. | Pc | √
3 | Pathology Department: Heamatology
3.1 | Cacl2 (0.02mol/L) | 10x15ml | √ | √ |  |  | Machine specific (closed system)
3.2 | Cuvette Rings 6 x 10 x 32 | 1920 tests/pkt | √ | √
3.3 | Washing Solution | 5 x 15 ml | √ | √
3.4 | Sample Cups 4 ml | 2 x 250 pcs | √ | √
3.5 | Sample Cups 2 ml | 1 x 500 pcs | √ | √
3.6 | Cleaner | 5 x 15 ml | √ | √
3.7 | Reagent Vials 5 ml with Caps | 50 pcs/pkt | √ | √
3.8 | Teflon Stirrer Bar | 10 pcs/pkt | √ | √
3.9 | Hemostat aPTT-EL | 6 x 4 ml | √ | √
4 | Hemostat D-Dimer | 2 x 1 ml | √ | √
4.1 | Hemostat Fibrinogen | 5 x 2ml | √ | √
4.2 | Hemostat Thromboplastin-SI | 6 x 10 ml | √ | √
4.3 | Hemostat Control Plasma Normal | 6 x 1 ml | √ | √
4.4 | Hemostat Control Plasma Abnormal | 6 x 1 ml | √ | √
4.5 | Hemostat Calibrator | 4 x 1 ml | √ | √
3.1 | COAGULATION ANALYZER (Hemostat and Costat)
3.1.1 | Cacl2 (0.02mol/L) | 10x15ml |  |  | √ |  | Open system
3.1.2 | Cuvettes and stirrer for -HEMOSTAT model | 100/Pkt. |  |  | √
3.1.3 | Cuvettes and stirrer for- Costat model | 100/Pkt. |  |  | √
3.1.4 | PT determination reagent | 10x2ml |  |  | √
3.1.5 | APTT determination reagent | 10x4ml |  |  | √
3.1.6 | Normal control plasma for coagulation study | 10x1ml |  |  | √
3.1.7 | Abnormal control plasma for coagulation study | 10x1ml |  |  | √
3.2 | MS95 & MS95s ANALYZER 5-PART DIFF
3.2.1 | Acti - diff for MS95s | 1.2ltr |  |  | √ |  | Machine specific (closed system)
3.2.2 | Acti - diff for MS95 | 1.2ltr |  |  | √
3.2.3 | Calibrator for MS95s (5-part) | 3ml/pkt |  |  | √
3.2.4 | Clot Filter | 2 pcs/Pkt |  |  | √
3.2.5 | Eo-Diff | 5ltr |  |  | √
3.2.6 | Hemoref | 250ml |  |  | √
3.2.7 | Isoflux | 20ltr |  |  | √
3.2.8 | 102 Control Diff (5-part) | 3ml each/Pkt. |  |  | √
3.2.9 | Transflux | 5ltr |  |  | √
3.2.10 | Cleaning Solution for MS Cell Counter | 250 ml |  |  | √
3.3 | MS4 ANALYZER 3-PART DIFF
3.3.1 | Actidiff for MS4Se Blood cell counter | 1.2ltr |  |  | √ | √ | Machine specific (closed system)
3.3.2 | Actidiff for MS4e Blood cell counter | 1.2 ltr |  |  | √ | √
3.3.3 | Isoflux for MS4 Blood cell counter | 20ltr |  |  | √ | √
3.3.4 | Transflux for MS4 Blood cell counter | 5ltr |  |  | √ | √
3.3.5 | 102 Control Diff (3-part) | 4ml each for low, normal and high /Pkt |  |  | √ | √
3.3.6 | Cleaning Solution for MS Cell Counter | 250 ml |  |  | √ | √
3.3.7 | Thermal Paper | 1roll/pkt |  |  | √ | √
3.3.8 | Clot Filter | 2pcs/Pkt |  |  | √ | √
3.4 | SYSMEX Kx 21 ANALYZER
3.4.1 | Cell Clean | 50ml |  |  | √ | √ | Machine specific (closed system)
3.4.2 | Cell Pack | 20ltr |  |  | √ | √
3.4.3 | Eight Check, 3WP, Low, Normal and High Control | 2ml each/Pkt |  |  | √ | √
3.4.4 | Thermal Printer paper roll | 1roll/pkt |  |  | √ | √
3.4.5 | Stromatolyzer - WH | 500ml |  |  | √ | √
3.5 | AUTOMATED ESR ANALYZER (VESMATIC 20)
3.5.1 | Normal ESR Control | 1x9ml | √
3.5.2 | Abnormal ESR control | 1x9ml | √
3.5.3 | ESR vacutte tube for Vesmatic 20, DIESSE | 2x100 | √
3.5.4 | Paper roll for vesmatic 20 | 1roll/Pkt | √
3.5.5 | Hematology general consumables
3.5.6 | Complete blood count QC material (three level-high, low and normal) | 2ml each/Pkt | √
3.5.7 | Rapid Malaria Test kit (ICT) | 50test/ pkt | √ | √ | √ | √
3.5.8 | G6PDH kit (Qualitative) | 20 test/ kit | √
3.5.9 | Reticulocyte stain solution (ready made solution) | 100ml | √ | √ | √ | √
3.5.10 | Safranin solution | 100ml | √
3.5.11 | Perl's stain solution (ready made solution) | 100ml | √
3.5.12 | New methylene Blue powder | 25gm | √ | √ | √ | √
3.5.13 | Sodium phosphate dibasic | 500g | √ | √ | √ | √
3.5.14 | potassium dihydrogen phosphate | 500g | √ | √ | √ | √
3.5.15 | Drabkin solution for Hb estimation | 5ltr | √ | √ | √ | √
3.5.16 | EDTA Dipotassium salt | 50gm | √ | √
3.5.17 | Giemsa stain | 100ml | √ | √ | √ | √
3.5.18 | May-grunwald-giemsa stain | 100ml | √ | √ | √ | √
3.5.19 | Leishman stain | 100ml | √ | √ | √ | √
3.5.20 | Sodium Metabisulphate | 50gm | √
3.5.21 | Trisodium citrate (3.8%) | 500g | √ | √ | √ | √
3.5.22 | Turks solution | 100ml | √ | √ | √
3.5.23 | Potassium ferrocyanide | 500g | √
3.5.24 | DLC counter | 1/pkt | √ | √ | √ | √
3.5.25 | Glass beads | 1kg/pkt | √ | √ | √ | √
3.5.26 | Heamocytometer with improved Neubauer rulings Counting Chamber | 1/pkt | √ | √ | √ | √
3.5.27 | Haemocytometer coverglass | 5/pkt | √ | √ | √ | √
3.5.28 | Sahli hemoglobinometer | 1/pkt | √ |  |  | √
3.5.29 | ESR tube with stand and scale- Westergren tube | 5/Pkt. | √ | √ | √ | √
3.5.30 | ESR vials with cap | 100 /Pkt. | √ | √ | √ | √
3.5.31 | Template for bleeding time (adult)-Depth: 1.0 mm. Length: 5.0 mm. Fully automated incision device for bleeding time determinations. Single use. Latex free. Sterile unless package is damaged or opened. Lot No. 2013-03 | 50/Pkt | √ | √ | √ | √
3.5.32 | Template for bleeding time (infant)-Depth: 0.5 mm. Length: 3-4.0 mm. Fully automated incision device for bleeding time determinations. Single use. Latex free. Sterile unless package is damaged or opened. | 50/Pkt | √ | √ | √ | √
3.5.33 | vacutainer tube coated with K2 EDTA for adults | 100 tube/Pkt | √ | √ | √ | √
3.5.34 | vacutainer tube coated with K2 EDTA for infants | 100 tube/Pkt | √ | √ | √ | √
3.5.35 | Buffered sodium citrate 0.109 M (3.2%) plastic vacutainer tube for adult | 100 tube/Pkt | √ | √ | √
3.5.36 | Buffered sodium citrate 0.109 M (3.2%) plastic vacutainer tube for infant | 101 tube/Pkt | √ | √ | √
3.5.37 | vacutainer tube (sodium/lithium heparin, no gel) | 102 tube/Pkt | √
3.5.38 | Timmer | 1pcs/pkt | √ | √ | √ | √
3.5.39 | Hb color scale | 50 test/pkt |  |  |  | √
3.6 | Flocytometry
3.6.1 | CD7-FITC | 1 vial each | √
3.6.2 | CD5-FITC | 1 vial each | √
3.6.3 | CD19-PE | 1 vial each | √
3.6.4 | CD10-FITC | 1 vial each | √
3.6.5 | CD33-FITC | 1 vial each | √
3.6.6 | CD117-PE | 1 vial each | √
3.6.7 | CD34-PE | 1 vial each | √
3.6.8 | Thiazole Orange | 1 vial each | √
3.6.9 | BD FACSFlow solution | 20 liters | √
3.6.10 | BD FACS Lysing solution | 50 ml | √
3.6.11 | BD FACS Calibrite 3 Beads | 4 vials/ pkt | √
3.6.12 | BD Pharmlyse solution (10x) | 100 ml | √
3.6.13 | BD Perm II (10x) | 25 ml | √
3.6.14 | Hu CD 19PerCP | 100 test | √
3.6.15 | cCD3PE | 1 vial each | √
3.6.16 | cCD79a PE | 1 vial each | √
3.6.17 | CD45PCP | 1 vial each | √
3.6.18 | MsIg GI FITC | 1 vial each | √
3.6.19 | MSIg GI PE | 1 vial each | √
3.6.20 | CD13 PE | 1 vial each | √
3.6.21 | TDT- FITC | 1 vial each | √
3.6.22 | Kapa FITC | 1 vial each | √
3.6.23 | Lambda PE | 1 vial each | √
3.7 | EQUIPMENT
3.7.1 | FULLY AUTOMATED COAGULATION ANALYZER: Parameters: Prothrombin Time (PT, PTR, INR), Activated Partial Thromboplastin Time (APTT), Fibrinogen (FIB), Thrombin Time (TT), Protein C, Protein S, ATIII; Coagulation Factors: Extrinsic Factors ( II, V, VII, X), Intrinsic Factors (VIII, IX, XI, XII); Detection Range : 50 - 450 mg/dL; Advanced Photo electronic-Double magnetic Testing; Built - in printer, Open system for reagents, 200 results memory capacity, No interference from abnormal samples (optical dense samples such as lipemic or icteric), Repeatability (CV%) <3%; Sufficient incubation stations for reagents and samples, Testing time 3 to 999 seconds, Inbuilt time for incubation. Should have the option to connect to LIS/HIS. Power supply: AC 220V- 230V, 50HZ; RS232 LIS/HIS connectivity port is mandatory. | Set | √ | √ | √
3.7.2 | FULLY AUTOMATED ESR ANALYZER: Features: Measures ESR according to Westergren Method. Works with EDTA blood. Built-in barcode reader. Sample volume: 1.3 ml whole blood. Analysis time: 30 minutes. The equipment should be complaince with infection control. | Set | √ | √
3.7.3 | ELECTRONIC DIFFERENTIAL COUNTER: – Technical data. Light weight and compact for easy transportation. This counter should have button to count all the important leukocytes, erythroid and megakaryocytic series. There should be short sound coming after every stroke of key and when the totalizer records to 100, 200, 300, etc. a long sound buzzs. The percentage of each type of cell is should be read directly from the window above each key. Function of CLEAR should be present. Function of CORRECTION should be present – To correct the possible counting error. Function of LOCK should be present – To keep the counting results from being touched carelessly. | Set | √ | √
3.7.4 | FULLY AUTOMATED HAEMATOLOGY ANALYZER - 5 part-diff, 21 parameter minimally: Differential counts:WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT, NEU%, LYMPH%, MONO%, EO%, BASO%, NRBC%, NEUT#, LYMPH#, MONO#, EO#, BASO#, NRBC#, RDW-SD, RDW-CV, MPV, P-LCR*, PDW*, PCT*, micro RBX%, macro RBC%, Mode Median, micro THR%, Macro THR%, Able to measure micro & macro platelets and RBC; Sample Volume: ≤20µl Sample Memory: 100-240 results with histogram The analyzer should have In-built printer, Barcode reader, External keyboard and QC soft ware’s for 3 control Throughput of the analyzer should be minimally 60-80 samples/hour. The Shelf life of reagents should be 12 -18 months. Analyzer should have auto Standby mode . Power Supply AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS must be provided. LIS/HIS connection should be there. Other accessories Program upgradable by CD/USB and results backup on CD/USB. Digital blood banks with programmable min/max ranges on all banks. Operating system windows based for user friendly. Cap piercing facilities for infants blood samples, Power supply Accuracy requirements for 5-part and 3-part Automated Hematology Analyzer WBC: Mean ±0.2x103/µl RBC: Mean ±0.04x106/µl PLT: Mean ±10 x103/µl Carry over requirements for 5-part and 3-part Automated Hematology Analyzer WBC: < 3% RBC: < 1.6% HGB: < 1.8% PLT: < 5% Precision requirements for 5-part and 3-part Automated Hematology Analyzer WBC: ≤3.5% RBC: ≤2.0% PLT: ≤6.0% Linearity requirements for 5-part and 3-part Automated Hematology Analyzer WBC: 0-100 x103/µl RBC: 0-7x106/µl HGB: 0.1-25.0 g/dl HCT: 10-67 % PLT: 0-999x103/µl | Set | √ | √ | √
3.7.5 | AUTOMATED HAEMATOLOGY ANALYZER: 3-part diff,18 parameter minimally. Differential counts:WBC, Lymp%, Lymp#, Mon%, Mon#, NEU%, NEU#, Gr#, RBC, Hb, MCV, Hct, MCH, MCHC, PLT, RDW-CV, RDW-SD, PDW, MPV, uRBC%, MacroRBC%, Pct, P-LCR, PDW, Mode, Median, uTHR%, Macro THR%. Sample Volume: ≤20µl Sample Memory: 50-200 results with histogram The analyzer should have In-built printer, Barcode reader, and External keyboard and QC soft wares for 3 controls Throughput of the analyzer should be minimally 40-60 samples/hour. The Shelf life of reagents should be 12 -18 months. Analyzer should have control Standby mode. Power Supply AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS must be provided. LIS/HIS connection should be there. Accuracy requirements for 5-part and 3-part Automated Hematology Analyzer WBC: Mean ±0.2x103/µl RBC: Mean ±0.04x106/µl PLT: Mean ±10 x103/µl Carry over requirements for 5-part and 3-part Automated Hematology Analyzer WBC: < 3% RBC: < 1.6% HGB: < 1.8% PLT: < 5% Precision requirements for 5-part and 3-part Automated Hematology Analyzer WBC: ≤3.5% RBC: ≤2.0% PLT: ≤6.0% Linearity requirements for 5-part and 3-part Automated Hematology Analyzer WBC: 0-100 x103/µl RBC: 0-7x106/µl HGB: 0.1-25.0 g/dl HCT: 10-67 % PLT: 0-999x103/µl |  |  |  | √ | √
3.7.6 | SWING ROLL MIXER: Speed - 0 to 60, Capacity - 50 to 100 tubes. The roller should have adequate vertical bend to ensure homogenous mixing. |  | √ | √ | √ | √
3.7.7 | FLOW CYTOMETRY for diagnosis of Haematological malignancies: At least 5-colored fluorescence, Air-cooled 488nm 15mw laser 10µ x 90µ light beam. Detector Range: 300nm-800nm. Flow-Cell Size: 250µ x 250µ BioSense. Max analysis rate: >3300 cells/sec. Min Cell Size, um: <0.5. Max analysis rate: >3300 cells/sec. Color dot density; contour, isometric, tomogram, overlay, legend and projection plots. PRISM displays; color precedence or color-blend gating; automatic Boolean logic gate creation; 32 gates per file, real-time statistics. Intel Pentium processor and 4 network transporters for digital signal processing. Signal-intensity display, biohazard-safe triple-fluid (sheath/cleanse/waste) system with automated functions and variable flow rate; auto setup; SYSTEM II software with acquisition, list mode, multigraph, reports and utilities. Features Fastest and Cell Tracker. |  | √
3.7.8 | HIGH PERFORMANCE LIQUID CHROMATOGRAPHY (HPLC) for hemoglobin typing- qualitative & quantitative: Injection range: Programmable from 0.1-1.0 μL at 0.1 μL increments, 1 μL to 2.5 mL at 1 μL increments. Precision: Typically < 0.5% RSD of peak areas from 3 μL, linearity > 0.999. Sample syringe size: 50, 100, 250, 500, 1000, 2500 μL. Carryover: < 0.02%. Flush volume: 0-2500 μL. External contacts: Two, start and stop time. Method storage (battery backed): Up to 20 methods including time programming. Stored methods: Editing possible during runs. Injection cycle time 2-3 samples per minute. Valve materials: In contact with Rheodyne® 7725 for analytical – stainless steel, Vespel®, ceramic. PEEK with sample with Rheodyne 9725 for bio applications – PEEK, Tefzel®, ceramic. Display: 7-line by 40-character backlit liquid crystal display. Integrated: For derivatization and dilution options serial communications – standard BCD. Communications interface – standard 5 relay outputs; 2 timed events, 2 inject, 1 start Ambient conditions: 10-30 °C, relative humidity 20-80%. Peltier sample cooling: 4-60 °C, ±1 °C. Power supply: AC220V, 50/60Hz. Should be supplied with UPS |  | √
3.7.9 | Barcode reader (or barcode scanner): The barcode reader should be able to read all the barcode and transfer the laboratory datas( values, histogram and scatter gram) accurately and precisely. | Pcs | √ | √ | √ | √
4 | Pathology Department: Histo/cyto Consumable
4.1 | Cover glass 22mmx50mm | 20 pkts of 10gm each | √ | √ | √
4.2 | DPX Mountant | 250 ml | √ | √ | √
4.3 | Ethanol (99 - 100%) 500 ml | 500ml | √ | √ | √
4.4 | Haematoxylin Powder | 25 gm | √ | √
4.5 | Slide staining rack, SS. (wxdxh) 100mmx100mmx25mm, use with staining trough (accepts 76x26mm microscope slides) | Piece | √ | √ | √
4.6 | Slide staining trough (wxdxh) 112x112x50mm | Piece | √ | √ | √
4.7 | Slide tray of 20 slides ( Aluminium ) | Piece | √ | √ | √
4.8 | Sodium bi carbonate | 500gm | √ | √ | √
4.9 | Xylene extra pure. | 500ml | √ | √ | √
4.1 | CYTO CONSUMMABLES
4.1.1 | Disposable Filter Paper for Cytocentrifuge (1ml) | 200/ pack | √ | √
4.1.2 | Specimen Chamber (1ml) for Cytocentrifuge | Piece | √ | √
4.1.3 | Specimen Chamber cap (1ml) for Cytocentrifuge | 200/Pkt | √ | √
4.1.4 | Specimen Chamber holder (1ml) for Cytocentrifuge | 200/Pkt | √ | √
4.1.5 | EA - 36 or 50 (Papanicolaous solution) polychromatic solution | 500ml | √ | √ | √
4.1.6 | Harris Haematoxylin Papanicolaus solution 1a | 500ml | √ | √ | √
4.1.7 | Orange Green-6, Papanicolaous solution, 2b Orange II Solution (contains methanol) | 500ml | √ | √ | √
4.1.8 | Polyethylene glycol 200 | 500 ml | √ | √ | √
4.1.9 | Dip quick I | 500gm | √
4.1.10 | Dip quick II | 500gm | √
4.2 | HISTO CONSUMMABLES
4.2.1 | Alcian blue powder | gm | √ | √
4.2.2 | Aluminium Chloride Crystal | 500gm | √ | √
4.2.3 | Aluminium hydroxide | 500gm | √ | √
4.2.4 | Aluminium Potassium Sulphate | 500gm | √ | √
4.2.5 | Ammonical silver powder | gm | √ | √
4.2.6 | Aniline blue powder | gm | √ | √
4.2.7 | Base mold SS, (reuseable), Size - 25Wx32Dx6Hmm | Piece | √ | √
4.2.8 | Base mold SS, (reuseable), Size - 25Wx38Dx6Hmm | Piece | √ | √
4.2.9 | Base mold SS, (reuseable), Size - 25Wx32Dx13Hmm | Piece | √ | √
4.2.10 | Borax powder (Sodium tetraborate decahydrate) | gm | √ | √
4.2.11 | Carmine Powder | gm | √ | √
4.2.12 | Chromic acid powder | gm | √ | √
4.2.13 | Citric acid | 500 gm | √ | √
4.2.14 | Concentrated Nitric acid | ml | √ | √
4.2.15 | Congo red powder | gm | √ | √
4.2.16 | Chloral hydrate | 500gm | √ | √
4.2.17 | Di Sodium Hydrogen Orthophosphate | 500gm | √ | √
4.2.18 | Embedding Ring | Piece | √ | √
4.2.19 | Eosin Yellow Powder (water soluble) | gm | √ | √
4.2.20 | Ferric Chloride crystal | gm | √ | √
4.2.21 | Ferrocynide powder | gm | √ | √
4.2.22 | Formaldehyde Solution 37-40% | 500ml | √ | √ | √
4.2.23 | Frozen Section Medium ( Neg - 50) | 125 ml | √ | √
4.2.24 | Gold Chloride powder | 1 gm | √ | √
4.2.25 | Gold Nitrate powder | gm | √ | √
4.2.26 | Iron alum powder | gm | √ | √
4.2.27 | Light green powder | gm | √ | √
4.2.28 | Liquid Paraffin | 500 ml | √ | √
4.2.29 | Lithium Carbonate powder | gm | √ | √
4.2.30 | Magnesium Sulphate | 500 gm | √ | √
4.2.31 | Metanil yellow powder | 100gm | √ | √
4.2.32 | Methenamine powder | gm | √ | √
4.2.33 | Methyl violet | ml | √ | √
4.2.34 | Microtome Blade (Low profile disposible, S35) | 50/Pkt | √ | √
4.2.35 | Naphthol powder | gm | √ | √
4.2.36 | N-N-dimethyl formamide solution | ml | √ | √
4.2.37 | N-N-dimethyl, m-phenylenediamine powder | gm | √ | √
4.2.38 | N-N-dimethyl, p-phenylenediamine powder | gm | √ | √
4.2.39 | Nuclear fast red | gm | √ | √
4.2.40 | Oxalic acid powder | gm | √ | √
4.2.41 | Paraffin wax with Ceresin (Congealing point 60 degree C) | 500gm | √ | √
4.2.42 | Pararosaniline (dye) | gm | √ | √
4.2.43 | Periodic acid powder | gm | √ | √
4.2.44 | Perspex box (museum Jar) | Piece | √ | √
4.2.45 | Phloxine powder | 25 gm | √ | √
4.2.46 | Phosphomolybdic acid powder | gm | √ | √
4.2.47 | Phospho Tungstic acid powder | gm | √ | √
4.2.48 | Picric acid powder | gm | √ | √
4.2.49 | Schiff`s reagent | 125 ml | √ | √
4.2.50 | Silver Nitrate | 100gm | √ | √
4.2.51 | Sodium barbital powder | gm | √ | √
4.2.52 | Sodium Di-hydrogen orthophosphate | 500gm | √ | √
4.2.53 | Sodium Iodate | gm | √ | √
4.2.54 | Sodium Thiosulphate | 500gm | √ | √
4.2.55 | Sodium metabisulphite powder | gm | √ | √
4.2.56 | Sudan IV powder (Black) | gm | √ | √
4.2.57 | Tissue cassette ( White colour only ) | 100/Pkt | √ | √
4.2.58 | Tissue grossing knife (Razor) No. 260 Type( L) | Piece | √ | √
4.2.59 | Tri-Sodium Citrate (Na3C6H507.2H2O) | 500gm | √ | √
4.2.60 | Toludine blue powder | gm | √ | √
4.2.61 | Uranium (Uranyl) Nitrate | gm | √ | √
4.2.62 | Whatman, Lens cleaning tissue(only for histopathology),(8X11.5cm). | 20 wallets of 50 sheets | √ | √
4.2.63 | Wide mouth specimen container with cap ( Medical grade PC/PP 125 ml ) | 12 Nos/pack | √ | √
4.2.64 | Wide mouth specimen container wit cap ( Medical grade PC/PP 250 ml ) | 12 Nos/pack | √ | √
4.2.65 | Wide mouth specimen container with cap ( Medical grade PC/PP 500 ml ) | 6 Nos/pack | √ | √
4.2.66 | Wide mouth specimen container with cap ( Medical grade PC/PP 1000ml ) | 6 Nos/pack | √ | √
4.2.67 | Wide mouth specimen container with cap ( Medical grade PC/PP 2500 ml ) | Piece | √ | √
4.2.68 | Wide mouth specimen container with cap ( Medical grade PC/PP 5000 ml ) | Piece | √ | √
4.3 | IMMUNOCHEMISTRY CONSUMMABLES FOR MANUAL METHOD
4.3.1 | Anti - Alpha-Fetoprotein ( AFP ) C3 , Ready to use | 6ml | √
4.3.2 | Anti - bcl-2 (Oncoprot), Ready to use | 6ml | √
4.3.3 | Anti - Beta HCG ( M94138 ), Ready to use | 6ml | √
4.3.4 | Anti - Calretinin (Polyconal), Ready to use | 6ml | √
4.3.5 | Anti - CD3 (T-Cell ) PS1, Ready to use | 6ml | √
4.3.6 | Anti - CD4(T4),Ready to use | 6ml | √
4.3.7 | Anti - CD5(4C7),Ready to use | 6ml | √
4.3.8 | Anti - CD7(D7-272),Ready to use | 6ml | √
4.3.9 | Anti - CD10(56C6),Ready to use | 6ml | √
4.3.10 | Anti - CD15 ( Tu9 ), Ready to use | 6ml | √
4.3.11 | Anti - CD19(EP169),Ready to use | 6ml | √
4.3.12 | Anti - CD20 (B cell ) L26, Ready to use | 6ml | √
4.3.13 | Anti - CD 30 (ki -1Ag) Ber-H2,Ready to use | 6ml | √
4.3.14 | Anti - CD34, Ready to use | 6ml | √
4.3.15 | Anti - CD45, cocktail LCA88, Ready to use | 6ml | √
4.3.16 | Anti - CD99(MIC2 antigen)-HO36.1.1, Ready to use. | 6ml | √
4.3.17 | Anti - CD117 ( T595 ), Ready to use | 6ml | √
4.3.18 | Anti - CDX-2 ( CDX2-88), Ready to use | 6ml | √
4.3.19 | Anti - CEA (CEA 88), Ready to use | 6ml | √
4.3.20 | Anti - Chromogranin A (LK2H10), Ready to use | 6ml | √
4.3.21 | Anti - Cytokrratin 7, (OV-TL1230), Ready to use | 6ml | √
4.3.22 | Anti - CK-19(RCK108) Ready to use | 6ml | √
4.3.23 | Anti - Cytokeratin 20 ( IT-Ks 20.8), Ready to use | 6ml | √
4.3.24 | Anti - CytoKeratin, Pan (AE1+AE3), Ready to use | 6ml | √
4.3.25 | Anti - Cytokeratin, High MW(34βE12) Ready to use | 6ml | √
4.3.26 | Anti - Desmin(D33),Ready to use | 6ml | √
4.3.27 | Anti - EMA (E29), Ready to use | 6ml | √
4.3.28 | Anti - GFAP(Polyclonal),Ready to use | 6ml | √
4.3.29 | Anti - Her 2/ Neu ( CB11 ), Ready to use | 6ml | √
4.3.30 | Anti - Kappa Light Chain ( K88 ), Ready to use | 6ml | √
4.3.31 | Anti - Lambda Light Chain ( HP6054 ), Ready to use | 6ml | √
4.3.32 | Anti - Melan-A(MART-1) clone A103 Ready to use | 6ml | √
4.3.33 | Anti - Melanoma (HMB45), Ready to use | 6ml | √
4.3.34 | Anti - MIB-1/Ki-67, Ready to use | 6ml | √
4.3.35 | Anti - Myogenin, Ready to use. | 6ml | √
4.3.36 | Anti - Myoglobin ( MG-1 ), Ready to use | 6ml | √
4.3.37 | Anti - Neurofilament ( NE-14, Ready to use | 6ml | √
4.3.38 | Anti - OCT4(EP143),Ready to use | 6ml | √
4.3.39 | Anti - Pax5(ZP007) | 6ml | √
4.3.40 | Anti - PLAP(PL8-F6),Ready to use | 6ml | √
4.3.41 | Anti - PSA (Polyconal) ErPr8, Ready to use | 6ml | √
4.3.42 | Anti - RCC(26),Ready to use | 6ml | √
4.3.43 | Anti - S-100 Protein (15E2E2), Ready to use | 6ml | √
4.3.44 | Anti - SMA (Smooth Muscle Actin) 1A4, Ready to use | 6ml | √
4.3.45 | Anti - TTF-1 (BGX-397A), Ready to use | 6ml | √
4.3.46 | Anti - Vimentin (V9 ), Ready to use | 6ml | √
4.3.47 | Antigen Retrival Citraplus HK081-5K | 250ml | √
4.3.48 | EZ-AR™ Common HK545-XOK | 1000ml | √
4.3.49 | EZ-AR™3 HK543-YOK | 500ml | √
4.3.50 | EZ-AR™4( RTU ) | 1000ml | √
4.3.51 | InSite™ ER Antibody ( ER 1D5) , Ready to use | 6ml | √
4.3.52 | InSite™ PR Antibody (PR88), Ready to use | 6ml | √
4.3.53 | Liquid Blocker super Pap pen | Piece | √
4.3.54 | Polymer HRP (DAB) Code no.QD400-60k sets | set | √
4.3.55 | Protease XXIV 3-Pack | 3x5ml | √
4.3.56 | Special positively charged glass with frosted end | 72/Pkt | √
4.3.57 | Tween® 20 (Polysorbate) for synthesis | 500ml | √
5 | Cytology
5.1 | Microscope: 6V 30W halogen bulb. Focusing •Stage height movement by roller guide (rack & pinion). •Stroke per rotation: 36.8 mm. •Full stroke range: 25 mm. •Upper limit stopped by simplified pre-focusing dial. •Tension adjustment on coarse focus adjustment knob Revolving nosepiece Fixed quadruple nosepiece with inward tilt Binocular observation tube Field number 20 Tube inclination 30° Interpupillary distance adjustment range 48-75 mm Stage Size 188 (W) × 134 (D) mm Movement range 76mm X-direction × 50mm Y-direction Specimen holder Double slide holder Rubber grip Equipped as standard Condenser Type Abbe condenser, with built-in daylight filter N.A. 1.25 with oil immersion Aperture iris diaphragm .Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. Right- hand stage control. The microscope should be installed and demonstrated for fuctionallity test. Physically damaged microscope will not be accepted. After sales and services should be provided. | Set | √ | √ | √
5.2 | Fume Hood Structure: All steel (1.0mm cold-rolling steel or 1.0mm stainless steel) covered with electrostatic spraying after folding, welding, polishing, acid cleaning, phosphorization, and chemical resistance. Inner liner panel: 6mm compact panel, acid and alkali resistance, temperature: resistance, anti-moisture. Guide plate: reasonable and knock down guide structure. Concentrated and effective discharge air, letting off different poisonous gas. Operation Table board: lab-grade phenolic resins worktop, or epoxy resin worktop, or ceramic worktop and stainless steel as worktop. Window: 6mm thickness steel safe glass, balancing form elevation and subsidence, it can be stopped at any position by pulling. Lighting equipment: 30W antisepsis daylight lamp, lamp house is insulated with gas. Lightness on table not lowers than 400LUX. Operation panel: lighting switch, ON/OFF switch, two sets of 220V/10A versatile electric socket. Ventilation equipment: selecting rust-resistance, low wasting, low noise and famous fan and special wind tube. Noise can be reduced at about 55 DB. Accessories: vertical and single test cock: made by yellow brass, the surface is treated by EPOXY baking paint. Cup slot: acid and alkali resistance, PP water cup.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. It should be installed and demonstrated for functionallity test. After sales and services should be provided. | Set | √ | √
5.3 | Cyto Centrifuge Application: Bench top centrifuge for cytological specimens that transfers cells from suspension to a glass slide for microscopic review. Configuration: Bench top design. Process Capacity: Up to 12 disposable chambers per process. Speed: Adjustable from 500 to 2,500 rpm in 100-rpm increments. Programmable Times: Range from 1 to 60 minutes in 1-minute increments. Chamber Sizes: 1, 6, and 12mL disposable chambers. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for fuctionally test. Physically damaged machine will not be accepted. | Set | √ | √
5.1 | Histology
5.1.1 | Automatic Tissue Processor, (With standard accessories) Specification : Temperature range of wax baths 45 degree to 65 degree celcius, excess temperature cut out 75 degree celcius Computerised freely selectable and storable programming facilities should be available. Easy editing and changing of programmes should be possible, even during a processing run. Infiltration time for each station should be separately programmable. Attached vacuum function device with fume control system. The vacuum device should be attachable both to the whole instrument as well as to individual stations Safely device for protection of drying of specimens in case of power failure. The bucket should go back inside the respective solution when power fails and not hang in mid air LCD display panel, audible alarm warning/error messages. Ergonomic control panel with full protected keyboard and LOC. Machine should also be able to cater to short time (quick) processing. Both immediate and delayed start functions should be available. Delayed start function atleast upto 72 hrs Interrupting an automatic process Carousel type with 12 stations (10 reagent stations, 2 wax baths). Metal containers at each station Capacity of each container 1.8 to 2 litres, Metal tissue baskets - 2 nos. capacity approx 80 - 100 cassettes to be run at one go. The tissue baskets should have less base diameter as compared to upper diameter to avoid basket getting stuck inside the containers. Firm bottom.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for fuctionally test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
5.1.2 | GROSSING STATION: Well ventilated grossing station (max 60inchs Wx32.5inch Dx92 inch H) plenum. Slotted rear of the workstation effectively that pulls fumes away from technician and exhaust to the outside. A rectangular collar mounted on top of workstation for connecting duct. Include ducts and exhaust fans. Fluorescent light and switch (mounted in the angled rear plenum behind a lens to maintain a bright work area. Work area divided between a sloped stainless work surface that drains into a stainless sink. A hand rinse spray nozzle located at the left end of the stainless work surface. A hot/cold faucet mounted behind the sink to rinse waste down the sink Include a garbage disposal can. Station has a lower base cabinet of welded stainless steel. Cabinet has two horizontal sliding doors to access the storage area and leveler. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for fuctionally test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
5.1.3 | ROTATORY MICROTOME: Rotary high precision manual microtome for paraffin sections. Mechanical automated feeding system. Integrated lockable hand wheel capable of being locked in any desired position. Section waste tray. Section thickness ranges 1to 60 micrometers. Trimming range 10 to 30 micrometer. Stable blade holder ( low profile disposable blade, S35) to ensure that no vibrations occur with section is being cut. Desirable to have arrangements to attach a cooling device for the paraffin block. Quick reclamp block holder(Universal block holder). The machine should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
5.1.4 | Tissue Floating bath (Thermostatic) Specification: Thermostatically controlled to 30 to 80˚C maximum Temperature control to +/- 0.5˚C Aluminium bath, polypropylene lid 450 W heating elements (include 2 spare heating elements), Pilot lights indicator for on/off and warn of overheating Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The tissue floating bath should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
5.1.5 | Slide Warmer, Temp. range 20-100˚C: Specification: Flattering table for stretching and drying of section tissue specimen. Temperature adjustable to 85ºC in 1ºC steps. Anodized dull black for good view of specimen. Large heating area accommodate 150-200 microscope glass slides.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer.The machine should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
5.1.6 | Cryostat: Temperature range: 0 to -35˚C, section thickness range 1 to 100 micron meter, Fine section thickness 1 to10 micron meter in 1 micron meter steps, Trimming section Thickness range 5 to 500 micron meter, Maximum specimen size 70x55mm, Total horizontal feed 28000micron meter, vertical cutting stroke 60mm, Sample retraction during return stroke 40micron meter can be deactivated, Power: 220-240volt, 50/60 Hz, Dimensions: (WxDxH) 640(w/o hand wheel) x 760x1200mm.The machine should be installed and demonstrated for fuctionallity test. Physically damaged machine will not be accepted. After sales and services should be provided. | Set | √ | √
5.1.7 | Tissue Embedding Station with wax Dispenser: Specification: Modular Paraffin Embedding Stations with Large heated work surface warmer with area for standard cable, free forceps, Cryo console with adjustable temperature of at least –50˚C, Isolated cold plate Paraffin dispenser with 3 litres or more paraffin reservoir. Manual and foot pedal controls for adjusting flow rate of paraffin. Should be programmable for preset time for melting paraffin and delay for weekend settings. Temperature setting for dispensing and thermal console. Cassette bath Mold warmer, Good illumination, should be able to operate on 22V. Temperature setting: 55 to 70˚C for cassette bath, mold warmer working area & forceps receptacle. Should supply 12-bulbs, 12- fuses, paraffin scraper-3 with each equipment. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for functionality test. After sales and services should be provided. Physically damaged equipment will not be accepted. | Set | √ | √
5.1.8 | Tissue Block cabinet (vertical): To store at least 20,000 paraffin blocks. Physically damaged tissue block cabinet will not be accepted. | Piece
5.2 | Histology & Cytology
5.2.1 | Automatic Slide Stainer ( With standard accessories) Specification: Should be suitable for at least 40 slides per batch. It should be designed for continous operation program with the storage capacity of method upto 20 methods in memory. Should be programmable upto 50 steps in each method. Operation condition: Temperature- 10 to 40 degree celsius. Relative humidity- 36% to 85%. Atmospheric pressure-70 to 106 Kpa. Capacity/Fill volumes: Solution reservoirs total (26 numbers), wash reservoirs-(5numbers), Drying stations-( 1 number). Fill level reservoirs-620-650ml. Batch operation- upto 40 slides in two baskets( two baskets=1 basket set). Facillities: Water supply-tap water, water pressure-Dynamic pressure:0.098 to 0.441MPa. Max. Static Presure:0.833MPa. Water temperature: Less than 30 degree celsius. Drain type:Method-gravity drain. Capicity -More than 15 liters. Power supply: 220-240 V, 50/60 Hz. UPS to be provided. The machine should be installed and demonstrated for functionaly test. After sales and services should be provided. Physically damaged machine will not be accepted. | Set | √ | √
5.2.2 | Slide Cabinet: To store atleast 20,000 slides of 75x25mm slide - Horizontal with anodized aluminum slide tray of 350x210x35mm, uniformly slotted rows are numbered 1 to 50 to identify the slides. Physically damaged cabinet will not be accepted. | Piece | √ | √
5.2.3 | Automated Cover slipper: Self-contained, automated instrument for glass cover slipping. Very precise and bubble-free cover slipping Consistent cover slipping quality. High throughput for up to 400 slides per hour. Convenient batch or single slide mode. Flexible mountant type control for volume and flow rate. Automatic detection for broken coverslips. Special histopathology adapter. Full self diagnostic function. Efficient fume control system. Automated rinsing system of DPX by xylene at end of the work..Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer. The machine should be installed and demonstrated for functionaly test. After sales and services should be provided. Physicallity damaged machine will not be accepted. | Set | √
6 | pathology Department: General Lab
6.1 | General Chemical
6.1.1 | Calcium chloride dihydrate | 500 gm | √ | √
6.1.2 | di-Potassium oxalate purified | 500 gm | √ | √ | √ | √
6.1.3 | Hydrochloric acid (Concentrated) | 500 ml | √ | √ | √ | √
6.1.4 | Immersion Oil | 125 ml | √ | √ | √ | √
6.1.5 | Potassium hydroxide pellets purified | 100 gm | √ | √ | √ | √
6.1.6 | Potassium Permanganate powder | 500 gm | √ | √
6.1.7 | Sodium chloride GR | 500 gm | √ | √ | √
6.1.8 | Sodium hydroxide pellets | 500 gm | √ | √
6.1.9 | Sodium hypochlorite solution GPR(5%) | 1 ltr | √ | √ | √ | √
6.1.10 | Sulphuric acid (concentrated) | 500 ml | √ | √ | √ | √
6.1.11 | Trichloroacetic acid | 500 gm | √
6.2 | Glass / Plastic Wares
6.2.1 | Bottle-Reagent (glass capacity 250ml) | pcs | √ | √ | √ | √
6.2.2 | Bottle-Reagent (glass capacity 500ml) | pcs | √ | √ | √ | √
6.2.3 | Bottle-Wash, narrow neck 20 - 30mm , capacity 500ml. High chemical resistance and contamination free despensing | pcs | √ | √ | √ | √
6.2.4 | Bottle-aspirator polypropylene. 25 litres, PP tap and screw closure. Wide shoulder for easy carrying and pouring. Supplied with tap closure and leakproof. | pcs | √ | √ | √ | √
6.2.5 | Cover glass 22mmx22mm | 10g/pkt | √ | √ | √ | √
6.2.6 | Cylinder - graduated, glass (1000ml) | pcs | √ | √ | √ | √
6.2.7 | Cylinder - graduated, plastic (1000ml) | pcs | √ | √ | √ | √
6.2.8 | Cylinder - graduated, glass (100ml) | pcs | √ | √ | √ | √
6.2.9 | Cylinder - graduated, plastic (100ml) | pcs | √ | √ | √ | √
6.2.10 | Cylinder - graduated, glass (500ml) | pcs | √ | √ | √ | √
6.2.11 | Cylinder - graduated, plastic (500ml) | pcs | √ | √ | √ | √
6.2.12 | Flask conical, glass (250ml) with screw cap | pcs | √ | √ | √ | √
6.2.13 | Flask conical (500ml) with screw cap | pcs | √ | √ | √ | √
6.2.14 | Flask conical, glass (1000ml) with screw cap | pcs | √ | √ | √ | √
6.2.15 | Flask conical, glass (2000ml) with screw cap | pcs | √ | √ | √ | √
6.2.16 | Flask conical, glass (3000ml) | Pcs | √ | √ | √
6.2.17 | Flask Volumetric, glass 50ml | pcs | √ | √ | √ | √
6.2.18 | Flask Volumetric, glass 100ml | pcs | √ | √ | √ | √
6.2.19 | Flask Volumetric, glass 500ml | pcs | √ | √ | √ | √
6.2.20 | Flask Volumetric, glass 1000ml | pcs | √ | √ | √ | √
6.2.21 | Test tube without rim, glass (10x75mm) | pcs | √ | √ | √ | √
6.2.22 | Test tube without rim, glass (12x100mm) | pcs | √ | √ | √ | √
6.2.23 | Test tube without rim, glass (12x75mm) | pcs | √ | √ | √ | √
6.2.24 | Test tube without rim, glass (15x125mm) | pcs | √ | √ | √ | √
6.2.25 | Test tube with rim, glass (15x150mm) | pcs | √ | √ | √ | √
6.2.26 | Volumetric Glass pipette 1ml, calibrated for delivery | pcs | √ | √ | √ | √
6.2.27 | Volumetric Glass pipette 2ml, calibrated for delivery | pcs | √ | √ | √ | √
6.2.28 | Volumetric Glass pipette 5ml, calibrated for delivery | pcs | √ | √ | √ | √
6.2.29 | Volumetric Glass pipette 10ml, calibrated for delivery | pcs | √ | √ | √ | √
6.2.30 | Test tube holder brass | pcs | √ | √ | √ | √
6.2.31 | Beaker, (polypropylene-transparent, squat form, 1000ml, can be used continuously at 100˚C with spout and graduations) | pcs | √ | √ | √ | √
6.2.32 | Beaker, (glass, squat form, 2000ml, can be used continuously at 100˚C with spout and graduations) | pcs | √ | √
6.2.33 | Tubes, Centrifuge, conical. 15ml tubes with cap, graduated in 0.5mL subdivisions from 2 to 14.5mL, 17x119mm polypropylene. | 50pcs/pkt | √ | √ | √ | √
6.2.34 | Tubes, Centrifuge, conical. 50ml tubes with cap, graduated in 0.5mL subdivisions from 2 to 49.5mL, 30x119mm polypropylene. | 50pcs/pkt | √
6.2.35 | Microscope slide, single frost, thickness 1 - 1.2mm. Dimensions: 75x25mm. The raised opaque coating can easily be marked with a pencil or ink, resistant to every commonly used solvent and it prevents the slides from adhering and becoming scratched. Ready to use | 50slide/pkt | √ | √ | √ | √
6.2.36 | Microscope slide, dimension 76mmx26mm, thickness 1 - 1.2mm, made from extra white glass with finely ground edges.Ready to use | 50slide/pkt | √ | √ | √ | √
6.2.37 | Capillary Tube, Heparinized, hematocrit (75mm length) | 100/pkt | √ | √ | √ | √
6.2.38 | Capillary Tube, non-heparinized, hematocrit (75mm length) | 100/pkt | √ | √ | √ | √
6.2.39 | Capillary sealing compound (Haematrocrit tube): Sealing compound provides perfect seal. Capillary Tubes are put into the sealing compound with a twesting motion. | 10trays/ pkt | √ | √ | √ | √
6.2.40 | Brushes, test tube. Head diameter: 11mm. For tubes diameter: 10 to 13mm | pcs | √ | √ | √ | √
6.2.41 | Brushes, test tube. Head diameter: 29mm. For tubes diameter: 16 to 25mm | pcs | √ | √ | √ | √
6.2.42 | Parafilm, roll (width 2 inches x length 250 ft) | pcs | √ | √ | √ | √
6.2.43 | Pipettes, Pasteur, maximum bulb capacity 3.5ml | 100/pkt | √ | √ | √ | √
6.2.44 | Basket, polished stainless steel for test tube: Dimensions (wxdxh),mm. 150x150x150 | pcs | √ | √ | √ | √
6.2.45 | Basket, polished stainless steel, test tube: Dimensions (wxdxh), mm. 200x200x200 | pcs | √ | √ | √ | √
6.2.46 | Basket, polypropylene, test tube, draining: Dimensions (wxdxh),mm. 400x300x100 | pcs | √ | √ | √ | √
6.2.47 | Test tube rack, polypropylene. Hole dia: 16mm. Nos. of hole: 60. Array: 5x12. Height: 70mm. | pcs | √ | √ | √ | √
6.2.48 | Test tube rack, polypropylene. Hole dia: 13mm. Nos. of hole: 72. Array: 6x12. Height: 57mm. | pcs | √ | √ | √ | √
6.2.49 | Pipettor tips, volume 5 -50 uL, optimal tightness fitting | 1000/ pkt | √
6.2.50 | Pipettor tips, volume 50-200 uL, optimal tightness fitting | 1000/ pkt | √ | √ | √ | √
6.2.51 | Pipettor tips, volume 200 -1000 uL, optimal tightness fitting | 1000/pkt | √ | √ | √ | √
6.2.52 | Pipettor tips, volume 1000 -5000 uL, optimal tightness fitting | 500/ pkt | √ | √
6.2.53 | Pipettor Stand: Holds 5 to 7 pipettors, fits most brands | pcs | √ | √ | √ | √
6.2.54 | Diamond pen | pcs | √ | √ | √ | √
6.2.55 | Filter paper, qualitative: diameter 110mm, provides fine retention and slow flow rate, suitable for general filtration of many types of samples. | 100/pkt | √ | √ | √ | √
6.2.56 | Filter paper sheet 46cmx57cm. provides fine retention and slow flow rate, suitable for general filtration of many types of samples. | 500/pkt | √ | √ | √
6.2.57 | Blotting paper. (Highly absorbent) effectively absorbs liquids including testing or diagnostic test kits, spill control and clean up of toxic and non-toxic sustances. Minimum dimension of 12 inches width x 5 metres length and thickness 0.21mm | rolls | √ | √
6.2.58 | Paper, lens cleaning tissue, high quality lens cleaning tissues, 80mmx100mm | 50sheet/pkt | √ | √ | √ | √
6.2.59 | Spirit lamp, brass | pcs | √ | √ | √ | √
6.2.60 | Digital count down timer.2 count down timers each for up to 23hr 59min 59sec. 12/24hr clock. Extra loud alarm. Bench top standing or wall mounted. Memory function. Powered by AAA battery (supplied) | pcs | √ | √ | √ | √
6.2.61 | Blood Tourniquet. Highly flexible, one hand operation with quick and secure release buckle. | pcs | √ | √ | √ | √
6.2.62 | Sample carrying tray with handle, autoclavable Polypropylene Tray to carry reagents or laboratory apparatus from one place to another, single handedly. | pcs | √ | √ | √ | √
6.3 | CLINICAL
6.3.1 | Acetic acid glacial | 500ml | √ | √ | √ | √
6.3.2 | Ammonium sulphate | 500g | √ | √
6.3.3 | Barium chloride | 500g | √ | √ | √ | √
6.3.4 | Benedicts reagent | 500ml | √ | √ | √ | √
6.3.5 | Ehrlichs aldehyde reagent | 100ml | √ | √ | √ | √
6.3.6 | Fouchet reagent | 100ml | √ | √ | √
6.3.7 | Liquid ammonium solution | 500ml | √ | √
6.3.8 | Pregnancy test kit (sensitivity as low as 20mIU/ml hCG) | 50strips/ pkt | √ | √ | √ | √
6.3.9 | Sodium nitroprusside | 100g | √ | √ | √
6.3.10 | Sulphosalicylic acid | 100g | √ | √ | √ | √
6.3.11 | Sulphur powder | 500g | √ | √ | √ | √
6.3.12 | Hydrogen peroxide | 500 ml | √ | √ | √ | √
6.3.13 | Urinary strip 10 parameters | 100 strips/ Vial | √ | √ | √
6.3.14 | Urinary strip 2 parameters (albumin and sugar) | 100strips/ Vial | √ | √ | √ | √
6.3.15 | LE (SLE) slide test | 50 test/ pkt | √ | √
6.3.16 | Occult Blood test kit | 50strips/ pkt | √ | √ | √ | √
6.3.17 | Semen diluting fluid | 100ml | √ | √ | √
6.3.18 | pH indicator paper (2-7pH) | 50strips/ pkt | √ | √ | √ | √
6.3.19 | pH indicator paper (8-10.5pH) | 50strips/ pkt | √ | √ | √ | √
6.3.20 | Slide mailer (5 Slides) | Piece | √ | √ | √ | √
6.3.21 | Ayre`s Spatula | pcs | √ | √ | √ | √
6.3.22 | Coplin Jar |  | √ | √ | √ | √
6.4 | EQUIPMENT (Every equipment supplied must be installed with end user training and demonstrated for functionality)
6.4.1 | Balance: Readability: 0.01g. Max. Load: 610 g. Repeatability: 0.01g. Linearity: 0.02g. Weighing pan: dia160mm. Power supply: AC 220-240V, 50/60Hz. Power cord with round three pin top | Set | √ | √ | √ | √
6.4.2 | Bench top Centrifuge: Swing out rotor with bucket with holders for 12x100mm test tube, and 15 ml tubes. Maximum RPM: 2000 - 5000 RPM. Maximum relative centrifuge force: 4640 x g. Main control: LCD microprocessor controlled. Power supply: AC 220-240V, 50 HZ. Power cord with round three pin top. Operating environmental condition: -10˚C to +40°C, RH >80%. | Set | √ | √ | √ | √
6.4.3 | Incubator: Temperature range: Ambient + 5°C to 70°C. Temperature control: Digital P.I.D. controller. Display: Dual LED display. Capacity: 100 - 300 liters. Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √ | √ | √ | √
6.4.4 | Bench top Centrifuge: Fixed angle rotor with holders for 12x100mm test tube, and 15 ml tubes. Maximum RPM: 2000 - 5000 RPM. Maximum relative centrifuge force: 4640 x g. Main control: LCD microprocessor controlled. Power supply: AC 220-240V, 50 HZ. Power cord with round three pin top. Operating environmental condition: -10˚C to +40°C, RH >80%. | Set | √ | √ | √ | √
6.4.5 | Deep Freezer. with microprocessor controlled temperature regulation with digital display. Temperature Range: - 20 to -40ºC. Capacity: 250 -300 Liters. Power: 220-240V, 50/60Hz.Power cord with round three pin top. Should be supplied with appropriate storage racks & Voltage stabilizer. Operating environmental condition: - 10 to +40°C, RH >80%. | Set | √ | √ | √
6.4.6 | Deep Freezer. Vertical type with microprocessor controlled temperature regulation with digital display. Temperature Range: - 70ºC. Capacity: 250 -300 Liters. Power: 220-240V, 50/60Hz.Power cord with round three pin top. Should be supplied with appropriate storage racks with individual storage compartments & Voltage stabilizer. Operating environmental condition: - 10 to +40°C, RH >80%. | Set | √ | √
6.4.7 | Distillation plant: (mono distiller) with all the required accessories and suitable for bench mounting. Thermostatic low water cut-off. Destillate withdrawal through outlet at the condenser. Evaporator easily accessible for cleaning purposes. economical energy consumption by distillation of the heated cooling water. Throughput: 4 liters/hour. Power supply: 220-240V 50/60Hz. | Set | √ | √ | √ | √
6.4.8 | Haematocrit Centrifuge. High speed bench-top centrifuge for haematocrit determinations with brushless drive. Sedimentation is completed at 16,060 x g in a little over 6 minutes. Quick entry foil keypad. Easy to read digital display. Easy entry and display of current parameter values. Metal housing. The lid of the rotor is the evaluation disk and cover in one. The haematocrit value can be read off the lid / evaluation disk after centrifugation. Motor overheating protection. Imbalance switch-off. Timer: 99 minutes max. Maximum speed: 13,000 rpm. (16,060 x g.) in 100 rpm increments. Noise level: 58 dB(A). Operating environmental condition: -10 to +40˚C, RH >80%. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √ | √
6.4.9 | Pipettors, variable volume 5 - 50µl, full volume range, thumb press tip ejector, calibration simple to perform by user. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
6.4.10 | Pipettors, variable volume 50 - 200µl, full volume range, thumb press tip ejector, simple calibration to perform user friendly. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
6.4.11 | Pipettors, variable volume 200-1000µl, full volume range, thumb press tip ejector, simple calibration to perform and user friendly. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
6.4.12 | Pipettors, variable volume 1000 - 5000µl, full volume range thumb press tip ejector, simple calibration to perform and user friendly. Can be used with tips from all manufacturers. | Piece | √ | √ | √ | √
6.4.13 | Microscope, binocular: (for routine work) Illumination inbuilt,Right-handed stage control. Objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X. Complete with power cord, 220/230 - 50/60HZ. | set | √ | √ | √ | √
6.4.14 | Microscope: Multi head with CCTV system and LCD monitor. Illumination: with objectives: 4X, 10X, 40X, & 100X. Eye piece: 10X Complete with round 3 pin power cord. Power supply: 220V, 50/60HZ. Right-handed stage control. Should be supplied with micophotography attachement and latest computer. | Set | √
6.4.15 | pH Meter: pH Range: 0.00 - 14.00pH. Resolution: 0.01pH. Relative Accuracy: +/- 0.01pH. Millivolt Range: -1999 to 1999mV. Temperature Range: 0.0 - 100.0.Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin round top | Set | √ | √
6.4.16 | Refrigerator with glass door for storage of laboratory reagents: Capacity: 1000 to 1300 liters. Temperature Range: 0 to+10° C. Temp. Accuracy: +/- 0.3 at 4°C. Temp. Controller: Digital P.I.D. controller. Display: Dual LED display. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √
6.4.17 | Refrigerator with glass door for storage of laboratory reagents: Capacity: 350 to 800 Liters. Temperature Range: 0 - 10° C. Temp. Accuracy: +/- 0.3 at 4°C. Temp. Controller: Digital P.I.D. controller. Display: Dual LED display. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √ | √ | √
6.4.18 | Vortex mixer. Speed Range: 0 - 3000 RPM. Operation Mode: Touch on and constant on Head: Rubber Mixing Cup. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top | Set | √ | √
6.4.19 | Water bath: Temperature range: Ambient + 5°C to 99°C. Temp. Accuracy: +/-2˚C at 37˚C. Temp. Controller: Digital P.I.D. controller, 4Zone PID. Display: Dual LED display. Capacity: 20 - 30liters. Power supply: AC 220V-240V, 50Hz. Power cord with round 3 pin top and appropriate Voltage stabilizer | Set | √ | √ | √ | √
6.4.20 | Hot air oven: Temp. Range: Ambient + 5°C to 220°C. Temp. Control: Digital P.I.D. controller. Display: Dual LED display. Capacity: 50 - 100 liters. Power cord with round three pin top, Voltage AC 220V-240V. | Set | √ | √ | √ | √
6.4.21 | BILIRUBIN METER: Measurement Parameters: 0 - 30 mg/dL (total bilirubin in serum) Light Source: 6V 1.5A tungsten lamp Wavelength: 461, 551 nm Formula: 2 wavelength differential formula Measurement cell: Capillary tube Sample Volume: 50 - 60uL Display: 3-digit LED Power: 220 - 240V, 50 Hz. power cord with round 3 pin top | set | √ | √ | √
6.4.22 | Hemoglobin color scale | pkt | √ | √ | √ | √

## pediatric
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
1 | Pediatric Department
1.1 | Bedside SpO2 patient monitoring system: Multicolor display screen with a black background for contrast, Backup audible alarm, Multiple-language graphical user interface to support variety of users, 96-hour trend memory with data captured every four seconds,Simple, intuitive operation with a space-saving design Displays lethysmographic waveforms, pulse amplitude and current measured SpO2 and pulse rate (wrap type sensor probe with 9 pins) | Set | √ | √ | √
1.2 | CP Monitor (Intensive care monitors) – Pediatric: NIBP, SPO2, ECG, Respiratory rate, heart rate & Temperature, Battery, 15" - 17" colour TFT Display, 6 channels, Modular in design, should have module for measuring, 5 lead ECG with electrocautery filter, multilead ST segment analysis, lethal arrhythmai alarms with 5 lead ECG cable Set, Spo2 measurement with plethysmograph-Neonate and pediatric Spo2 sensors, pitch should vary with saturation, NIBP measurement with pediatric and neonate cuffs, Respiration measurement by impedance method, Provision for invasive pressures ( CVP and IBP) simultaneously with 20 disposable transducers and 2 reusable pessure CVP and IBP pressure cables , Temperature measurement, Co2. | Set | √ | √
1.3 | Capnography (ETCO2) monitor, Pediatric : Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. Suitable voltage stabilizer. EtCO2 Measurement Range: 0-100 mmHg (display dependent); 0-13%; 0-12.5kPa. EtCO2 Accuracy 0-40 mmHg ±2 mmHg, 41-70 mmHg ±5%, 71-100 mmHg ±8%. LoFlo 0-40 mmHg ±2 mmHg, 41-70 mmHg ±-5%, 71-100 mmHg ±8%. ±12% for respiration rate above 80 BPM.EtCO2 Resolution: 1 mmHg 0-100 mmHg. EtCO2 Stability: Short Term Drift: Drift over four hours shall not exceed 0.8 mmHg maximum. | Set | √ | √
1.4 | Long Term Drift: Accuracy specification will be maintained over a 20-hour period after zeroing. EtCO2 Noise: RMS noise of the sensor less than or equal to 0.25 mmHg at 7.5% CO2 EtCO2 Rise Time (10-90%): Mainstream: Less than 60 ms (Adult reusable or single patient use); Less than 50 ms (Infant reusable for single patient use). Sidestream: less than 200 ms. Respiration Rate (RR) Range: 0-150 breaths per minute (BPM). Respiration Rate (RR) Accuracy: ±1 breath. Compensations: Barometric pressure 550 -780 mmHg (EtCO2 Alarm Limits: User-selectable/High 5 to 100 mmHg, Low 0 to 95 mmHg/OFF. Respiration Rate (RR) Alarm Limits: User-selectable, High 5 - 150 respirations per minute, Low 0 - 100 respirations per minute/OFF. Airway Adapter Deadspace: Environmental: Operating Temperature: -10° to 40°C, RH >80%, height 2000m.
1.5 | Child ACLS Manikin with Interactive Arrhythmia Simulator : PR; Palpable and visual landmarks; • Fully articulated head, neck and jaw; Airway Management; • Realistic anatomy of the mouth, tongue, oral pharynx, larynx, epiglottis, vocal cords, trachea and esophagus; • Separate left and right lungs for auscultation; • Suctioning capabilities; • Oral, nasal and digital intubation capabilities; • Inflatable stomach bladder to indicate esophageal intubation; I.V. Injection (at arm): • Realistic size of an child’s veins; • Articulated at the biceps for antecubital and dorsal access; • Realistic flashback from a pressurized system; Blood Pressure Measurement: Instructor determines systolic and diastolic levels, heart rate and sound volume; Speaker in arm reproduces real blood pressure sounds; 5 Korotkoff phases can be turned on and off; Auscultatory gap can be turned on and off; | Set | √ | √
1.6 | Intraosseous Infusion: • Palpable landmarks include the patella, tibia and tibial tuberosity; - Pressurized system allows aspiration of fluid; Femoral Access: • Palpable arterial pulse; • Palpable symphysis pubis and anterior superior iliac spine; • Realistic flashback from a pressurized system; Defibrillation Chest: • Internal load box absorbs full strength of every shock;• Manual, semi-automatic, automatic defibrillation; • Monitor manikin like a real patient at 4 ECG sites and 2 defibrillation sites; • Compatible with all standard brands and types of defibrillators and ECG-monitors; Interactive ECG-Simulator: This external unit allows defibrillator and pacer training directly on the manikin. (A training without a manikin is also possible). Various ECG rhythms can be generated and displayed on an ECG-monitor (not included) for diagnostic and defibrillation training. All arrhythmias can be displayed in adult or paediatric mode so that a total of 34 rhythms are available.
1.7 | Digital weighing scale capability to measure 1 milligram to 1000 milligram with flat top | Set | √ | √ | √
1.8 | Digital Pediatric NIBP and Spo2 (Portable) monitor, Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. LCD display with heart rate, Spo2 , pulse rate, min. 2 hours battery back-up, with all accessories. | Set | √ | √ | √
1.9 | Endotracheal tube pressure cuff inflator | Set | √ | √
1.1 | Electro chest physiotherapy device (vibrator/occillator) for Pediatric | Set | √ | √
1.11 | ICU Pediatric Bed : Height-adjustable electric pediatric intensive care bed (4 sections, with column motors, all side tilt) | Set | √ | √
1.12 | ICU Ventilator (Respiratory Ventilator) Pediatric : 1. Should have facility for Invasive and Non-Invasive ventilation; 2. Microprocessor Control suitable for Pediatric (one onwards) and adult ventilation; 3. Electromagnetic Compatible Hinged arm holder for holding the circuit; 4. Should have built in touch color screen TFT display of minimum 10” or more for display; Of waveforms and Monitored value; 5. Should have inbuilt facility to upgrade with EtcO2; 6. Facility to Measure and display:- a) Status indicator for ventilator mode; b) Battery indication; c) Pressure Vs time Vs volume Vs time, flow Vs time 3 curves/ waveforms; d) Alarm Setting; 7. Automatic compliance and leakage compensation for circuit and ET Tube. 8. Should have facility of log book, for events and alarms with date & time; 9. Should have following Settings; a) Tidal volume (Minimum at least 20 ml, Maximum up to 2000ml); b) Inspiratory Pressure (0 upto 80 cm of H20); c) Respiratory rate 1 to 80 bpm; d) Apnoea back up rate; e) CPAP/PEEP; f) Pressure support; g) Fi02; h) Pause Time; i) Pressure & flow Trigger; j) Inspiratory flow up to 120 Lpm; Temperature of delivered Gas on LED display; c) Temperature should be adjustable; d) Jar should be autoclavable; 17. Quality Certification : Valid CE/BIS/US FDA; 18. Demonstration of the quoted model is must, preferable on site; 19. Nebulization assembly compatible with ventilator and circuit; 20. Should have interface facility; 21. Flow sensor-Should have life more than 1 year; 22. Expiratory Unit- Life should be more than 3yrs; 23. Data storage facility for at least 24hrs; 24. Internal rechargeable battery at least 60 min. backup; 25.; 26. Should have flow sensors having long life and the company shall specify the life cycle and the cost of the flow sensors at the time of quoting the tender; 27. CMC/ AMC for the atleast 5yrs and Cost of consumables spares; 28. Source :Indigenous/Imported; 29. Warranty 2-5 years from the date of installation; 30. Standard Accessories along with: a) Patient breathing circuit of silicone for Paediatric (reusable) 3 Sets; Disposable breathing circuits 100 nos. | Set | √ | √
1.13 | Peak pressure, mean pressure, CPAP/PEEP, expiratory tidal volume, total frequency, spontaneous frequency, lung function monitoring including C20/C,R,C,Tc, RVR, have CE mark and FDA approval for use on nenonates, Standard scope of supply must include Basic Unit, Trolley, 2 silicone Hose Sets, flow sensors/Exp. Valves - 2 Sets, each unit), O2 cell, oxygen connecting hose, F & P servo control Humidifier with two reusable chamber, Nebuliser circuit support arm, Instruction manual, input power : AC 220-240 volts, suitable voltage stabilizer, operational environmental condition : humidity above 80%, temperature -5 to +40, altitude 2000m, backup source, servo controlled stabilizer to be included. Power supply: 220V-230V, Power cord of round 3 pin type plug top , Suitable UPS.
1.14 | Medication cart : Material: Aluminum -Alloy, ABS plastic & stainless steel, Overall size: 850mm*520mm* 1010mm, The whole trolley is including 6 layers of drawers, can be pulled out from both sides, self locked when closed. Totally 50medicine boxes on all drawers. Other components: 2 pcs of garbage bins, 1 pc of transparent file box, 1pc of sliding side board and 1 set of five hooks hanging device | Set | √ | √ | √
1.15 | Nasal oxygen high flow system with integrated humidifer with flow generator with regulatable FiO2 with nasal cannula and circuit for Pediatric | Set | √ | √
1.16 | Oxygen hood : ideal round shape & flat top enhancing visibility & optimum circulation of oxygen, Polycarbonate, unbreakable, autoclavable & should have oxygen inlet nozzle : Extra Large- 13.0”L x 10.4”W x 8.5”H; Nele opening dimension: 6.3”W x 5.0”H | Piece | √ | √ | √ | √
1.17 | Oxygen hood : ideal round shape & flat top enhancing visibility & optimum circulation of oxygen, Polycarbonate, unbreakable, autoclavable & should have oxygen inlet nozzle : Large- 13.0”L x 10.4”W x 8.5”H; Nele opening dimension: 6.3”W x 5.0”H | Piece | √ | √ | √ | √
1.18 | Oxygen hood : ideal round shape & flat top enhancing visibility & optimum circulation of oxygen, Polycarbonate, unbreakable, autoclavable & should have oxygen inlet nozzle : Medium- 7.0”L x 7.0”W x 5.4”H; Nele opening dimension: 3.9”W x 2.5”H | Piece | √ | √ | √ | √
1.19 | Oxygen analyzer: Range 0-100% oxygen. Display accuracy 0.1%, electrical connector: battery operated, operating temperature 0-50° C, output signal Oxygen in %, sensor type galvanic (electro chemical), response time < 5 sec. to T90, accuracy ± 1 % at full scale, repeatability ±1% of signal, temperature compensation: compensated, maximum storage temperature -20° C to + 60° C, power: alkaline batteries, Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. | Piece | √ | √
1.2 | Pediatric Otoscope | Set | √ | √ | √
1.21 | Pocket pen tourch for pupil examination with Pupil Gauge | Piece | √ | √ | √
1.22 | Stethoscope - Paediatric | Piece | √ | √ | √ | √
1.23 | Ultrasonic nebuliser, heavy duty jet nebuliser: heavy duty and provided with support for tubings, fitted with an electronic timer and should operate continuously. The output of drug delivered should be electronically adjustable between 0.1-0.5ml/min. The medium size of particles generated should be 5 microns, facility for automatic stoppage when liquid drug is exhausted. The nebulizer chamber and tubing should be sterilizable. 4 masks (2 each for infants and older children) must be supplied with the unit 2 extra chambers must be supplied with each unit. Input power: AC 220-240 volts, servo controlled stabilizer to be included. | Set | √ | √ | √
1.24 | Beneview T5 CP monitor for Spo2 probe : wrap type | Piece | √ | √
1.25 | Beneview T5 CP monitor for BP cuff sizes 6-11 cm, , 10- 13 cm, 10-15 cm, 14-21.5 cm | Set | √ | √
1.26 | Beneview T5 CP monitor for Spo2 probe: boot type: Paediatric size | Piece | √ | √
1.27 | Beneview t5/8 (Mindray) patient monitor reusable NIBP cuff (arm) small infant circumference (cm) 7 to 13, bladder width (cm) 5.8 | Piece | √ | √
1.28 | Beneview t5/8 (Mindray) patient monitor reusable NIBP cuff (arm) infant circumference (cm) 10 to 19, bladder width (cm) 9.2 | Piece | √ | √
1.29 | Beneview 5/8 (Mindray) patient monitor reusable NIBP cuff (arm) paediatric circumference (cm) 18 to 26, bladder width (cm) 12.2 | Piece | √ | √
1.3 | Beneview t5/8 (Mindray) patient monitor wall mount bracket | Piece | √ | √
1.31 | Beneview t8 (Mindray) patient paediatric ICP transducer | Piece | √ | √
1.32 | CO2 sample line, watertrap, airway adapter for Mindray SynoVent E3 ventilator paediatric | Piece | √ | √
1.33 | Disposable breathing tube kid (Child) for Mindray SynoVent E3 ventilator with heating wire | Piece | √ | √
1.34 | Disposable filter (round) for Mindray SynoVent E3 ventilator | Piece | √ | √
1.35 | Infant reusable comfort NBP cuff 10 to 15 cm for Philip MP 60 monitor | Set | √ | √
1.36 | Oxygen sensor probe compatible for Datex Ohmeda | Piece | √ | √
1.37 | Pediatric circuit tubings with heated wire, silconised and autoclavable for Paediatric ventilator compatible with puritan bennett 840. | Piece | √ | √
1.38 | Paediatric reusable inspiratory bacteria filter (Re/Flex)compatible with Puritan bennett 840 | Piece | √ | √
1.39 | Paediatric reusable comfort NBP cuff 14 to 21.5 cm for Philip MP 60 monitor | Piece | √ | √
1.4 | Phoenix nCPAP 200 breathing circuit (paediatric) starter kit ( Phoenix nCPAP 200 pressure generator) | Piece | √ | √
1.41 | Reusable silicon Breathing tube kit(Child) (115-00 4502-00) for Mindray SynoVent E3 ventilator | Piece | √ | √
1.42 | Reusable silicon Breathing tube kit(Infant) (115-00 4503-00) for Mindray SynoVent E3 ventilator | Piece | √ | √
1.43 | Renovar CP monitor SpO2 paediatric probe (reusable | Piece | √ | √
1.44 | Schillers oxywave Spo2 monitor compatible probe -Paediatric size | Piece | √ | √
1.45 | Ultrasonic atom sanitizer 303 spare tubings with mask | Piece | √ | √
1.46 | Beneview t5/8 (Mindray CP monitor) reusable SpO2 probe, infant wrap type | piece | √ | √
1.47 | Philips MP60 CP montor, reusable SpO2 probe, wrap type. | Piece | √ | √
1.48 | Philips MP60 CP montor, reusable SpO2 probe, Boot type. | Piece | √ | √
1.49 | Pediatric ophalmoscope | Piece | √ | √
1.5 | UPS for ventilators | Nos | √ | √
1.51 | Accessories for AIRVO2
1.52 | Filter 2 pack for AIRVO2 | set | √ | √
1.53 | Optiflow junio neonatal circuit | set | √ | √
1.54 | Optiflow junior nasal canula (OPT 316) | set | √ | √
1.55 | Optiflow junior nasal canula (OPT 318) | set | √ | √
1.56 | Optiflow junior wiggle pads (OPT 012) | pieces | √ | √
1.57 | Beneview t5/8 (Mindray CP monitor) reusable SpO2 probe, ped clip type | piece | √ | √
1.58 | ECG cable Philips MP-60 CP monitor, reusable ECG cable (M1971A) |  | √ | √
1.59 | Beneview t5/8 (mindray CP monitor) reusable ECG cable (EL6501-B) |  | √ | √
1.6 | Mindray ventilator face mask for CPAP |  | √ | √
1.61 | Fisher and Paykel Health Care(pediatric size).
2 | General consumable for Pediatric/Neonate
2.1 | Atom CPAP Nasal Canula,SS | Piece | √ | √
2.2 | Atom CPAP Nasal Canula,S | Piece | √ | √
2.3 | Atom CPAP Nasal Canula,M | Piece | √ | √
2.4 | Atom CPAP Nasal Canula, L | Piece | √ | √
2.5 | Bubble CPAP system starter kits : CPAP Mask, harness, gas supply line two flow meters (BC161-10), nasal prong(different sizes), infant bonnet(different sizes) and nasal tubing. | Set | √ | √
2.6 | Central peripherial line (PICC), neonate | Set | √ | √
2.7 | Central peripherial line (PICC), pediatric 22G | Set | √ | √
2.8 | Coban elastic self-adherent wrap tape | Piece | √ | √
2.9 | Disposable Paediatric ECG chest electrode(FOAM): Application : Paediatric; Shape: Teardrop ,Backing : Foam Gel : Solid Gel or Wet Gel; Sensor : Low impedance Ag/Agcl sensor: Size: 27mm | Piece | √ | √
2.1 | Disposable ECG leads for neonate compatible for Philips CP monitor MP60/MX600 latex free product code:13953D | Piece | √ | √
2.11 | EZ intraosseous needle size 15 mm needle | Set | √ | √
2.12 | EZ intraosseous needle size 25 mm needle | Set | √ | √
2.13 | EZ intraosseous needle stablizer | Packet | √ | √
2.14 | EZ intraosseous Connector | Packet | √ | √
2.15 | Extra-Spike® Plus: • Withdrawal cannula for multiple withdrawal from stock bottles and for supplementary injections, safe and handy; • Slim, short spike for easy piercing of stronger stoppers; • Easy opening and closing with one hand by grip plate centrally arranged to spike with integrated, tight snap with colour code; • Easy withdrawal and supplementary injection, reliable protection against microbial contamination by integrated airvent filter; • Luer lock-connector. | Piece | √ | √
2.16 | ECG paper compatible with Philips CP Monitor MP60 & Mx600 | Piece | √ | √
2.17 | Filter paper for ventilator humidifier compatible with Mindray | Piece | √ | √
2.18 | Filter paper for ventilator humidifier compatible with Drager | Piece | √ | √
2.19 | Heparinized insulin 1 ml syringe for ABG without needle | Piece | √ | √
2.2 | IV extension lines internal diameter 1mm, length 11 cm | Piece | √ | √
2.21 | IV extension lines internal diameter 1mm, length 30 cm | Piece | √ | √
2.22 | IV extension lines internal diameter 1 mm, length 50 cm | Piece | √ | √
2.23 | IV extension lines internal diameter 1mm, length 100 cm | Piece | √ | √
2.24 | IV Set (60 drops per ml) | Piece | √ | √ | √ | √
2.25 | Incentive spirometer (for Lung exercise) pediatric and adult use 3 chamber of pf 600, 900 and 1200 cc/sec inhaling volume, three coloured balls for easier indentification of inhaled capacity | Set | √ | √
2.26 | Intravenous line filter (bacterial) 0.2 micro supor, filters bacteria and particles, luer lock connector, latex free port for supplement medication, effective air ventilation regardless of position | Set | √ | √
2.27 | Intraosseous infusion needle, 14 gauge | Piece | √ | √ | √ | √
2.28 | Intraosseous infusion needle, 16 gauge | Piece | √ | √ | √ | √
2.29 | Intraosseous infusion needle, 18 gauge | Piece | √ | √ | √ | √
2.3 | Microdrip Set with burrete | Piece | √ | √ | √ | √
2.31 | Meconium aspirator (Pengiun Type) | Piece | √ | √ | √ | √
2.32 | Octopus is a polyurethane multi-lumen (two) extension tube for I.V. catheters, transparent and with low priming volume. Priming volume (0.3 to 0.5 ml) with non-return valve. Length 10 cm, internal diameter 1 mm and external diameter 2 mm. | Packet | √ | √
2.33 | Paediatric double-lumen central venous catheter: Radiopaque central venous polyurethane catheter for injecting or infusing different solutions and/or medications simultaneously; Tray presentation containing; 1 double-lumen catheter (3Fr) with centimetre markings from 4 cm from distal tip; 1 short I.V. cannula 24G - Length 19 mm; 1 introducer needle 22G - Length 40 mm; 1 Nitinol straight guidewire (particularly kink-resistant) with guide advancer - Length 50 cm; 2 dilators 3Fr (Length 3 cm and 5 cm); 1 secondary fixation wing; 2 injection caps; 1 scalpel; 1 5 ml syringe | Set | √ | √
2.34 | Paediatric double-lumen central venous catheter: Radiopaque central venous polyurethane catheter for injecting or infusing different solutions and/or medications simultaneously; centimetre markings from 5 to 12 cm; Tray presentation containing: 1 double-lumen catheter; 1 introducer needle (Ext. Ø 0.8 mm – 21G -Length 40 mm); 1 short I.V. cannula 22G -Length 25 mm; 1 “J” or straight guidewire (Length 40 cm -Ø 0.46 mm) 4,5Fr; 2 dilators (length: 38 mm and 62 mm); 1 secondary fixation wing; 2 injection caps; 1 scalpel; 1 5 ml syringe. | Set | √ | √
2.35 | Spare nebuliser chamber with tubings and pediatric mask | Set | √ | √ | √ | √
2.36 | Self adhesive urine sample bag - Pediatric | Piece | √ | √ | √
2.37 | Sterile IV water bottle (one litre) | Piece/Packet | √ | √
2.38 | Sterile Skin cleansing alcohol preparation (webcol) saturated with 70% isopropyal alcohol | Packet | √ | √
2.39 | Steri-strip antimicrobial skin closures | Piece | √ | √
2.4 | Single channel disposable sensor kit (20) part no M1567A for MP60 Philips CP monitor | Set | √ | √
2.41 | Treacheal suction Set (specimen collection for sampling of secretions from tracheobronchial tree, transparent container) | Piece | √ | √
2.42 | Umbilical catheters 3.5 F | Piece | √ | √ | √
2.43 | Umbilical catheters 5 F | Piece | √ | √ | √
2.44 | Vaccu suck. Crown tip-Yankaur suction set. Tube length 2.5 meter. | Packet | √ | √
3 | Consumables of ABG Machine
3.1 | Cartridges, Test Cards (OPTI CCA TS) | 25 test/ Packet | √ | √
3.2 | Thermal Printing Paper (OPTI CCA TS) | Roll | √ | √
3.3 | Calcium Gas, Bottle (OPTI CCA TS) | Bottle | √ | √
3.4 | Mindray VS 800 BP cuff size 18 x 26 cms, arm |  | √ | √
3.5 | circumference bladder. | Set
3.6 | Mindray VS 800 BP cuff size 9.8 x 13.3 cms, arm |  | √ | √
3.7 | circumference bladder. | Set
3.8 | Welch Allyne, NIBP monitor, BP cuff, |  | √ | √
3.9 | size 12 x 16cms | Packet
3.1 | Welch Allyn, NIBP monitor, BP cuff, size 9.8 x 13.3 cms. |  | √ | √
4 | Additional for Neonatal consumables
4.1 | Coban 3M 1inch | box | √
4.2 | Tegaderm 1 1/2 inch | box | √
4.3 | Extrathin CGF dressing 10cm | box | √
4.4 | Sterilised prepacked umbilical tape for umbilical catheterisation | box | √
4.5 | X-ray gonad shields for preterm and neonate | box | √
4.6 | Phototherapy eye shields for preterm and neonate | box | √
4.7 | Sterilised endotracheal tube adaptor with lure lock side port for use with surfactant for ET size 2.0-4.0 | box | √
4.8 | Double lumen umbilical catheter 4Fr and 5fr | box | √
4.9 | neonatal extension tubes with filters liquid infusion lines | box | √
4.1 | Neonatal liquid filters | box | √
4.11 | double lumen octopus with auto flush and non-return valves | box | √
4.12 | Triple lumen octopus with auto flush and non-return valves | box | √
4.13 | Anti UV lure lock infusion lines (Lectroguard) | box | √
4.14 | Cap/bonnet for neonate | box | √
4.15 | Cap/bonnet for preterm | box | √

## pharmacy
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Pharmacy Department
1 | Equipments
1.1 | Pharmacy Refrigerator: Temperature range display: +2° C to +8° C LED display should be easy to read for precise control. Capacity, Shelves & Doors: 350-500 ltrs, 5 shelves and 2-4 doors with adjustable wire shelves and sliding wire baskets, doors should be transparent. Dimension and design: As per capacity and design with chip and rust resistant paint finish and should be easy to clean and 100% stainless steel. Alarm & safety device: Both high and low temperature alarm, refrigerator or heating system should shut off if temperature drops below 2° C or rise above 20° C with keyed door lock. Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. Suitable voltage stabilizer. . Certification: CE certification with minimum of 1 year warranty. |  | √ | √
1.2 | Weighing balance for extemporaneous compounding: Capacity: 120-300 grams, Sensitivity: 100 mg, Linearity: ± 0.1 g, Operation temperature range: -10° C to +40° C, Display: LCD, Calibration: Internal, Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. stainless steel weighing pan. | Set | √ | √
1.3 | Digital weighing scale – 500 mg to 5 kg for compounding | Piece | √ | √
1.4 | Pill counting tray: Size of the tray: 20 cm x 14.5 cm x 3.8 cm, material of the set: Plastic (ABS + PS), Material of spatula: Wooden handle + stainless steel. | Set | √ | √
1.5 | Spatula, ointment, metal 12 inch SS | Piece | √ | √ | √ | √
1.6 | Ceramic Mortar and pestle - size 10 inch: Material: Ceramic, Colour: white | Set | √ | √ | √
1.7 | Ceramic Mortar & pestle - size 4 inch: Material Ceramic, colour: white | Set | √ | √ | √
1.8 | Tablet counting machine: Portable table top automatic tablet counting machine; Counts coated and uncoated tablets, hard and soft gelatin capsules and dragees; Digitally controlled count adjustment; Accurate counting and dispensing (100%) with excellent repeatability; Tool free removal and change of product contact parts. Weight of the machine: 3-10 kgs. Dimensions of the machine: 30-40 cm height x 30-40 cm Length x 15-20 cm width. Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. Throughput: 15-18 tablets/seconds. Product size: Should enable to count product of dimension 3-20 mm diameter and 2.5-20 mm thickness. Build standard: Should comply with cGMP or equivalent standards. CE marked in accordance with machinery. Operation: Quiet operation (should not have a noise level of more than 80 dBA); Simple operations with fast product change over. Cleaning procedure: The machine should offer easy access for cleaning. Count capacity: Should enable to count up to 9,999. Product contact parts: All product contact parts should be made of SS 316 or better. Partial pill indicator: Alerts when an object <25% the size of the tablet/capsule is detected. Spares and accessories: Should include all the required parts for product change over, spares & consumables and accessories such as air compressor, hoses, etc. | Set | √
1.9 | Wash Room Trolley SS | Piece | √ | √ | √
1.1 | Pharmacy shelve Dimension-2080x1000x330mm 5shelves with adjustable feet Unit should be assembled with the minimum requirement of tools Shelving config are easily adjusted to suit stocking requirements Ticket railing to allow clear product identification label | set | √ | √ | √ | √
1.11 | Ointment Blender | Piece | √ | √ | √
1.12 | Laminor work flow bench | Set | √
1.13 | Label printer | Pcs | √
1.14 | Hot air oven | Set | √
1.15 | Autoclave | Set | √
1.16 | Digital Hygro-thermometer: displays temperature and relative humidity; wall mountable; Temperature range: -10 to 60°C; Humidity range: 10% to 99% | Pcs | √ | √ | √ | √
2 | Consumables
2.1 | Bottle for dispensing extemporaneous preparations: Amber colour PET bottles with child resistant screw cap and 100 ml capacity. | Piece | √ | √
2.2 | Personal protective equipment kit: Gown with cuffs and back closure (should be water impermeable fabric). Pair of gloves (specially designed for protection). Pair of goggles. Pair of shoe covers (disposable). Mask (disposable). Absorbable towel | Set | √
2.3 | Cytotoxic Drugs spill kit: Gown with cuffs and back closure (should be water impermeable fabric). Pair of gloves (specially designed for protection). Pair of goggles. Pair of shoe covers (disposable). Mask (disposable). Mask (N-95), Absorbable towel, disposable sponges, puncture-proof, leak proof container, container of desiccant powder or granules to absorb wet contents, disposable dust pan, plastic scraper. | Set | √
2.4 | label paper for Label Printer | Pcs | √
2.5 | Envelope, plastic dispeensing 10cm X 12cm | Pkt of 1000 | √ | √ | √ | √

## physio
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Physiotherapy Department
1 | Equipments
1.1 | Traction Machine for Cervical & Lumbar spine with Versatility, Tension Main Timer, Hold Timer, Patient control switch, Traction pilot light Safety Fuse with table and cervical and Lumbar harness, Thoracic Harness, Spreader bar, Pillow and triangular flexion stool. Power cord with round 3 pin top. Appropriate voltage stabilizer should be provided with the unit. a. Cervical & Lumbar Traction Machine Cables (chains). b. Cervical & Lumbar Traction Machine: Cervical. Environmental operating condition: 2Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. Suitable voltage stabilizer. RH >80%. | Set | √ | √ | √
1.2 | Electric muscle stimulator: Diagnostic cum therapeutic unit, suitable for taking S.D curve, producing different types of stimulating currents i.e Galvanic and faradic current . The stimulating current outputs are: Galvanic - continuous DC upto 135 volts for a maximum of 0-75mA load. Interupted Galvanic - Pulse durationof 0.01, 0.03, 0.1, 0.3, 1, 3, 10, 30, 100, 300m sec. Faradic - 80 pulse per second duration with 0.6 ms width. Accessories - Inactive and active plate electrodes, ball tip electrode with handle, 2 pair electrode wires. Environmental operating condition: Relative humidity >80%. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. Suitable voltage stabilizer (delete). | Set | √ | √ | √
1.3 | Hot Pack Machine: Stainless steel with on/off switch and thermostatic control. Preset normal operating temperature control, Preset normal operating temperature control. a. Hot Pack Contour cervical Pack with cover. b. Hot Pack Half size Pack with cover. Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. RH >80%. | Set | √ | √ | √
1.4 | Infrared Lamp with watt range from 150 - 500 watts,Automatic timer, thread type bulb, . Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. RH >80%. | Set | √ | √ | √
1.5 | Interferential Machine: Two oscillators of medium frequency are generated and cross- applied to the body by means of four electrode pads/self adhesive electrodes. Oscillator frequency fixed to 4000 Hz and second can be ranged to 4000 – 4150 Hz. a. Interferential Machine cable to the pads. b. Interferential Machine pad-large with cover/self adhesive electrodes. c. Inferential machine pad -medium with cover/self adhesive electrodes. d. Inferential machine pad small with cover/self adhesive electrodes. e. interferential machine switch cable. Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin type plug top. >80%. | Set | √ | √ | √
1.6 | Short Wave Diathermy SWD: High power short wave Diathermy with patient’s switch & internal cooling system along with additional voltage stabiliser 450 watts. Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. RH >80%. | Set | √ | √ | √
1.7 | Therapeutic LASER: Mid laser used by physiotherapist. Helium-neon laser producing red light at 630nm combined with infrared laser at 904nm, or other physiotherapy laser. Environmental operating condition: RH >80%. Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √
1.8 | Therapeutic Ultra Sound Machine: Ultra sonic Deep Therapy Unit with complete set including Transducer.1-3 MHz. Environmental operating condition: RH >80%. Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √ | √
1.9 | Pocket Transcutaneous Electrical Nerve Stimulator (TENS): Small home stimulation pack (dual channel) with wire attachments, electrodes, 9 V batteries. a. Transcutaneous Electrical Nerve Stimulator (TENS): 9V battery. b. Transcutaneous Electrical Nerve Stimulator. 1. Dual channel: individually isolated circuits. 2. Amplitude: 0 mA - 80 mA; indication only; actual mA will tend to be less than indicated due to electrode impedance. 3. Type: Constant Current. 4. Waveform: Asymmetrical, rectangular bi-phasic with zero DC current. 5. Selectable pulse width: 50 µS - 300 µS [2% accuracy]. 6. Pulse Rate selection: in the continuous mode 2 Hz – 200 Hz [2% accuracy]. 7. Mode: Continuous, Burst or Modulated. 8. Burst mode: Bursts of 9 pulses [200 µS] at 150 Hz, repeating twice every second. 9. Modulation mode: 6-second cycle of concurrent width modulation and pulse repetition rate modulation. Width starting at 200 µS and decreasing exponentially to 100 µS in three seconds and then returning back to 200 µS in the next three seconds. Rate starting at 100 Hz, decreasing exponentially to 65 Hz and then returning to 100 Hz. 10. Time duration of the treatment selectable: 1 minute to 12 hours. 11. Low Battery Indicator: If the battery goes below 6.9 volts +/- 0.2 volts the battery symbol will flash on/off once every second. 12. Environmental conditions for storage & transport: -10 to +40 degrees Centigrade>80% Humidity. 13. Self adhesive electrodes.Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √ | √
1.1 | Adaptive sitting chair: A sturdy and comfortable seating system, the High Low Chair (for children up to 45" in height) is a combination high chair and mobile floor sitter for children. Comfortable, Secure Positioning. Easily converts to a mobile floor sitter by detaching the upper frame from the pedestal base. Adjusts in height with the use of a single lever that activates a pneumatic cylinder. The mobile floor sitter frame is constructed of 1" powder coated tubular steel with locking casters. The cushioned seat is covered with a water proof, easy to clean polymer coated fabric. The durable melamine tray remains securely in place with a quick release locking lever. Include H-strap harness, washable cotton quilted trunk support, removable abductor and tray. An accessory head support and lateral supports are available. The high low chair, seat depth: 7" - 10", seat back height: 22", seat width: 12", High chair seat height: 19", Floor sitter seat height: 24", Tilt in space angle: 15°, 20°, 25°, Weight of seat and frame: 46 lbs, Weight capacity of Hi Low chair:75 lbs, Wheel diameter: 2", Foot rest dimensions: 6" x 12", Tray dimensions: 12" x 18". Input voltage 220-230V, Power cord with round 3 pin top. | Set | √ | √
1.11 | Assorted square & round pegboard: wooden board with 20 multi colour wooden pegs (10 round & 10 square) in four rows. Peg size - 76 x 16 mm | Set | √ | √
1.12 | Corner Chairs: Provide support for symmetrical posture and mid-line control of head and trunk and improve breathing for children who have low trunk tone. Prevents scapular and shoulder retraction, a pattern frequently seen in children who cannot stabilize against gravity. Age/Size: 8 months - 4 years, 4 years - 8 years. Height of chair with head supports, adjusts from 51 cm - 69 cm, adjust from 74 cm - 104 cm. Removable tray: 38 cm x 43 cm, 46 cm x 51 cm. Seat depth with insert: 20 cm, 26 cm. Seat depth without insert: 30 cm, 36 cm. Removable abductor: 8 cm wide, 10 cm wide. Chair with 4" swivel, locking casters. | Set | √ | √
1.13 | Delux revolving hole punch, length 22 cm, throat depth 3.8 cm with replacement tubes of 2.4 mm tube, 2.8 mm, 3.2 mm, 4 mm, 4.4 mm, & 4.8 mm | Piece | √ | √
1.14 | Dumbell of Stainless steel with weight of 1 kg, 1.5 kg, 2 kg, 2.5 kg, 3 kg, 3.5 kg, 4 kg, 4.5 kg, 5 kg (one pair each) in a set with stand frame. | Set | √ | √ | √
1.15 | Treadmill: 0.1mph to 11 mph adjustable speed (0.1 mph increments). Incline range from 0 -15 % in 1% increments. Support bars on both sides. Surface area - 51 x 140 cm, safely stop button, hand rails, weight capacity 400Ibs. Environmental operating condition: RH >80%. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. | Set | √ | √ | √
1.16 | EMG bio-feedback, pressure feedback, electrotherapy and electro diagnostics. The myofeedback unit is provided with separate part for myofeedback, pressure feedback and combination with electrotherapy. Several surface and cavity electrodes can be connected. This makes the unit extremely suitable for all kinds of exercise and relaxation therapies including Pelvic floor re-education. Environmental operating condition: RH >80%. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. | Set | √
1.17 | Heat gun: A free standing, light weight heat gun for use with all types of thermoplast with 2 hot air settings 300°C - 500°C. Weight 550 gm. Environmental operating condition: RH >80%, Voltage: AC 210-240V, 50Hz . | Set | √ | √
1.18 | Magnetic Field Therapy Unit, MG-Wave is a portable unit for therapy with low-frequency pulsating electromagnetic field. Specification: Field Pilot Signal: rectangular. Pulse Duration: Upto 5ms, Frequency: 1 + 750Hz, Max B Induction: 150 gauss (peak) (± 15%) on Piece coil, Timer: 0+240min, end-of-treatment acoustic signal and automatic interruption of operation. Power Supply: 100-200 Vac, 50-60Hz wall-mounted power supply. Input: 75VA, Safety Class: II- BF type (CEI EN 60601-1). Environmental operating condition: Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. RH >80%. | Set | √
1.19 | Treatment Couches 6-Section (Manumed Optimal) : Hydraulic high-low/Electric high-low, head section, flexion position, separate leg section, table top size 280 x 67cm, height adjustment 48-98cm. | Set | √ | √
1.2 | Treatment Couches 3 - Section (Manumed Optimal): Hydraulic high-low/Electric high-low, table top size 280 x 67cm, height adjustment 48-98cm. | Set | √ | √
1.21 | Mirror (with wheels) for therapeutic purpose - size (27 inch W x72 inch H) | Piece | √ | √ | √
1.22 | Parallel Bars adjustable height and width by crank arm 3 m long. with base included - Paediatric | Set | √ | √
1.23 | Parallel Bars adjustable height and width by crank arm 4.5 m long. with base included - Adult | Set | √ | √ | √
1.24 | Paediatric positioner: Half roll with two lateral supports. Wedge, Log roll, Tray with mirror for creative positioning, Carrying bag, Modules allow for a variety of creative positioning, Three positioning belts help to secure child, Soft covering is easy to clean and is nontoxic. Size: 34"x 19" (86 cm x 48 cm) base | Set | √ | √
1.25 | Recumbent Cycle Trainer: Adjustable seat height and distance with back rest and adjustable resistance. Electric or battery operated display to show RPM, time, resistance level, distance, speed. | Set | √ | √ | √
1.26 | Shape sorting cube with 10 multi colour wood blocks: size of box - 5.5 inch L x 5.5 inch W x 5.5 inch D. | Set | √ | √
1.27 | Compact Shoulder Wheel | Piece | √ | √ | √
1.28 | Stacking cones with sequencing & colour identity, size - 4.5 inch H x 2.5 inch diameter | Set | √ | √ | √
1.29 | Standing frame - Adult | Set | √ | √ | √
1.3 | Standing frame - Paediatric | Set | √ | √
1.31 | Tilt Table: Adult size, manual/electric, dimension 78" lengthX 28" widthx 33" height, weight capacity 230 lbs with angles from 0-90 degrees | Set | √ | √ | √
1.32 | Wire cutter, heavy duty, length 22 cm | Piece | √ | √
1.33 | Crank adjustable Hi-Lo Mat platform. 183cm x 244 cm | Set | √ | √ | √
1.34 | Therapeutic Exercise ball , red: Diameter: 20-100cm Should come with inflation valve plug remover. | Piece | √ | √
1.35 | Therapeutic Exercise ball, blue: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
1.36 | Therapeutic Exercise ball, green: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
1.37 | Therapeutic Exercise ball, orange: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
1.38 | Therapeutic Exercise ball, yellow: Diameter: 20-100cm Should come with inflation valve plug remover | Piece | √ | √
1.39 | Peanut Therapeutic Exercise ball, Diameter:40-80 cm | Piece | √ | √
1.4 | Two-way Air pump: Manually operated pump inflates balls, toys air mattresses etc on up and down strokes cutting effort in half. | Piece | √ | √
1.41 | Trampoline: With an easy grip handle, Suitable for both indoor and outdoor use, with tough weatherproof mat and tubular steel frame with a durable coated finish, padded cover around the mat. The handle simply screws off and the legs fold down for ease of storage. Weight: 8 Kg, from handle to floor: 32.25", dimensions: 33.75" x 33.75", maximum weight: supports up to 68Kgs. | Set | √ | √
1.42 | Cushion: wedge shaped with firm, wipe down cover, two wrap straps, one abductor module for pediatric, size: Height 4-12 inch, Length:22-28 inch, width:20-24 inch. | Set | √ | √
1.43 | Cushion: Bolster shaped with firm, wipe down cover, cylindrical shape, diameter: 8-18 inch, Length: 36-48 inch | Set | √ | √
2 | Consumables
2.1 | Self adhesive electrodes for TENS and Interferential Machine | Packet | √ | √ | √
2.2 | Cold pack 23 inch neck contour with cover(It should be compatible for Cold pack machine) | Piece | √ | √ | √
2.3 | Cold pack Oversize with cover (It should be compatible for Cold pack machine). | Piece | √ | √ | √
2.4 | Cold pack Standard size with cover (It should be compatible for Cold pack machine). | Piece | √ | √ | √
2.5 | Electric muscle stimulator Electrodes - medium (It should be compatible for Electric Muscle Stimulator machine) | Piece | √ | √ | √
2.6 | Electric muscle stimulator Electrodes - small (It should be compatible for Electric Muscle Stimulator machine) | Piece | √ | √ | √
2.7 | Finger Hooks, 6.4 mm, steel used in splinting (Packet of 100) | Pkt | √ | √
2.8 | Hot Pack Contour cervical Pack with cover (It should be compatible for Hot Pack Machine) | Piece | √ | √ | √
2.9 | Hot Pack Oversize Pack with cover (It should be compatible for Hot Pack Machine) | Piece | √ | √ | √
2.1 | Hot Pack Standard size Pack with cover(It should be compatible for Hot Pack Machine) | Piece | √ | √ | √
2.11 | Interferential Machine Electrodes - medium (It should be compatible for Interferential machine). | Piece | √ | √ | √
2.12 | Interferential Machine Electrodes - small (It should be compatible for Interferential machine). | Piece | √ | √ | √
2.13 | Interferential Machine Electrodes cover – medium (It should be compatible for Interferential machine). | Piece | √ | √ | √
2.14 | Interferential Machine Electrodes cover- small (It should be compatible for Interferential machine). | Piece | √ | √ | √
2.15 | Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Yellow: Light) | Piece | √ | √ | √
2.16 | Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Red: Medium) | Piece | √ | √ | √
2.17 | Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Green: Heavy) | Piece | √ | √ | √
2.18 | Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Blue: Extra Heavy) | Piece | √ | √ | √
2.19 | Exercise Band: Natural rubber latex resistive exerciser, color-coded levels of progressive resistance. Size: 15 cmx5.5 m long (Black:2X Heavy) | Piece | √ | √ | √
2.2 | Air cushion: Providing additional shock and vibration absorption provides a more pain relieving ability and a more stable seating environment. Inflates by mouth or with a small portable hand air pump that comes with the cushion. Standard size: 22" x 22"/56 cm x 56cm, Weight: 1 Kg, Cell Height: 2"/5cm fully inflated (due to the required immersion of the user in the cushion, it will raise the user only (1/2")/(1 cm) above the seat), Cushion material: Neoprene rubber, Cover material: Polyester blend with a non-skid polyester and vinyl base, Water Resistance: Cushion and cover are water resistant. | Set | √ | √
2.21 | Gutter Splint/Gaiter: Large | Piece | √ | √ | √
2.22 | Gutter Splint/Gaiter: Medium | Piece | √ | √ | √
2.23 | Gutter Splint/Gaiter: Small | Piece | √ | √ | √
2.24 | Gutter Splint/Gaiter: Pediatric | Piece | √ | √ | √
3 | Consumables - (Mobility Aids)
3.1 | Adjustable Axillary Crutch (Aluminum) - Large | Pair | √ | √ | √ | √
3.2 | Adjustable Axillary Crutch (Aluminum) - Medium | Pair | √ | √ | √ | √
3.3 | Adjustable Axillary Crutch (Aluminum) - Small | Pair | √ | √ | √ | √
3.4 | Adjustable Elbow Crutch (Alumunium) - adult | Pair | √ | √ | √ | √
3.5 | Adjustable Elbow Crutch (Alumunium) - Small | Pair | √ | √ | √ | √
3.6 | Elbow Crutch Rubber Tips | Pair | √ | √ | √
3.7 | Walker Rubber Tips | Pair | √ | √ | √
3.8 | Adjustable Walker with wheels (Alumunium) - adult | Piece | √ | √ | √
3.9 | Adjustable Walker with wheels (Alumunium)- Small | Piece | √ | √ | √
3.1 | Adjustable Walker without wheels (Alumunium) - adult | Piece | √ | √ | √
3.11 | Adjustable Walker without wheels (Alumunium) - Small | Piece | √ | √ | √
3.12 | Adjustable Platform Crutch (Aluminum) - Large | Pair | √ | √
3.13 | Adjustable Platform Crutch (Aluminum) - Medium | Pair | √ | √
3.14 | Adjustable Platform Crutch (Aluminum) - Small | Pair | √ | √
3.15 | Adjustable Platform Walker with wheels (Alumunium) - Large | Piece | √ | √
3.16 | Adjustable Platform Walker with wheels (Alumunium) - Medium | Piece | √ | √
3.17 | Adjustable Platform Walker with wheels (Alumunium) - Small | Piece | √ | √
3.18 | Adjustable Platform Walker without wheels (Alumunium) - Large | Piece | √ | √
3.19 | Adjustable Platform Walker without wheels (Alumunium) - Medium | Piece | √ | √
3.2 | Adjustable Platform Walker without wheels (Alumunium) - Small | Piece | √ | √
3.21 | Adjustable Posterior Walker with wheels (Alumunium) - Adult/ Two fixed direction 5 inch front wheels & two fixed direction 3 inch. One way ratchet rear wheels. | Piece | √ | √
3.22 | Adjustable Posterior Walker with wheels (Alumunium) - Pediatric/ Two fixed direction 5 inch front wheels & two fixed direction 3 inch. One way ratchet rear wheels. | Piece | √ | √
3.23 | Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 12.5" (inches); Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.24 | Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 14"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set
3.25 | Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 15.5"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.26 | Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 17"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.27 | Four Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Four wheel, Cross folding frame with Seat Upholstery (removable & washable) and tension adjustable straps, Back support/Upholstery (removable & washable) and tension adjustable straps, Push handle, Arm support, Foot support, Pressure-relief Cushion with waterproof, thin and stretchy cushion cover (washable), 2 large rear wheels with pneumatic tyres, 2 small front casters, and removable & length-adjustable anti-tip levers. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 18.5"; Seat dephts range: 14", 16", 18"; Seat angle (tilt): 5 - 12 degrees; Back height: adjustable at 14'', 16", 18"; Foot rest height (below seat): adjustable at 12" - 17"; (Upper) back support angle: adjustable at/or 8 - 10 degrees; Wheel camber: 3 degrees; Rear axle position: multiple; Caster Size: 4.5” x 3”; Seat sag (maximum): 2cm; Back sag: 1cm; Push rim clearance: 2cm. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.28 | Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (Seat width): 12 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.29 | Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.3 | Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims attached to the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.31 | Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims on the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.32 | Three Wheel Manual wheelchair (Basic) Manual wheelchair, User propelled, Long wheelbase, Three wheel, Rigid frame with Solid Seatbase, Push handle, Removable Armrest, height adjustable and foldable Back support, Pressure-relief Cushion with waterproof, thin & strtechy cushion cover, Large front caster wheels with long wheelbase, height and angle adjustable footrest, adjustable axle position, Pneumatic rear wheel tyres, Removable rear wheels with a quick release mechanism and folding backrest for easy transport and storage. Push rims on the rear wheels. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 20 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.33 | Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 12 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.34 | Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.35 | Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.36 | Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.37 | Four Wheel Heavy Duty Wheelchair Manual wheelchair, User propelled, Heavy duty, Four wheel, Cross folding frame. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Curved armrest. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Height adjustable push handles. Spoked & removable rear wheels with a durable quick-release mechanism and foldable backrest for ease of storage and transportation. Push rims attached to the rear wheels. Sturdy and medium sized castor wheels to help manoeuvre over uneven terrain. Tension adjustable straps to provide good postural support. Strong and pneumatic rear wheel tyers. Choice of safe or active rear wheel positions. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 20 inches. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.38 | Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 12 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.39 | Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.4 | Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.41 | Four Wheel Light Rigid Wheelchair Manual wheelchair, User propelled, Light weight, Four wheel, Rigid frame. Light and easy manoeuvrability. Adjustable seat depth, footrest height and angle, backrest height and angle to provide a customised fit. Tension adjustable back straps. Contoured pressure relief cushion with waterproof, thin & stretchy cushion cover. Folding backrest and quick-release rear wheels and castor wheels for quick and easy storage and transportation. Spoked rear wheels. Push rims attached to the rear wheels. Small castor wheels and 24” rear wheels to help negotiate small spaces. Multiple rear wheel axle positions. Bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A pressure relief cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.42 | Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims; Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 12 inches (for children) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.43 | Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 14 inches (small adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.44 | Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 16 inches (medium adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.45 | Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 18 inches (standard adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.46 | Four Wheel Transfer Wheelchair Manual wheelchair; Attendant propelled (for use as transfer aids in the hospital/institutions); Crossfolding frame (for storage when not in use); Strong & removable armrests; Detachable, swing away & height adjustable footrest hanger; (Elevating) adjustable angle of footrest hanger (for transferring patients who may not be able to bend their knees); Leg strap: nylon strap (4 to 5 cm wide) with Velcro, the leg strap should have Velcro fitted to its full length to allow maximum adjustability (to rest the leg during tranfer); Foot rest: footplate to be plastic, aluminium or mild steel & height adjustable; Belt for back stability; Seat angle (tilt): 5 - 8 degrees; Back to seat angle: 90 degrees; (Upper) back bend post: 8 - 10 degrees; Standard back height; Gripped push handle; Nylon or vinyl back & seat upholstery; Comfort cushion with water resistant cover; Strong swivel fork and sturdy casters; Standard or spoke rear wheels with strong semi-solid tyres; 12mm thick axle; Smooth push rims. Strong & bushed brakes with adjustable toggle mechanism. Size (seat width): 20 inches (large adult) Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √ | √
3.47 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 8” Seat 200 mm wide x 175-250 mm adjustable depth Seat and frame configuration: If the chair is modular the seat system (combined seat and backrest with optional accessories such as headrest) will be separate from the base frame. The chair or the base frame to be rigid or folding. Infinite adjustable seat system tilt between 0 and minimum 30 degrees by means of a quick release mechanical or hydraulic option. The design should allow for COG adjustment either as a specific wheel/seat adjustment or by means of a mechanism which will always result in COG being safely within base of support. The design of the COG adjustment should allow for appropriate access to the rear wheel. Seat module: The seat to back angle to be adjustable between 90° and 140° with a minimum of 5 settings.The design should always allow the foot rests to be in line with the front of the seat after adjustment of seat depth. Seat base must provide rigid support. The seat will have a length adjustable pelvic strap at 60°- 90° angle over hips. Width of pelvic strap for 12” and smaller = 30mm. Width of pelvic strap for 14” and bigger = 50mm. 12” short and smaller sizes should allow for age appropriate hip abduction. Back post / Push handle requirements: Top of Push handles must be a minimum 93.5cm from ground to top of push handles with 10cm height adjustability. Push handle may in no way restrict self-propelling, affect posture or set-up of back support system. Armrests: Armrests supplied as standard for chairs larger than 12” wide but can be removed as an option. Must be height adjustable with a minimum range of 100mm with 15mm increments (or infinitely adjustable) with the lowest height being 150mm from the top of the seat. Must be adjustable forward/back and must allow for support directly below the elbows. Footrest: Height adjustable hanger with minimum 150 mm range (100mm on 12” chair and smaller) to allow for variations in user profile and seat length adjustments. Must flip-up / fold up or must be detachable (not required on chairs smaller than 12”). Adequate rigidity and must have a non-slip surface. Tray-table: Lap tray to be supplied as standard for chairs smaller than 12” wide. Lap tray Optional device for chairs over 12”. Fitting for tray table to be provided on frame as standard. See specifications below under options. Head support: Standard accessory. Adjustable forward/back and up/down with a range of at least 100mm and 200mm respectively. Front castors and forks: Front castor wheel diameter range: 100 mm and 200 mm (+2mm). Front castor wheel width range: 25 mm and 80 mm. Stem to be fastened to the fork with thread and nut system to allow for re-fastening and adjusting. Rear wheels and axles: Size 12” standard chair and larger: minimum 24” x 1 3/8 diameter with one part semi-solid tyre. Size 12” short chair and smaller: minimum 20” x 1 3/8 diameter with one part semi-solid tyre or foam insert. Optional pneumatic tyre. To be fitted with quick-release axle. Anti-tip wheels: Must be removable. Should allow for adjustability to match the level of independence and balance of the user. Range of adjustability to allow for stability in static position and to adjust for the required degree of stability in propulsion and if applicable, when front castors are raised for obstacle clearance. Weight: Not to exceed 18 kg (complete, but without back system, tray table or head extension). 12” and under not to exceed 14 kg (complete, but without back system, tray table, or head support/extension). Repair / tool kit: All tools required to adjust frame and back rest systems, as well as foot rest system will be supplied with the chair. Back system: Tension adjustable backrest with rigid, height adjustable full length laterals to fit both basic and adjustable folding frame wheelchairs. Standard back height 42cm for chair sizes above 12" and 30cm for 12" & smaller chairs. Tall back height 50cm for chair sizes above 12" and 42cm for 12" & smaller chairs. Lateral depth: 20cm for 18" & 20" chairs, 15cm for 14" & 16" chairs, and 10cm for 12" chairs. Each wheelchair unit submitted and ordered must be supplied with the following: o Any necessary tools for assembly and user maintenance with each unit. o A comfort cushion with specified cushion cover o A printed user manual o A printed assembly manual o Pricing and ordering instructions for common spare parts All wheelchairs submitted for review must include documentation from manufacturer including product specifications, parts lists, accessory list, product catalogue, assembly manual, user manual relative to that product or product line. | Set | √ | √
3.48 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 10” Seat 250 mm wide x 250-350 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
3.49 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 12” short Seat 300 mm wide x 250-320 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
3.5 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 12” standard Seat 300 mm wide x 300-400 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
3.51 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 14” short Seat 350 mm wide x 250-350 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
3.52 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 14” standard Seat 350 mm wide x 350-450 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
3.53 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 16” standard Seat 400 mm wide x 350-450 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
3.54 | Adjustable posture support chairs, attendant/ self-propelling, with tension adjustable back with rigid, height adjustable full length laterals. Size: 18” standard Seat 450 mm wide x 350-450 mm adjustable depth. Specifications described in No. 3.47 | Set | √ | √
4 | Consumables - (Supportive)
4.1 | Strapping and Taping Tape: Breathable porous athletic tape protects joints from sprains and injuries. Provides support to help speed recovery from injuries, high strength, versatile tape for a wide range of uses. Easy to tear, no scissors necessary. Size: 3.8 cm x 13.7 - 32 per case | Roll | √ | √
4.2 | Strapping and Taping Tape: Breathable porous athletic tape protects joints from sprains and injuries. Provides support to help speed recovery from injuries, high strength, versatile tape for a wide range of uses. Easy to tear, no scissors necessary. Size: 5 cm x 13.7 - 24 per case | Roll | √ | √
4.3 | Ankle Support/brace: Extra strong highly porous orthopedic webbing, containing heat resistant rubber with high modulus of elasticity. Hook loop closure - Large | Piece | √ | √ | √
4.4 | Ankle Support/brace: Extra strong highly porous orthopedic webbing, containing heat resistant rubber with high modulus of elasticity. Hook loop closure - Medium | Piece | √ | √ | √
4.5 | Ankle Support/brace: Extra strong highly porous orthopedic webbing, containing heat resistant rubber with high modulus of elasticity. Hook loop closure - Small | Piece | √ | √ | √
4.6 | Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - below 26 cm (child size) | Piece | √ | √ | √
4.7 | Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - 26-30 cm (small) | Piece | √ | √ | √
4.8 | Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - 30-34 cm (medium) | Piece | √ | √ | √
4.9 | Pouch arm sling with ring and straps around the neck support and stabilization of shoulder joint. For shoulder dislocation and subluxation, prevention and treatment - 34-38 cm (Large) | Piece | √ | √ | √
4.1 | Back supports: 80 cm | Piece | √ | √ | √
4.11 | Back supports: 85 cm | Piece | √ | √ | √
4.12 | Back supports: 90 cm | Piece | √ | √ | √
4.13 | Back supports: 95 cm | Piece | √ | √ | √
4.14 | Back supports: 100 cm | Piece | √ | √ | √
4.15 | Back supports: 105 cm | Piece | √ | √ | √
4.16 | Back supports: 110 cm | Piece | √ | √ | √
4.17 | Soft Cervical collar, high density polyurethane foam cover with cotton stockknitte, plastic bushes to provide enough ventilation and Velcro closures for easy application, removal and adjustment, polythylene sheet reinforcement for better immobilization: size neck circumference in cm 38-42 (large) | Piece | √ | √ | √
4.18 | Soft Cervical collar, high density polyurethane foam cover with cotton stockknitte, plastic bushes to provide enough ventilation and Velcro closures for easy application, removal and adjustment, polythylene sheet reinforcement for better immobilization: size neck circumference in cm 34-38 (medium) | Piece | √ | √ | √
4.19 | Soft Cervical collar, high density polyurethane foam cover with cotton stockknitte, plastic bushes to provide enough ventilation and Velcro closures for easy application, removal and adjustment, polythylene sheet reinforcement for better immobilization: size neck circumference in cm 30-34 (small) | Piece | √ | √ | √
4.2 | Crepe Bandage - 5 cm x 450 cm (4.5 m stretch length) | Roll | √ | √ | √
4.21 | Crepe Bandage - 10 cm x 450 cm (4.5 m stretch length) | Roll | √ | √ | √
4.22 | Crepe Bandage - 15 cm x 450 cm (4.5 m stretch length) | Roll | √ | √ | √
4.23 | Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (Size: small - 12" - 14") | Piece | √ | √ | √
4.24 | Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (Size: Medium - 14" – 15" | Piece | √ | √ | √
4.25 | Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (Size: Large - 15"-17") | Piece | √ | √ | √
4.26 | Knee Cap/Brace: Elastic construction provides effective joint control for weak or injured knees. Aids stiff, swollen and painful knees. Ideal for sports, works, or home. Feature 11" long thermal neoprene sleeve with nylon facing on both sides and 4-way stretch. Heavy duty outside for durability and is contoured and darted for optimum fit. Retains warmth, promotes healing, and reduces abrasions. Fits right or left knee. Measure around knee. (SizeX-Large: 17" - 20") | Piece | √ | √ | √
4.27 | Silicon Polymer Heel Pad/cup/Shoe Lift : shoe size 9 and 10 (Large) | Pair | √ | √
4.28 | Silicon Polymer Heel Pad/cup/Shoe Lift : shoe size 7 and 8 (Medium) | Pair | √ | √
4.29 | Silicon Polymer Heel Pad/cup/Shoe Lift : shoe size 5 and 6 (Small) | Pair | √ | √
4.3 | Top Layer Foam: High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m3. Size range: 46 cm wide; 43 cm long; 5cm high. 41 cm wide; 43 cm long; 5cm high. 38 cm wide; 43 cm long; 5cm high. 36 cm wide; 43 cm long; 5cm high. 30 cm wide; 43 cm long; 5cm high. Variance allowed on overall measurement for cushions: +1cm. The entire size range will remain one series. | Series | √ | √ | √
4.31 | Top Layer Foam: High-density foam 40/24; Hardness factor 21.6-25.2; Density range 36.1-40kg/m3. Size range: 61 cm wide; 43 cm long; 5cm high.  56 cm wide; 43 cm long; 5cm high.  51 cm wide; 43 cm long; 5cm high.. Variance allowed on overall measurement for cushions: +1cm. The entire size range will remain one series. | Series | √ | √ | √
4.32 | Base layer Foam: High density compressed (reconstituted flexible PU) foam 1300; Hardness factor –52-63; Density 130+10% kg/m3. Size range:  61 cm wide; 43 cm long; 5cm high  56 cm wide; 43 cm long; 5cm high  51 cm wide; 43 cm long; 5cm high  46 cm wide; 43 cm long; 5cm high  41 cm wide; 43 cm long; 5cm high  38 cm wide; 43 cm long; 5cm high  36 cm wide; 43 cm long; 5cm high  30 cm wide; 43 cm long; 5cm high Variance allowed on overall measurement for cushions: +1cm The entire size range will remain one series. | Series | √ | √ | √
4.33 | Incentive Spirometer: 500 - 4000 ml capacity Latex + DNR free. | Piece | √ | √
4.34 | Moleskin Padding with self adhesive backing, roll of 0.8 mm x 300 mm x 4.6 m | Roll | √ | √
4.35 | Splinting material (Thermoplast/ Aquaplast/ EZEform): Perforated, thickness 2.4 mm, size - 61 x 91 cm | Sheet | √ | √
4.36 | Splinting material (Thermoplast/ Aquaplast/ EZEform): Perforated, thickness 3.2 mm, size - 61 x 91 cm | Sheet | √ | √
4.37 | Splinting material (Thermoplast/ Aquaplast/ EZEform): Solid, thickness 2.4 mm, size - 61 x 91 cm | Sheet | √ | √
4.38 | Splinting material (Thermoplast/ Aquaplast/ EZEform): Solid, thickness 3.2 mm, size - 61 x 91 cm | Sheet | √
4.39 | Tennis Elbow Strap - Large (30-35 cm) | Piece | √
4.4 | Tennis Elbow Strap - Medium (25-30 cm) | Piece | √
4.41 | Tennis Elbow Strap - Small (18-25 cm) | Piece | √
4.42 | Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Extra Large (55-61 cm) | Pair | √ | √
4.43 | Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Large (49-55 cm) | Pair | √ | √
4.44 | Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Medium (43-49 cm) | Pair | √ | √
4.45 | Varicose vein support stocking: Measures circumference around mid thigh-approximate 6 inches above. Small (37-43 cm) | Pair | √ | √
4.46 | Velcro loop, size - 25 mm, reel 25 meter | Roll | √ | √
4.47 | Velcro self adhesive hook, size - 25 mm, reel 25 meter | Roll | √ | √
4.48 | Wrist brace - strong porous, orthopaedic webbing. Double hook loop locking system enhances the immobilization and provides a firm grip to the wrist. Size 20.0-23.8 cm (Large) | Piece | √ | √
4.49 | Wrist brace - strong porous, orthopaedic webbing. Double hook loop locking system enhances the immobilization and provides a firm grip to the wrist. Size 16.3-20.0 cm (Medium) | Piece | √ | √
4.5 | Wrist brace - strong porous, orthopaedic webbing. Double hook loop locking system enhances the immobilization and provides a firm grip to the wrist. Size 12.5-16.3 cm (Small) | Piece | √ | √
4.51 | Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.25 x 75mm; Packet of 100 needles. | Packet | √
4.52 | Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.30 x 50mm; Packet of 100 needles. | Packet | √
4.53 | Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.20 x 30mm; Packet of 100 needles. | Packet | √
4.54 | Accupuncture/Dry Needles: accupuncture/dry needle with guide tube; sterilized & disposable; size: 0.25 x 25mm; Packet of 100 needles. | Packet | √

## Radiology
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Radiology Department
1 | X-Ray Unit
1.1 | X-Ray Machine 300 mA (specification) | set | √ | √ | √ | √
1.2 | X-Ray Machine 600mA (specification) | set | √ | √ | √ | √
1.3 | X-Ray Machine 60 mA, portable (Moltimobile 2.5 )(specification) | set | √ | √ | √ | √
1.4 | Film Auto Processor: Automatic films processors for processing of all types of X-Ray and Imaging Films It should have the following technical features: Should accept film size from 10x10 inches to 14x17 inches, It should process minimum of 60 films per hour both in rapid cycle and standard cycle, The processing time should be variable from 1.5min to 5 min, It should have a developing tank of minimum 10 litres capacity. The container assembly should be made of material, which is non-corrosive, and of latest technology, The main drive should stop automatically during isle time to prevent wear and tear and remain ready in standby, Microprocessor controlled switches and thermostatic control should be there for maintenance of temperature of developer between 30 deg. C and 35 deg, C. The replenishment containers should have capacity of at least 25 liters and the replenishment should be automatic, Wall mounting facility with film output through wall should be possible with the unit, the unit should be quoted with appropriate voltage 220 stabilizer and all essential & optional accessories. Specify the power requirement. | Set | √ | √ | √ |  | where there is X-ray services in DH
1.5 | X- Ray film dryer : 25 films drying capacity, 306/316 SS make for anti corrosion and rust free, power supply- 220V-230V, 50Hz,power cord of round 3 pin plug. MCB controlled and on / off switch. Extra heater to be provided. | set | √ | √ | √ | √
1.6 | Computed Radiography (CR): GENERAL: Drop-and-go cassette buffer: 5 cassettes of mixed sizes input buffer and 5 cassettes of mixed sizes output buffer. Throughput:35 x 43 cm (14 x 17 inch) = approx. 83 plates/hour. Display for status and error indication: • LCD touch screen.• LED status indicator. Greyscale resolution: • Output to processor: 16 bits/pixel square root compressed. Dimensions and weight: • Covered floor space: (W x D x H): 66 x 51 x 123 cm (26 x 20 x 48.4 inch). • Output buffer included: (W x D x H): 115 x 51 x 123 cm (5.3 x 20 x 48.4 inch). • Weight: approx.: 180 kg (397 lbs). Configuration requirements: • NX, • ID tablet, • CR HD5.0 Detectors and Cassettes, • CR MD4.0R Plates and Cassettes. Power: • 220 - 240V/50-60Hz Standby 87W, peak. 590W, fuse 16A. • 120V/60Hz (USA) Standby 92W, peak. 621W, fuse 15A. • 100V/60Hz (Japan) Standby 92W, peak. 621W, fuse 15A. Environmental conditions DX-G digitizer: • Temperature: 15 - 30 ˚C (59 - 86°F). •Humidity: 15 - 75% RH. • EMC compliant with IEC 60601-1-2. • Rate of change of temperature: 0.5°C/minute (0,9 °F). Environmental effects: • Noise level: max. 65 dB (A). •Heat dissipation: standby 92W, continuous operation 242W. SAFETY Approvals: • ETL classified CUS, CE. Transport details: • Temperature: -25 to +55°C (-4 to 131°F), -25°C for max. 72 hours, +55°C for max. 96 hours. • Humidity: 5 - 95% RH. | Set | √ | √ | √
1.7 | Digital Radiography (DR) Machine: 1. Microprocessor controlled high frequency generator (80kW) with Automatic Exposure Control – 1. 2. High speed 600 KHU Rotating Anode X-ray Tube – 1 No. 3. Multi-Leaf collimator – 1 No. 4. Vertical stand for the Digital Detector – 1 No. 5. 3D Ceiling Mounted Tube (CMT) with Servo – 1 No. 6. IDC X3C (3K X 3K) Digital Detector with flat Cesium lodide Scintillator Panel fully integrated with the system – 1 No. 7. Patient table – mobile carbon fiber flat table top with locking wheels – 1 No. 8. Grid – 1 No. A) 180 cm carbon fiber grid (70 Line/cm and ratio of 13:1) – 1 No. B) 100 cm carbon fiber grid (70 Line/cm and a ratio of 13:1) – 1 No. 9. Technologist work station consisting of i) Pentium 4 Duo-processor 3.2 Ghz. ii) 2 GB RAM. iii)160 GB SATA Raud 1 hard Disk drive. iv) 0 Colour flat panel display monitor 1600 x 1200. v) Ethernet 100/1000 Mbps for networking with PACS/RIS/HIS. vi) DV cum CD writer for image exporting. vii) Magellan 3.0 work List software. viii) Built-in Dicom 3.0 enabled functionally. ix)Auto stitching software for special X-rays such as scoliosis and long limbs. Accessories: 10.Konica Dry Laser Camera with 3 online film sizes – 1 No. 11. Voltage stabilizer for the X-ray system – 1 No. 12. Online UPS for the computer and workstation with 30 min. Backup – 1 No. 13.Titanium Apron – 2 Nos. 14. Lead Glass 100 cm x 120 cm 1 No. 15. Book Clarks 100 cm x 120 cm – Each. 16. 1 Box of films for the standard sizes chosen – 1 box for each size. 17. Writable CD/R -1000 Nos. 18. Footrest – 1 No. 19. Hand Grip – 1 No. 20. Abdominal compression binder – 1 No. 21. Footstep – 1 No. 22. Arm rest – 1 No | Set | √ | √
1.8 | Cassette 12 inch x 10 inch | Piece | √ | √ | √
1.9 | Cassette 12 inch x 15 inch | Piece | √ | √ | √
1.1 | Cassette 14 inch x 17 inch | Piece | √ | √ | √
1.11 | Lead goggle | Piece | √ | √
1.12 | Thyroid shield | Piece | √ | √ | √
1.13 | Digital X-Ray film, 14x17 (125 sheets in one packet), compatible with dry view camera) | Packet of 125 sheets | √ | √
1.14 | Digital X-Ray film, 11x14 (125 sheets in one packet), compatible with dry view camera) | Packet of 125 sheets | √ | √
1.15 | Digital X-Ray film, 8x10 (125 sheets in one packet), compatible with dry view camera) | Packet | √ | √
1.16 | Cassette pass box - 2 doors ( when one door opens othecassette to be fit in. 14x17 r shuold be automatically locked)size: | piece | √ | √ | √ | √
1.17 | Chest stand ss | set | √ | √ | √ | √
1.18 | Safety panel ( Three panels) | set | √ | √ | √ | √
1.19 | View box, double films | set | √ | √ | √ | √
1.2 | view box, 3 films | set | √ | √ | √ | √
1.21 | washing tank with lid, ss 13. 5 Ltr | piece | √ | √ | √ | √
1.22 | washing tank with lid, ss 22 Ltr | piece | √ | √ | √ | √
1.23 | washing tang with lid, 45 Ltr | piece | √ | √ | √ | √
1.24 | Film 12 inch x 10 inch (packet of 50 sheets) | packet | √ | √ | √ | √
1.25 | Fiml 12 inch X 15 inch (packet of 50 sheets) | packet | √ | √ | √ | √
1.26 | Film 14 inch X 17 inch (packet of 50 sheets) | packet | √ | √ | √ | √
1.27 | Fixer 13 . 5 Ltr | packet | √ | √ | √ | √
1.28 | Developer 13.5 Ltr | packet | √ | √ | √ | √
1.29 | Replenisher 4.5 Ltr | packet | √ | √ | √ | √
1.3 | Intensifying screen (speed 400)10inch X 12inch | piece | √ | √ | √ | √
1.31 | Intensifying screen (speed 400)12inch X 15inch | piece | √ | √ | √ | √
1.32 | Intensifying screen (speed 400)14inch X 117inch | piece | √ | √ | √ | √
1.33 | Clip Hanger 10 inch X 12 Inch SS | piece | √ | √ | √ | √
1.34 | Clip Hanger 12inch X 15 Inch SS | piece | √ | √ | √ | √
1.35 | Clip Hanger 14 inch X 17 Inch SS | piece | √ | √ | √ | √
1.36 | Sationary grid 6:1 10inch X 12 inch | piece | √ | √ | √ | √
1.37 | Sationary grid 6:1 12inch X 15 inch | piece | √ | √ | √ | √
1.38 | Sationary grid 6:1 14inch X 17 inch | piece | √ | √ | √ | √
1.39 | Half film blocker 12 inch X 15 inch | piece | √ | √ | √ | √
1.4 | Half film blocker 14 inch X 17 inch | piece | √ | √ | √ | √
1.41 | Abdominal shield | piece | √ | √ | √ | √
1.42 | lead apron | piece | √ | √ | √ | √
1.43 | Lead figures, A-Z | set | √ | √ | √ | √
1.44 | Lead figures, R-L | set | √ | √ | √ | √
1.45 | Lead numbers, 0-9 | set | √ | √ | √ | √
1.46 | Glass marking pencil, white |  | √ | √ | √ | √
1.47 | Safe light holder : with filter and bulb | set | √ | √ | √ | √
2 | Computed Tomography
2.1 | Consumables
2.1.1 | Film - dry view laser imaging film: Size 14 x 17 inch (compatible with the machine) | Packet of 125 sheets | √
2.1.2 | Extension line CL Line 155 cm | piece | √
2.2 | Equipment
2.2.1 | CT Machine 16/64 slice Specification: Brilliance CT 16 slice: Infant cradle, Extended Brilliance Workspace, Bone mineral analysis package PC: HP dx 2000PC, Color laser printer HP 3800n, Kodak Dry view laser camera 8900, Stabilizer, UPS with 30 minutes backup – 125KVA, Pressure injector: MEDRAD Stellant, Defibrillator: Heart start XL, Separate PC for conferencing and teaching aids, Anesthetic machine with pediatric and adult ventilator. Cardiac monitor with stand and Pulse oximeter compatible to the machine. Scan tools Pro, RRCV tools kit, Functional CT/MR image fusion, CV pro analysis package, Virtual colonoscopy, Dental planning package, Lung disease quantification package, Bone mineral analysis, (Installed in Computer). | Set | √
2.2.2 | Operator’s Manuals/ Documentations: Bone mineral analysis program (Q-MAP II) CT operator’s manual, Brilliance CT – Instruction for use – English Volume -1, Brilliance CT – Review – Volume – 2, Brilliance CT – Basic analysis – English – Volume – 3, Brilliance CT – Advanced analysis – English – Volume – 4, Extended Brilliance Workspace – Instruction for use – Volume – 1, Extended Brilliance Workspace – Viewing modes – Volume – 2, Extended Brilliance Workspace – Analysis – Volume – 3, Extended Brilliance Workspace – View forum – Volume – 4, Extended Brilliance Workspace - Supplement version 2.1 Instruction for use, Extended Brilliance Workspace – Cardiac application guide, Brilliance CT – 16 power configuration: Cardiac applications, CT patient log, Kodak Dry view 8900 laser imager – Release 3: User guide, Stellant CT injection system: operation manual, M3/M4 Monitors (M3046A) Measurement server: Instruction for use, Heart start XL - Instruction for use, Heart start XL – User training workbook, Oxywave pulse oximeter: Operation manual, Drager Anesthesia Ventilator, Model Fabius-CE: Operation manual, Drager Anesthesia Ventilator, Model Fabius-CE: Medibus for Fabius CE. | Set | √
2.2.3 | CDs/MODs: Brilliance_ Host_ OS_2.0.0.10, Brilliance CT- Host 2.0.2.13005, Brilliance CT- incremental 2.0.2.13006, Brilliance CT- versions 2.0.2. 13011, CIRS – 4.1.19.5, CIRS- Service- 4.1.19.5, Brilliance CT- Firmware-2.0.2.13001, A plane collimator calibration files: S/N 1758, Systems Diagnostic utilities: Revision:014, Dell Precision Workstation resource CD. | Set | √
2.2.4 | Extended Brilliance Workspace V2.1.1 –Software CD set: P/N 4550 122 02131: Unattended V1.25 – EBW-V2.01: Disk: 1of 6, Environmental creator –EBW-V2.0.1: Disk: 2 of 6, COTS (3rd party) V2.0.1- EBW-V2.0.1: Disk: 3 of 6, Application software – EBW-V2.0.1: Disk: 4 of 6, View forum – EBW- V2.0.1: Disk: 5 of 6, Mandatory Add-on – EBW – V2.0.1: Disk: 6 of 6, Update version –V2.0.11- EBW, EBW-V2.1.0, EBW-V2.1.1, Brilliance Hot fix, NVIDIA FX540 Drivers CD For Dell 670-EBW. | Set | √
2.2.5 | Drivers & Utilities – Dell Precision Workstation Resource CD, Q-MAP II Version 3.15, Q-MAP II Instructional CD: Rev. A, Merge DSP 1.0.2 Dicom storage provider, Microsoft Windows XP Professional Service Pack2: Operating system CD – For HP PC, Microsoft Windows XP Professional Service Pack2: Version 2002 – Compaq PC, HP/Compaq DX 2000 MT drivers, HP restore Plus: For HP dx 2000 MT PC, HP Business Desktop documentation library CD: dx2000 model, HP color LaserJet 3800, Norton Antivirus, LCD monitor – Electronics user’s manual: Version – H6195, Kodak Dry view 8900 laser imager – Version 1.0, Default grid voltage file: Tube No. 105702 MRC 600, Heart start XL: User training CD ROM, Sony magneto optical Disk – 9.1GB: Model EDM 9100C. | Set | √
2.2.6 | Cables and Hardware accessories/spares: Spike buster/Extension box with six outlets, Computer power cables, RJ 45 straight network – long, RJ 45 straight network – short, RJ 45 cross straight network cables – short, DVI Video cable for Philips LCD monitor, VGA cable for Philips LCD monitor, USB cable, Audio cable, Examination table lifting service bracket, Kodak dry view laser camera connector, C7 PCB for consul servo voltage stabilizer, C7 trip PCB for consul servo voltage stabilizer, Transformer CC10 for consul servo voltage stabilizer, Transformer CC11 for consul servo voltage stabilizer, 12 teeth pinion for consul servo voltage stabilizer, Carbon brush for servo voltage stabilizer, Packing sims-1mm, Packing sims-2mm, CT tube repairing kit, Laser aperture tool, Cable for QMAP PC, Parallel cable, Earth cables for ECG Monitor, Binder, Dongle for QMAP PC, VGA adaptor HD15/DB15. | Set | √
2.2.7 | Dry view camera compatable to CT (laser imaging) | piece | √
2.2.8 | Patient Positioning straps (Belts) | Set | √
2.2.9 | UPS with 30 minutes backup-125 KVA | Set | √
2.2.10 | UPS Battery: US-12-125: 12V-125AH. Valve regulated: Constant voltage recharge. Lead acid: Boost-equalize – 14.5-15 OV. Rechargeable battery: Float – 13.5-13 BV. Max. – 37.50 A | Set | √
2.2.11 | Automatic injector | Set | √
2.3 | CT Accessories
2.3.1 | Extended Phantom performance kit | Kit | √
2.3.2 | Eye Phantom | Kit | √
2.3.3 | Phantom holder | Set | √
2.3.4 | CT slicker cushion | Set | √
2.3.5 | Foot rest extension | Set | √
2.3.6 | Coronal head rest assembly | Set | √
2.3.7 | Standard head rest assembly | Set | √
2.3.8 | Kit couch cushion | Set | √
2.3.9 | Patient restraint upgrade kit | Kit | √
2.4 | CT and MRI
2.4.1 | Lazer camera compatible to CT-MRI | Set | √
2.4.2 | Stellant syringes: content, 2-200ml syringes, 1-60 inch T tube connector, 1 quick filled tube | set | √
3 | Interventional Radiology
3.1 | For Angiography (Tace, Bronchial, Uterine and Renal Artery Embolization)
3.1.1 | Puncture Needle 18, 19, 20G sulflo IV Catheter with radio-opaque ETFE. (Length- 2 ¾ inch.), double wall puncture. | Set | √
3.1.2 | Guide wire (puncture) with size 0.035. | Set | √
3.1.3 | Vascular sheath 5, 7 Fr. | Set | √
3.1.4 | Single wall puncture needle 18-21G. | Piece | √
3.1.5 | Cobra catheter 4 & 5 Fr (length 65, 100 cm) | Piece | √
3.1.6 | Simon (Simon II) catheter 4. 5 Fr. (length 65, 100 cm) | Piece | √
3.1.7 | Hydrophilic guide wire size0.032, 0.035 inches (length 80, 100 & 125 cm) | Piece | √
3.1.8 | Angiography suite with Digital Subtraction facility | Set | √
3.1.9 | Simon (I) catheter (4Fr, 5 Fr) length 65, 100 cm. | Piece | √
3.1.10 | Yashiro catheter (4, 5Fr) length 70, 100 cm | Piece | √
3.1.11 | Head hunter (4, 5Fr) length 70, 100 cm. | Piece | √
3.1.12 | Pigtail (intravascular) 6- 8 Fr. | Piece | √
3.2 | For Ablation of liver tumour
3.2.1 | RFA generator | Set | √
3.2.2 | RFA electrodes | Set | √
3.3 | Percutaneouse Drainage
3.3.1 | Chiba Needle 18 Gauge (15 cm) | Piece | √
3.3.2 | Chiba Needle 18 Gauge (20 cm) | Piece | √
3.3.3 | Chiba Needle 20 Gauge (15cm) | Piece | √
3.3.4 | Chiba Needle 20 Gauge (20cm) | Piece | √
3.4 | Guide wires
4 | Bentson J TIP sizes (0.035, 0038 inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
3.4.2 | Bentson straight tip (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
3.4.3 | Bentson supper stiff J TIP (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
3.4.4 | Amplatz extra stiff J tip (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
3.4.5 | Amplatz extra stiff straight tip (0.035inches) 80-145 cm-any standard GW length for specific procedure. | Piece | √
3.4.6 | Facial dilators (plastic) standard length: 6 Fr, 7 Fr, 8 Fr, 9 Fr, 10 Fr, 12 Fr. | Piece | √
3.4.7 | Self retaining loop catheter (Pigtail) standard length: 8 Fr, 10 Fr, 12 Fr, 14 Fr. | Piece | √
3.4.8 | Micro punctures (billiary drainage set) consist of: acoustic needle, J-tip guide wire, coaxial introducer, Needle with stylet, extra stiff G W with a fluffy tip | Set | √
3.4.9 | Percutaneous B: spring activated core biopsy needle with coaxial system, 18 gauge (15 cm), 18 gauge (20 cm), 20 gauge (15 cm), 20 gauge (20 cm), spinal needle (15, 20 cm) 18, 20, 22, 23, & 25 G | Set | √
3.4.10 | Mammography Machine, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Set | √
4 | Magnetic Resonance Imaging
4.1 | Equipments
4.1.1 | MRI friendly Suction machine 1/2 HP, mobile fibre casting pressure tube with pressure regulator tube. Twin suction bottle of 2-3 Ltrs capacity. Display of vacuum from 100 to 700mm HG. (MRI compatible) Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √
4.1.2 | CO2 cylinder for MRI 'A' type (Re-filling) | Piece | √
4.1.3 | MRI compatible wheel chair | Piece | √
4.1.4 | MRI compatible Patient trolley | Piece | √
4.1.5 | MRI compatible IV Stand | Piece | √
4.1.6 | MRI compatible Anaesthesia machine Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √
4.1.7 | MRI compatible Pulse Oximeter, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top | Set | √
4.1.8 | Dry view camera, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top. | Piece | √
4.1.9 | MRI UPS | Set | √
4.1.10 | MRI Machine with automatic injector | Set | √
4.1.11 | MRI machine 1.5 Tesla specifications: ADW CPU, ADW monitor, mouse and mouse pad, metal detectors, Kodak dry view 8900 camera, report card CPU, Monitor, key board, CCTV, newmatic slimline head sets, microphone, Onkyo music system with speakers, AM/FM antenna, GOC computer, Monitor, PAC monitor, spectris solaris pressure injector 2 batteries and charger, SCSI tower, keypad, mouse and pad, gradient chiller, 1.5T sumitomo compressor, Cryo chiller, Quad knee coil, 8 channel brain coil, 1.5T split head coil, 8 channel body array, 8 channel Neuro Vascular array, 8 channel CTL coil, 8 channel anterior bridge, general purpose flex coil, phased array shoulder coil, 8 channel breast coil, QD plug coil adaptor extremity, coil ID connector, 1.5T peripheral Vascular coil, 1.5T 5 inch round coil, 1.5T 3 inch round coil, Dual array adaptor, MR compatible adaptor, MR compatible anaesthesia m/c, phantom cart with LV shim phantom, DQA phantom with loader, TLT head sphere, ball phantom, upper anular phantom, CTL phantom, SNR phantom head SNR phantom, ball phantom with base, 8 channel NV array phantom, loading phantom, service ladder, newmatic slimline head sets, invivo pulse oxymeter, ERU, Oxygem monitor with sensor, Ethemet awitch, 1.5T CX K4150 magnet for Twinspeed gradient module, GEDE 100KVA UPS, PDU/HFD cabinet, TAC cabinet, RFS cabinet, CO2 cylinder for door locking, magnet monitor, system main disconnect panel, UPS batteries, MOD media 2.3 GB, patient squeeze bulb audio alarm, signa excite 1.5T phased array cables for Twinspeed, FR cage, viewing box, MRI 3 TR ductable split, TR, 1.5 TR, CTT Linux 4.3.13, excite HD M4 Applications, Gamma correction software, 11.0 M4 patch 00-PA-CD, HD MR Driver Module HVPS, Report Card M3A Installation 2.0, 11.0 M4 patch 00-PA-CD, Functool demo exams, Volume Viewer Demo exam, Functool Applications, Reporting tool, Volume Viewer Applications, MR pating 1.1, AW fast load DVD, AW 4.2P software, 11.0 M4 Prose/Functool Patch CD, 10.0 8 Channel NV Array QA tool, AW demo exam on CDROM, 1.5T Signa Excite 11.0 LRG, Split Head Coil LRG, AW software Config Floppy, Laryngoscope ADULT, Non-magnetic Oxygen Cylinder, Emergency light, HP Color Laser Jet 3800n Printer, Non-Magnetic IV Stand, Stethoscope, Sphygnomanometer, Diagnostic Pro Advantage, MR Machine Datex Ohmeda. | Set | √
4.1.12 | UPS 80KVA | Set | √
4.1.13 | MR compatible Laryngoscope for adult and children | Set | √
4.1.14 | MR compatible anesthesia machine for adult and paediatrics | Set | √
4.2 | Consumables
4.2.1 | Film - dry view laser imaging film: Size 14 x 17 inch (compatible with the machine) | Packet of 125 sheets | √
4.2.2 | Ear Plugs (spongy foam) | Piece | √
4.2.3 | MRI compatible ECG leads | Piece | √
4.2.4 | Patient Positioning straps (Belts) | Set | √
5 | Accessories
5.1 | MRI UPS Battery: Highlit, 12xLB420, 12v 416w/cell/15 min to 1.67v/cell at 25 degree C (77 degree F), Float voltage 13.62v at 20 degree C (68 degree F), 13.56v at 25 degree C (77 degree F), connection torque 8-9Nm/70-80 mlb. | Set | √
5.2 | Hand Held Meta Detector: Hand held super scan metal detector weighing 10 oz. | Set | √
5.3 | 1.5T 8 Channel Brain Coil: This eight element head coil for the use with 8 channel EXCITE 1.5T MR system provides unprecedented high SNR. Critically important for high throughput, the coil is optimized for ASSET parallel imaging with a speedup factor of two. With a 24 cm diameter and S/I coverage of 30 cm, the coil features an open, patient friendly design with a lookout mirror to let patients view the world outside the system. This coil will change your expectations for clinical brain imaging intracranial MRA and paediatric imaging (8-ch option is mandatory) | Set | √
5.4 | 1.5T Channel CTL Array: This 12 element, multi-station array for use with 8-channel EXCITE 1.5 T MR system delivers extraordinarily high SNR, spatial resolution and coverage over the full range of spine, soft tissue neck and carotid studies. It’s packaged to conform to the spin’s normal curvature and includes a patients comfort pad and restraint. Patients can be positioned head first or feet first and the coil’s extensive coverage 75 cm in the S/I direction-accommodates imaging of the entire spine, making it a truly high productivity solution for your site. This coil’s unique split top design extends its clinical flexibility. The removable top facilities routine neck imaging, including the anterior cervical spine and neck and eliminates the need to switch to a dedicated neck coil. Coil markers make accurate positioning at imaging isocenter surprisingly quick and easy. The result improved productivity and patient throughput. This package includes applications manuals, operating instructions, SNR testing phantoms and formfitting patients comfort pads. | Set | √
5.5 | 1.5 T Open Neurovascular Array: With a four-channel, quadrature phased array design and extended, 48 cm S/I coverage, this receive only coil is specially suited to high-SNR imaging of the neck, head, brain and vascular system. It features a two piece removable top with many openings and an adjustable mirror system to reduce claustrophobia. | Set | √
5.6 | 1.5T 8 Channel body Array: This is a 12 element coil designed for use in creating extraordinary images of the chest, abdomen and pelvis on 8 channel EXCITE 1.5T MR system. This coil offers such extensive coverage that repositioning of the coil for combined abdomen-pelvis exams typically is not necessary. It provides higher sensitivity than anybody coil offered previously by GE, to enable the highest possible signal-to-noise ratio. And particularly important for body imaging, the coil was optimized for the use of ASSET98 ch option is mandatory | Set | √
5.7 | USAI 1.5T Phased Array shoulder coil: The USAI 1.5 T shoulder array coil offers the increased signal of the phased array coil along with a unique sleeve design that gives the user exceptional joint imaging capabilities. Added benefits of teh coil include. This shoulder coil is able to image the shoulder region providing images with clear definition of the shoulder joint, specifically the head of the humerus, clavical, acromion bone, supraspinatus muscle and ligaments. This shoulder coil is able to image the shoulder region providing images with clear definition of the shoulder joint, specifically the head of the humerus, clavical, acromion bone, supraspinatus muscle and ligaments. This shoulder coil is able to image the shoulder region providing images with clear definition of the shoulder joint, specifically the head of the humerus, clavical, acromion bone, supraspinatus muscle and ligaments. The shoulder coil is able to image both shoulders (not at the same time). The shoulder coil provides patients comfort for patient imaging of up to 40 minutes procedures. The shoulder coil is ASSET compatible and delivers optimum FATSAT imaging. | Set | √
5.8 | Peripheral Vascular phased Array Coil: Coil is a specially designed 12 elements phased array coil with 1.26 meter extended coverage for peripheral vascular run-offs from renal to dorsal digital arteries with smartstep2000. The other applications include visualization of lower abdomen and lower extremities. High signal uniformity. 1.26 meter coverage in S/I direction (3 x 44 cm steps with 3 cm overlap). Physical dimension 1.37 m long and 74 cm wide. Flexible and 2 part coil structure for close proximity to patient. Received only phased array design. Wide positioning strips with Velcro captures is provided with the coil to hold the coil closely to the peripheral leg vessels into a common plane. Must ensure that M1087CY pre amp kit also ordered with the system. Includes signal to noise test scanning phantom for image quality assurance procedures. | Set | √
5.9 | MAI QUAD EXTREME COIL W ID | Set | √
5.1 | Open breast phased Array coil: Coil is specially designed for imaging and visualization of the female breast and auxiliary tissue. It is possible to perform either unilateral imaging of the either left or right breast or bilateral imaging of the both breast. | Set | √
5.11 | 3 Inch circular general purpose coil: This is a 3” circular general purpose coil for applications in orbits, IAC, ankle and wrist. Can be combined with another 3 inch GP coil for simultaneous high resolution dual imaging of TMJ, Orbits and IAC (Dual array adapter is mandatory for such applications). | Set | √
5.12 | 1.5T GP flex coil: Coil is designed to optimize imaging of irregular anatomy such as the branchial plexus, hip, thigh, elbow, large knees, ankle and foot. | Set | √
5.13 | 5 inch circular general purpose coil: This coil is useful for a wide range of small parts applications-including the orbits, IAC, ankles and wrist. And with the help of a dual array adaptor, two of the coils can be used concurrently for simultaneous high-resolution of TMJ, orbits or IAC. | Set | √
5.14 | Dual Array Adapter: Allows for combining two receive only non phased array coils. This helps in high SNR imaging of symmetrical anatomy such as TMJ | Set | √
6 | Radiology Ultrasound
6.1 | Equipments
6.1.1 | USG Machine - Type II: 'USG Machine - Type II: Specification: Full backlit Alphanumeric keyboard, telescoping, tilt, swivel, high adjustment. The system includes patient entry preset menus archiving, edge enhancement, frame averaging, storage and recall. Features: Truscan Architecture - software that gives computational power, image manipulation capability workflow flexibility and product upgradeability. Truaccess - a raw data processing technology for image enhancement. Codescan - Improve image qua. Transducer : 3.5 - 5 MHz Convex probe for ABD, OB GYN, Linear probe 7 - 10 MHz Additional probe Port. Accessories: Footswitch, cable set for B/W printer, Fixture for B/W printer, 2KVA UPS. Measurement capabilities: Distance, Circumference, area and volume. Details: Multipurpose imaging System design for abdominal, vascular, obstetrics, gynecology, cardiology, neonatology, urology, transcranial and small parts application. Servo controlled Voltage stabilizer. | Set |  |  | √
6.1.2 | USG Machine - Type III:'USG Machine - Type III: Transducer : 3.5 - 5 MHz Convex probe for ABD, OB GYN, Linear probe 7 - 10 MHz Additional probe Port. Details: 17" Flat screen colour doppler with comprehensve archiving facilities. Recording facilities: VCR, CDROM/FFD/DVD/UBS thum drive. Features: State of the art system with DVD and standby mode, coded contrast imaging, cross beam, report writer, stress echo package and PDF. Scanning modes include B-mode, M-mode, colour flow, pulsed wave, power doppler and real time triplex mode. Accessories : Footswitch, cable set for B/W printer, cable set for VCR, Fixture for B/W printer, VCR-TV Monitor and CD writer built in, 2KVA UPS and Fixture for VCR/Colour printer. Endovaginal probe, endorectal probe with 7 to 10 MHz. TRUS guided biopsy probe.Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top and suitable UPS | Set | √ | √ | √ |  | where there is X-ray services in DH
6.1.3 | Ultrasound Machines (Portable with laptop type), with color Doppler facility and atleats 2 probes (convex and linear), with complete set of printer. Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top. | set | √ | √ | √
7 | Consumables
7.1 | Thermal printing paper type I (Sony), ultra sound (110mmx20m) | Roll | √ | √ | √
7.2 | Ultrasound gel | 250 ml | √ | √ | √
7.3 | Probe: curvi linear for trans abdomin ultrasound (compatable with existing machine) |  | √ | √ | √
8 | Probe: linear (compatable with existing machine, mention the Specs)
8.1 | Probe: transvaginal (compatable with existing machine) |  | √ | √ | √
8.2 | Printer for ultrasound machine |  | √ | √ | √

## Surgery
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Surgery Department
1 | Endoscopy
1.1 | Equipments
1.1.1 | Argon Plasma Coagulation (APC) unit: For open Endoscopic and Laparoscopic. Power supply 220-240v, 50Hz. All power cord should be of round 3 pin plug top. | Piece | √ | √
1.1.2 | Argon Gas Cylinders: Compatible to the APC Generator | Piece | √ | √
1.1.3 | Argon Plasma Coagulation (APC) hand piece: compatible to the energy generator | Piece | √ | √
1.1.4 | Video endoscope Medical Monitor: Flat screen, 14” to 16” | Piece | √ | √
1.1.5 | Video endoscope Processor, 3 chips camera high definition | Piece | √ | √
1.1.6 | Gastro-Scope compatible to above processor | Piece | √ | √
1.1.7 | Colono scope compatible to the above processor | Piece | √ | √
1.1.8 | UPS for the entire system | Piece | √ | √
1.2 | Accessories
1.2.1 | Biopsy forcep for upper GI ENDOSCOPY | Piece | √ | √
1.2.2 | Biopsy forcep for Colonoscope | Piece | √ | √
1.2.3 | Hemoclip for GI bleeding | Set | √ | √
1.2.4 | Multiband Ligator for Esophageal verices | Piece | √ | √
1.2.5 | PEG (Endoscopic Gastrostomy Set) | Set | √ | √
1.2.6 | Esopageal dilator set | Set | √ | √
1.2.7 | Esopageal stents | Piece | √ | √
1.2.8 | Esopageal varices sclerotherapy injector (needle) | Piece | √ | √
2 | Laproscopy
2.1 | Equipments
2.1.1 | Laproscopic camera 3 chips high definition: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
2.1.2 | Laproscopic camera unit: flat screen, 18 inches | Piece | √ | √
2.1.3 | Laproscopic Medical Monitor: Power supply 220-240v, 50Hz. Suitable Voltage stabilizer. Power cord with round 3 pin plug top | Piece | √ | √
2.1.4 | Laproscopic CO2 insufflator: Power supply 220-240v, 50Hz. Power cord with round 3 pin plug top | Piece | √ | √
2.1.5 | Laproscopic Light source: compatible to the above light source | Piece | √ | √
2.1.6 | Laproscopic Light cable: 0 degree,10 mm | Piece | √ | √
2.1.7 | Laproscope lens: 30 degree,10 mm | Piece | √ | √
2.1.8 | Laproscope lens | Piece | √ | √
2.2 | Consumables
2.2.1 | Cleaning brush: Size: 16 mm | Piece | √ | √
2.2.2 | Disposable 12 mm abdominal trocar | Piece | √ | √
2.2.3 | Disposable 10 mm abdominal trocar | Piece | √ | √
2.2.4 | Reducer: (from 10 mm to 5 mm ) compatible to Trocar | Piece | √ | √
2.2.5 | Reducer: (from 12 mm to 5 mm ) compatible to Trocar | Piece | √ | √
2.2.6 | Disposable 5mm abdominal trocar | Piece | √ | √
2.2.7 | Laparoscopic Bipolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
2.2.8 | Laparoscopic Monopolar Diathermy cord: Compatible to the hand piece in use | Piece | √ | √
2.2.9 | Liga clip: LT 300 | Piece | √ | √
2.2.10 | Liga clip: LT 400 | Piece | √ | √
2.2.11 | Oil for the instruments maintenance | ml | √ | √
2.2.12 | Single clip applicator: Titanium clip mediun for LT 300 (Ethicon) | Piece | √ | √
2.2.13 | Single clip applicator: LT 400 ( large Titanium clip) | Piece | √ | √
2.2.14 | Claw extractor with monopolar with Ratiched handle 33.2 cm, size 5mm | Piece | √ | √
2.2.15 | Curved dissector with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
2.2.16 | Grasper with monopolar cautery 33.2 cm, size 5mm | Piece | √ | √
2.2.17 | Laparoscopic Aspiration needle: size bigger than 20 Fr. | Piece | √ | √
2.2.18 | Curved dissector with Bipolar cautery 33.2 cm: size 5mm | Piece | √ | √
2.2.19 | Laproscopic Babcock (10 mm), size 10mm | Piece | √ | √
2.2.20 | Curved scissor with monopolar cautry | Piece | √ | √
2.2.21 | Gall bladder holding forceps fenestrated (grasping forceps). 5mm, 33.2 cm | Piece | √ | √
2.2.22 | Irrigation and suction cannula | Piece | √ | √
2.2.23 | Laproscopic Hook dissector : 5mm diameter shaft length 34 cm | Piece | √ | √
2.2.24 | Merry lamp forcep: 32.2 cm, 5 mm | Piece | √ | √
2.2.25 | Metal trocar: size 5mm | Piece | √ | √
2.2.26 | Metal trocar: size 10mm | Piece | √ | √
2.3 | Accessories
2.3.1 | CO2 regulator for D-type cylinder | Piece | √ | √ |  |  | for D-type cylinder
2.3.2 | CO2 regulator for B-type cylinder | Piece | √ | √ |  |  | for B-type cylinder
2.3.3 | CO2 cylinder D type | Piece | √ | √ |  |  | D-type
2.3.4 | CO2 cylinder B type | Piece | √ | √ |  |  | B-type
2.3.5 | High pressure Hose (CO2), connection length 10m | Piece | √ | √ |  |  | connection length10 m.
3 | Neurology
3.1 | Equipment
3.1.1 | Intra Cranial Pressure Monitoring system with hydrophobic filter with two convient sites, sliding chamber, locking bracket, with drainage bag of capacity 700cc and system panel and patient connecting wire with barium impregnated silicon catheter tubing and ventricular silicon catheter of internal diameter 1.5 mm, outer diameter 2.8 mm, length 35 cm with CSF flow control valve of medium pressure 12 cm size. | Set | √
3.1.2 | Endovenous LASER System: Laser Type Diode laser, CW, class IV, Centre Wavelength 1470 nm, Maximum Power 12 W to the laser output port, Aiming Beam Red Class IIIa diode laser, ≤5 mW at port, wavelength 635–655 nm, Pulse Interval 100 ms–1000 ms, Power Supply 100–240 V AC, 50–60 Hz, 500 VA max, Dimensions (H x W x D) 150 mm x 370 mm x 325 mm (± 5 mm), Weight 12 kg max. DIODE LASER:VenaCure 1470 nm Laser MMC Card-64 MB Interlock Connector Foot Switch Travel Case Spare Key-VenaCure 1470nm Laser. ACCESSORIES: USB Memory Card Reader/Writer Warning Signs Premium Laser Safety Glasses Standard Laser Safety Glasses Emergency Reset Device User Manual Spare Fuses. Procedure Kit set, Micro Introducer Kit set and catheters | set | √
3.1.3 | May Field Cranial stabilization skull clamps, Codman 360 degree rotation of rocker arc with clamp and stainless pins. | Piece | √
3.1.4 | Micro Chair (Hydraulically-operated): Drive System: DC 24V, Elevation/lowering: 44cm-86cm, Back/forth & right left sliding: free movement. Rotation counter and clockwise reclining up to 6 degree from vertical with arm rest. Power consumption 400VA. Power source 100-110V, 50/60Hz. Insulation resistance of 400M ohm, lekage current 50 micro amperes. weighs approximately 130kgs (Should be compatible with microscope leica M651) | Set | √
3.1.5 | Multi-Purpose Head Frame: Basal frame/E.Semicircular frame, Hand rest with slider /instrument support, Spring scalp HoOK retractors (6nos), Cotton plate, Head holder with 4 standard pins, Table attachment, Head pin standard, Self retaining retractors (Large, Small), Soft brain spatulas (Large tip width 6mm, medium tip width 4mm ,small tip width 2mm), Universal attachment 1no, Slide adjuster 1no, -shaped head rest (large and small) 1no Piece, Self retain retractor attachment 1no. | Set | √
3.1.6 | Neuro Surgery Operating Table: Power Source voltage: AC 100V, 220V, Driving voltage: DC 24V, Battery: DC 12V x 6, Power ConsumptionVA (MST- 7000N), Table top: 450mmx1895mm, Elevation Stroke: 480-980mm (High/low speeds), Trendelenburg (high/low speeds): head –down 40 degree Head –up 20 degree. a. Operating Table Horseshoe Head pad - adult (20.5 OD x 7.5 ID x 4.5cm) b. Operating Table Horseshoe Head pad - pediatric (14 OD x 5.7 ID x 3.2cm) c. Operating Table Supine Head pad - 23x25x5.7cm d. Operating Table Lateral Head pad with centre dish - 23x25x2.5cm e. Operating table prone head rest - Medium 22.2x19.7x12cm f. Operating table prone head rest - Medium 22.2x19.7x12cm g. Operating table prone head rest - Large 27x22.9x13.4cm h. Positioning Chest rolls - Adult 50.8x10cm i. Positioning Chest rolls - Paediatric 30.5x5cm. Operation theatre stirrup pads - boot/ stirrup pad set 68.5x36T/12.5B/6cm. | Set | √
3.1.7 | Neurosurgery operating microscope (portable) | Set | √
3.1.8 | Nitrogen Gas cylinder - D type | Piece | √
3.1.9 | Ultrasonic aspirator system : 1. universal handpiece 2. disposable tips for soft tissue -superlong straight tip, spetzler barracuda large diameter tip, straight tip, straight large diameter tip, straight micro diameter tip, spetzler barracuda tip, ring curette tip; 3. accessories - step torque wrench offset 7mm, cleaning wire, sterilization tray, console 110V(including foot pedal and IV pole), plastic suction canister reusable, disposable suction canister liner, disposable tubing set & extender filter tubing, cart. | Set | √
3.1.10 | Neuro-endoscopy set consisting of adult and child:Neuroendoscopy Instrument Set includes Injection Needle 1.7mm; Deflecting Mechanism for LASER probe and ring-grip handle, diameter 2.9mm,length 38cm; Grasping Forceps-diameter 2.7mm, working lenght 30cm; Biopsy Forceps with working length 30cm; Scissors- pointed with single action jaws(diameter 2.7mm) and doble action jaws(1.7mm) with working length 30cm; Forceps for ventriculostomy,flexible with diameter 1mm; Coagulating Electrodes-both unipolar and bipolar with diameters 1.7mm; Autoclavable Holding Systems consisting of socket, straight articulated stand, and clamping jaw; Hopkins Wide Angle and Straight Forward Telescopes 6degree,with angled eyepiece, with instrument channel diameter 3mm, lenght 15cm, autoclavable, fiber optic light tranismission incorporated, color code:green; Hopkins Wide Angle forward-oblique telescope 30degrees diameter 4mm, length 18cm color code:red; and Straight Forward TelescopesOperating Sheaths- outer diameter 6.5mm, working length 13cm; Obturators; Telescope Bridges; Irrigation Tube with Luer-lock; suction cather disposable; and Adapters; containers for sterilization for teklescopes,angled telescopes and instruments. | Set | √
3.1.11 | Surgical Clipper with fixed head new, unique cordless/ coded operation | Set | √
3.1.12 | Electric Craniotome: Motor speed should be atleast 75,000 rpm at 8 bar pressure, motor should be light(preferable 85g or less), main motor unit should be detachable from cable, straight and angled attachment of various lengths should be available for cranial and spinal surgery, change of hand piece with mounted tool should be possible, motor should be converted to an angulated position with or without an adaptor, sound level should be very low, no intra-operative oiling of motor should be necessary, duraguard should be detachable from craniotome hand piece, so that same hand piece can be used for paediatrics, adults and spine surgery, quick coupling attachment should be available, single use and reusable burrs should be available, sterilization through flash or regular steam autoclave, perforator driver with cutter should be available , should be able to use saw hand piece with same system, in-built or optional irrigation pump should be available, five years warranty followed by AMC quote for next five years. Power supply 210-240V. The following accessories should be included: (Should be compatible with Leica microscope 651) | Set | √
3.1.13 | Following accessories should be included with the machine (Electric Craniotome): A. Hand Piece: Hand Piece, straight, 120 mm. Hand Piece, straight, 90 mm. Hand Piece, straight, 160 mm. B. Craniotomy attachment: Craniotome hand piece. Fixed duraguard, adult. Fixed duraguard, paediatric. C. Craniotome cutter: Craniotome cutter, paediatric. Craniotome cutter, adult. D. Perforator: Perforator driver. Cranial perforator, 9 x 12 mm, Hudson type. Cranial perforator, 6/9 mm, Hudson type. Hudson chuck. Spare cutter for perforator, 9 x 12 mm. Spare cutter for perforator, 6 x 9 mm. E. Burrs: Rosen Burr D 3.1 mm for 120 mm hand piece. Diamond burr D 3.2 mm for 120 mm hand piece. Barrel burr D 4.0 mm for 160 mm hand piece. Barrel burr D 4.0 mm for 120 mm hand piece. Neuro cutter D 2.3 for 120 mm hand piece. Neuro cutter D 3.1 for 120 mm hand piece. Neuro cutter D 3.1 for 160 mm hand piece. Acorn burr D 6.0 mm for 90 mm hand piece. Pin point cutter D 1.5 for 120 mm hand piece. Twist drill D 1.5 mm for 90 mm hand piece. F. Micro sagital saw attachment: Micro sagital saw pencil shape. Saw blade for micro sagital saw 9/13/0.3/0.3 mm. G. Storage and maintenance: Oily spray for high speed motor and hand pieces. Oily spray for perforator. Perforator basket with covering lid with holders for motors, all hand pieces, hose tools and all other accessories. | Piece | √
3.1.14 | Operating binocular Loupes with 4X times magnification | piece | √
3.1.15 | Skull bone fixation system (cranial fixation system) Instruments consisting of applier,holding forceps, pin cutter, removal forceps, storage tray; instrument oil (aerosol);Sterilit instrument oil(bottle) ; Implants with clamp, 11mm (box of 12), clamp, 16mm (box of 6), clamp, 20mm (box of 6) Sterilit instrument oil (bottle) Container for storage tray | set | √
3.1.16 | Negative Pressure Wound Therapy system or Vacuum assisted closure system:Unit Weight with empty canister3.35kg (7.4 lbs with a 500cc empty canister and no solution bag) Dimensions with empty canister217mm x 260mm x 191mm (8.55in x 10.25in X 7.5in) Power SupplyClass I Equipment Ordinary Equipment Battery TypeCustom lithium battery Battery LifeApproximately 6 hours Battery Recharge TimeApproximately 4 hours Battery IndicatorThree-segment; low-battery alarm when <2 hours Electrical Data (Power Supply)Voltage 100 - 240 VAC Frequency 50Hz - 60Hz Power 60W. Other accessories includes Foam dressings , Pad and Tubings. | Set | √
4 | Consumables
4.1 | Aneurysm Clips 1/1(mm): Bayonet (s) opening width 11.0(10mm)bayonet (l) opening width 13.0(12mm) | Piece | √
4.2 | Aneurysm Clips 1/1(mm): Bayonet opening width 9.0(7mm) | Piece | √
4.3 | Aneurysm Clips 1/1(mm): Bent opening width 8.0 (8mm) | Piece | √
4.4 | Aneurysm Clips 1/1(mm): Bent, opening width 11.0(12mm) | Piece | √
4.5 | Aneurysm Clips 1/1(mm): Side ward bent, opening width 11.0(12mm) | Piece | √
4.6 | Aneurysm Clips 1/1(mm): Slightly curved, opening width 10.0(12mm) | Piece | √
4.7 | Aneurysm Clips 1/1(mm): Straight (large) opening width 9.0(10mm) | Piece | √
4.8 | Aneurysm Clips 1/1(mm): Straight (small) opening width 8.0 (7mm) | Piece | √
4.9 | Aneurysm Clips 1/1(mm): Straight, opening width 11.0(12mm) | Piece | √
4.1 | Fenestrated large type 1/1 (mm) – Straight (L) opening width 12.0 (12 mm) | Piece | √
4.11 | Fenestrated large type 1/1 (mm) – Straight (M) opening width 11 (9 mm) | Piece | √
4.12 | Fenestrated large type 1/1 (mm) – Straight (S) opening width 10 (6 mm) | Piece | √
4.13 | Cottonoids with X-ray detectable markers & string attached, small 1.5 x 5 cm | Pkt | √
4.14 | Cottonoids with X-ray detectable markers & string attached, large 2 x 5 cm | Pkt | √
4.15 | Surgical patties for tissue protection (x-ray detectable), size1/2 x 3 | Piece | √
4.16 | Operating microscope draping sheet plastic, disposable, compatible with microscope Leica M651 | Piece | √
4.17 | Artificial dural graft, expanded polytetrafluoroethylene 4 x 4 cm | Piece | √
4.18 | Artificial dural graft, expanded polytetrafluoroethylene 5 x 5 cm | Piece | √
4.19 | Low adherent absorbant dressing – 10 x 10 cm | Piece | √
4.2 | Low adherent absorbant dressing – 10 x 20 cm | Piece | √
4.21 | Hemostatic glue solution made of gelatin or fibrin 2-5ml | Piece | √
4.22 | Pure cotton bandage with cohesive effect 15 cm x 3.5 m (stretched) | Piece | √
4.23 | Pure cotton bandage with cohesive effect 10 cm x 3.5 m (stretched) | Piece | √
4.24 | Neurosurgical patties-sterile packages 5 x 5 cm | Piece | √
4.25 | Neurosurgical patties-sterile packages 10 x 10 cm | Piece | √
4.26 | Neurosurgical patties-sterile packages 20 x 20 cm | Piece | √
4.27 | Cranioplastic kit, sterile (methylmethacrylate medical grade in powder form with liquid polymerizing agent) | Set | √
4.28 | Peripheral vascular patch prepared from bovine pericardium, sterile packed, 1 cm x 6 cm | Piece | √
4.29 | Peripheral vascular patch prepared from bovine pericardium, sterile packed, 1 cm x 10 cm | Piece | √
4.3 | Carotid vascular patch made of collagen coated knitted polyster 2 x 8 cm | Piece | √
4.31 | Carotid vascular patch made of collagen coated knitted polyster 12 x 25 cm | Piece | √
4.32 | Bard inflation device with 30 atmosphere and 30 cc syringe | Piece | √
4.33 | Carotid endarterectomy shunt- T-shaped silicone with 15 cm length and shunt outer diameter of 3 mm distal, 4 mm proximal and bulb proximal and distal outer diameter 4.3 mm | Piece | √
4.34 | Fogarty Balloon catheters 2F, with 60 cm length | Piece | √
4.35 | Fogarty Balloon catheters 3F, with 60 cm and 80 cm lengths | Piece | √
4.36 | Fogarty Balloon catheters 4F, with 40 cm and 80 cm length | Piece | √
4.37 | Fogarty Balloon catheters 5F, with 80 cm length | Piece | √
4.38 | Fogarty Balloon catheters 6F, with 80 cm length | Piece | √
4.39 | Fogarty Balloon catheters 7F, with 80 cm length | Piece | √
4.4 | Dacron or PTFE prosthesis grafts 30 cm long and 6 mm diameter | Piece | √
4.41 | Venticulo Peritonial (VP) Shunt System Catheters (Silicone rubber) - length 62cm, Internal Diameter 1.3mm, Outer Diameter 2.5mm, Medium Pressure | Piece | √
4.42 | Venticulo Peritonial (VP) Shunt System Catheters (Silicone rubber) - length 90cm, Internal Diameter 1.3mm, Outer Diameter 2.5mm, Medium Pressure | Piece | √
4.43 | External cerbrospinal fluid drainage system: consisting of one ventricular catheter, one collecting system consisting of three way stop cock, connecting tube, drip chamber and collection bag consisting of two chambers and a folding scale. | set | √
4.44 | Transparent film dressing frame style 6x7cm | piece | √
4.45 | Transparent dressing with label 10x12cm | piece | √
4.46 | Sacrum shaped adhesive hydrocellular foam dressing 17x17cm | piece | √
4.47 | Adhesive hydrocellular foam dressing 12.5x12.5cm | piece | √
4.48 | Vascular graft for arterio-venous fistula (6mm x 40cm) | piece | √
4.49 | Vascular graft for arterio-venous fistula (5mm x 40cm) | piece | √
4.5 | Bone wax 2.5gm | piece | √
4.51 | Ommaya reservoir set: burr hole style with dome size 2.5cm and with one ventricular catheter | piece | √
4.52 | Hydrocellular foam heel dressing | piece | √
5 | Instrunments
5.1 | Adson Hypophyseal Forceps, stainless: 5 mm diameter 20.5 cm | Piece | √
5.2 | Adson Hypophyseal Forceps, stainless: 8 mm diameter 20.5 cm | Piece | √
5.3 | Hypophyseal forceps – 3 mm diameter 17 cm | Piece | √
5.4 | Specula for transphenoidal hypophysectomy diameter 15 mm x 70 mm | Piece | √
5.5 | Pituitary curette- horizontal 45 °, 18.5 cm | Piece | √
5.6 | Pituitary curette- horizontal 90 °, 18.5 cm | Piece | √
5.7 | Pituitary curette- vertical 45 °, 18.5 cm | Piece | √
5.8 | Pituitary enucleator blunt 24.5 cm | Piece | √
5.9 | Bipolar Forceps (non-stick), Bayonet - tip 0.8 mm, 20 cm | Piece | √
5.1 | Bipolar Forceps (non-stick), Bayonet right curved - 45 degree tip 0.8, 21cm | Piece | √
5.11 | Bipolar Forceps (non-stick), Curved upward - tip 1.8 mm, 20 cm | Piece | √
5.12 | Brain spatula, Malleable Large 6 mm with 20 cm | Piece | √
5.13 | Brain Spatula, Malleable Medium 4mm with 20cm | Piece | √
5.14 | Brain Spatula, Malleable Small 2mm with 20cm | Piece | √
5.15 | Cloward Kerrison Rongeur- 40° angled 3.5 mm bite, 17 cm | Piece | √
5.16 | Complete clip set with the sterilizable metal case for all the clip sizes | Piece | √
5.17 | Cranial Elevator Stainless – Large 21.5 cm, 2 cm, 8.5 mm | Piece | √
5.18 | Cranial Elevator Stainless – Small 21.5 cm, 1.7 cm, 7 mm | Piece | √
5.19 | Dissecting tumor and tissue grasoing forcep, size-30.48 cm | Piece | √
5.2 | Double action Rongeur- Angular, bite 5 mm, length 22 cm | Piece | √
5.21 | Dura forceps- 12 cm | Piece | √
5.22 | Dura Hook – 13.5 cm | Piece | √
5.23 | Dura Hook – 2 hooks 13.5 cm | Piece | √
5.24 | Dura Raspatorium, Elevator stainless – A length 20 cm | Piece | √
5.25 | Dura Raspatorium, Elevator stainless – B length 18 cm | Piece | √
5.26 | Dura Raspatorium, Elevator stainless – C length 21 cm | Piece | √
5.27 | Roux retractors double-ended stainless steel, with curved blades: 3 sizes, size 1 with handle length 14 cm with one blade 22 x 20 mm and other blade 28 x 25 mm; size 2 handle length 15 cm with one blade 26 x 23 mm and other blade 38 x 29 mm; size 3 with handle length 16.5 cm and one side blade 29 x 29 mm and other side blade 44 x 34 mm. | Set | √
5.28 | Baby-mixter dissecting and ligature forceps | Piece | √
5.29 | Tangential occlusion clamps 25.5 cm | Piece | √
5.3 | Multipurpose vascular clamps with toothing 15 cm long | Piece | √
5.31 | Tunnelling instruments for bypass grafting 10 mm diameter tunnelling tube and 680 mm length with handle | Piece | √
5.32 | Intimal spatulas and dilator: Robb 20.5 cm long with tip diameter 2 mm and 4 mm flat side | Piece | √
5.33 | Vascular dilators: Debakay 19 cm long, malleable diameters 0.5 mm, 1 mm, 1.5mm, 2mm, 2.5 mm, 3mm, 3.5mm, 4mm, 4.5mm, 5mm, 7mm, 9mm. | Piece | √
5.34 | Gelpi laminectomy retractors 18.5 cm | Piece | √
5.35 | Vascular scissors-dissecting with 1 probe-pointed blade 12 cm long | Piece | √
5.36 | Hudson Brace with Cranial Burs and drills, Brace, with snap-luck chuck, length 25.5 cm stainless with following parts: | Set | √
5.37 | Bur with 12 blades, 8 cm, 15 cm width | Piece | √
5.38 | Bur with 8 blades, 8 cm, 15 cm width | Piece | √
5.39 | Bur with 8 blades, 11 cm, 15 cm width | Piece | √
5.4 | Drill 7 cm, 2 mm width | Piece | √
5.41 | Drill 8 cm, 3 mm width | Piece | √
5.42 | Perforator drill – 11 cm,15 mm width | Piece | √
5.43 | Perforator drill - 8cm,15mm width | Piece | √
5.44 | Micro Forceps - 11cm | Piece | √
5.45 | Micro needle holders, SS - Length 13 cm with catches | Piece | √
5.46 | Micro needle holders, SS - Length 13 cm without catches | Piece | √
5.47 | Micro Scissors Slightly curved, SS – 13 cm | Piece | √
5.48 | Penfield dura dissector, SS 19.5 cm length | Piece | √
5.49 | Penfield dura dissector, SS 21.5 cm length | Piece | √
5.5 | Penfield dura dissector, SS 24 cm length | Piece | √
5.51 | Penfield dura dissector, SS 29 cm length | Piece | √
5.52 | Raney Carbon steel clips | Piece | √
5.53 | Raney Disposable plastic clips | Piece | √
5.54 | Raney Scalp clip applier, SS | Piece | √
5.55 | Raney scalp clip Frame - length 19 cm, stainless | Piece | √
5.56 | Raspatorium double ended, SS - Large 21cm, 16 mm, 7 mm | Piece | √
5.57 | Raspatorium double ended, SS - Small 14 cm, 10.5 cm, 4.5 mm, | Piece | √
5.58 | Retractors, Beckman – Adson, with Hinged arms, SS, 5 prong, blunt blades length 29 cm 3 cmW - 1 pair, 5 prong, blunt blades length 24.5 cm 3 cmW, 2.8 cmH – 1 pair | Piece | √
5.59 | Rongeur Small type, Curved, SS - Length 15 cm | Piece | √
5.6 | Self retaining retractors (titanium) Double – 37 cm | Piece | √
5.61 | Self retaining retractors (titanium) Double – 42 cm | Piece | √
5.62 | Self retaining retractors (titanium) Single – 37 cm | Piece | √
5.63 | Self retaining retractors (titanium) Single – 42 cm | Piece | √
5.64 | Suction Tube, Frazlersuction Tube, Malleable – 1.5 mm diameter, 23 cm | Piece | √
5.65 | Suction Tube, Frazlersuction Tube, Malleable – 2.5 mm diameter, 23 cm | Piece | √
5.66 | Suction Tube, Frazlersuction Tube, Malleable – 3.5 mm diameter, 23 cm | Piece | √
5.67 | Suction Tube, Frazlersuction Tube, Malleable – 5 mm diameter, 23 cm | Piece | √
5.68 | Sugita dissectors silver tip, malleable L-18 cm | Piece | √
5.69 | Sugita dissectors silver tip, malleable L-23 cm | Piece | √
5.7 | Sugita dissectors silver top, slim type malleable L-18 cm | Piece | √
5.71 | Sugita dissectors silver top, slim type malleable L-23 cm | Piece | √
5.72 | Tumor forceps - Large 10 mm width 18 cm | Piece | √
5.73 | Tumor forceps - Small 7 mm width 18 cm | Piece | √
5.74 | Venticulo Peritonial (VP) Shunt System; Malleable catheter passer length 55 cm - Child | Piece | √
5.75 | Venticulo Peritonial (VP) Shunt System; Malleable catheter passer length 90 cm to 120 cm - Adult | Piece | √
5.76 | Ventricular Needle - Length 12.5 cm, 1.5 mm diameter | Piece | √
5.77 | Ventricular Needle (silver) - Length 13.5 cm | Piece | √
5.78 | Ventricular Puncture Needle - Large 3 mm length 12.5 cm | Piece | √
5.79 | Ventricular Puncture Needle - Small 2 mm length 12.5 cm | Piece | √
5.8 | Wire Saw Guide Single, SS - 5mm width 30cm | Piece | √
5.81 | Wire Saw Guide Single, SS - 6mm width 30cm | Piece | √
5.82 | Wire Saw, Olive-krona, SS - length 35cm | Piece | √
5.83 | Gigli saw handle, SS | Piece | √
5.84 | Aneurysm Clip applying forceps with locking device-Angled 30 °. | Piece | √
5.85 | Aneurysm Clip applying forceps with locking device-Angled 15 °. | Piece | √
5.86 | Aneurysm Clip removal-Angled 30° | Piece | √
5.87 | Aneurysm Clip removal straight | Piece | √
5.88 | Hydrocellular cavity dressing | Piece | √
5.89 | Hydroactive gel for slough removal 30gm Tube | piece | √
5.9 | Hydrocellular sacral wound dressing | Piece | √
5.91 | Saline solution spoits-14 cm long | Piece | √
6 | Urology
6.1 | Equipments
6.1.1 | Karl stroz light source 485B stand cold light, Power supply: 220V-230V,50Hz, Power cord of round 3 pin plug top and suitable Stabilizer | UE | set | √
6.1.2 | Telescope 30 deg. 27005 B | UE | Piece | √
6.1.3 | Telescope zero deg. 27005 A | UE | piece | √
6.1.4 | Hamou Endoment for Laproscopy and Hysteroscopy, Power supply: 220V-230V,50Hz, Note: All Power cord should be of round 3 pin plug top. suitable stabilizer with all its accessories should be provided | LE | set | √
6.2 | Consumables
6.2.1 | 27651B cleaning brush, flexible, long for6 Fr channel | UE | Piece | √
6.2.2 | Ball electrode, 4Fr, length 73 cm | UE | Piece | √
6.2.3 | Ball electrodes,length53cm :27160AA - 3fr, 27160A -4fr, 27160B- 5fr, 27160D - 7fr ,27160E - 8fr, 27160F-10fr | UE | Piece | √
6.2.4 | Bard U.S.C poly urethome ureteral catheter 4fr | UE | Piece | √
6.2.5 | Bard U.S.C poly urethome ureteral catheter 3fr | UE | Piece | √
6.2.6 | Bard U.S.C poly urethome ureteral catheter 5fr | UE | Piece | √
6.2.7 | Cidex tray with cover (10 litres capacity) | UE | Piece | √
6.2.8 | Glycine solution (Iso-osmotic) 1.5% | UE | litres | √
6.2.9 | Hand piece for Valley Lap (Ligasure) vessel sealing system R3M859702 LS 2070 | UE | Piece | √
6.2.10 | Integral ureter stent (Double J stent), size 26 mm, Ch.5 | UE | Piece | √
6.2.11 | Integral ureter stent (Double J stent), size 26 mm, Ch.6 | UE | Piece | √
6.2.12 | Integral ureter stent (Double J stent), size 26 mm, Ch.8 | UE | Piece | √
6.2.13 | Pig tail percutanous nephrostomy catheter with needle (romsons)size:5,6,7 | UE | Piece | √
6.2.14 | Silicone foleys catheter, size - 14 fr | OT G | Piece | √
6.2.15 | Silicone foleys catheter, size - 16 fr | OT G | Piece | √
6.2.16 | Silicone foleys catheter, size - 18 fr | OT G | Piece | √
6.2.17 | Suprapubic catheterisation set disposable, catheter 10M1Cr 12Ch | UE | Piece | √
6.2.18 | Toomey syringe glass | UE | Piece | √
6.2.19 | Urostomy pouch with Colostomy WAFFER - Size 45mm |  | Piece | √
6.2.20 | Urostomy pouch with Colostomy WAFFER - Size 57mm |  | Piece | √
6.3 | Instrunments
6.3.1 | 27023FE grasping forceps for small fragments single action jaw, flexible, 5Fr. Length 73 cm | UE | Piece | √
6.3.2 | 27023VK stone basket,5Fr. Length 60 cm consisting of: 1x27023 VT3-ring handle, 3 x 27023VR basket ,3x27023VS coil | UE | set | √
6.3.3 | 27023ZE biopsy forceps,single action jaw,flexible,5Fr. Length 73 cm | UE | piece | √
6.3.4 | 27040BK resectoscope sheat with LUER lock stopcock,including connecting tube for irrigation,24fr,short beak with obturator 270400D colour code yellow. | UE | set | √
6.3.5 | 27040SI resectoscope sheath including connecting tubes for in and outflow, 2 LUER-lock adaptors 26fr ,oblique back,inner tube with ceramic insulation,for use with working element 27040 L/E and 27050 C/D/E/EN colour code yellow | UE | set | √
6.3.6 | 27074B stone crushing forceps, single action jaw. 27074SZ adapters for use with resectoscope sheath, 2704SL/SD/SE/SM outer sheath only. | UE | set | √
6.3.7 | 27078AJ cystoscopic rongeur 24Fr,single action jaw with obturator 270778AO,including 2 LUER lock adapter for continous suction and irrigation | UE | set | √
6.3.8 | Biopsy forceps, double action jaws, flexible 5fr,length 40cm (271562H) | UE | Piece | √
6.3.9 | Biopsy forceps, through cutting, double action jaw for unipolar coagulation (27072DD) | UE | Piece | √
6.3.10 | Catheter deflecting mechanism with 1instrument channel with rachet optical(27026F) | UE | piece | √
6.3.11 | Cysto-urethescope sheath 24fr | UE | piece | √
6.3.12 | Half cut trocar for urlolgy, 5mm | UE | piece | √
6.3.13 | Half cut trocar for urlolgy, 6mm | UE | piece | √
6.3.14 | Half cut trocar for urlolgy, 8mm | UE | piece | √
6.3.15 | Instrument channel 6Fr for adult.Direction of view 0degree,angle of view110 degree,working length 35cmand OD;15.5Fr | UE | piece | √
6.3.16 | Lithotrite 27076A incorporating both handle and turning screw action 24fr including 2 LUER lock adaptors | UE | Piece | √
6.3.17 | Optical jawbiopsy forceps,double action jaws (27072BL) | UE | Piece | √
6.3.18 | SACHSE MEATOTOME with scale to set cutting sizes 10fr-36fr | UE | Piece | √
6.3.19 | Stone punch:27077A punch working element,27077B punch sheat with central valve,including connecting tubes for suction band irrigation25Fr, straight beak obturator270077DO | UE | set | √
6.3.20 | Stonebasket: basket with 6 wires,5fr (27186A), store basket 5fr,length 60cm (27023VK), ZEISS urethral stone dislodger 5fr(27190) | UE | set | √
6.3.21 | Urethroscope sheath size 22 fr with obturator(026B) | UE | Piece | √
6.3.22 | Urethroscope sheath size17fr with obturator(27026U) | UE | Piece | √
6.3.23 | Urethroscope sheath size19 fr with obturator(27026D) | UE | Piece | √
6.3.24 | Working Elements(27050EH),sets consisting of:27050E 1 working element also to use with optical urethrotome, 2 cuttuing loops (27050), 2 coagulating electrodes(27050), 2 high frequency cords (277)and 1 protection tubes(280) | UE | set | √
6.4 | Accessories
6.4.1 | 11025E pressure compensation cap for ventilation during gas sterlization | UE | Piece | √
6.4.2 | 27170 connecting cord for electrodes | UE | Piece | √
6.4.3 | 27224 ELLIK evacuator | UE | Piece | √
6.4.4 | 27770E coagulating electrode 8Fr,unipolar,length53cm | UE | Piece | √
6.4.5 | 495NL fibre optic light cable size3.5mm, length180cm | UE | Piece | √
6.4.6 | Adapter for use with resectoscope sheaths (27094) | UE | Piece | √
6.4.7 | Argon gas | UE | litres | √
6.4.8 | Coagulating electrode 8fr(27770E) | UE | Piece | √
6.4.9 | Coagulating electrode,pointed 24fr.colour code yellow (27050L) | UE | Piece | √
6.4.10 | Cold knife,straight (27069K) | UE | Piece | √
6.4.11 | cutting loop,angled 24fr.colour code yellow(27050 G) | UE | Piece | √
6.4.12 | Luovik cold knife,straight waveform(27969M) | UE | Piece | √
6.4.13 | Telescopes bridge 27025G with one instrument channel | UE | Piece | √
7 | Heamodialysis
7.1 | 5 micron filter compatible with R/O machine | Piece | √ | √
7.2 | AV Fistula needle 15G | Piece | √ | √
7.3 | AV Fistula needle (16G X 1) | Piece | √ | √
7.4 | Blood line with bubble catcherpart No.N-A00/V003 with filter 8mm | Set | √ | √
7.5 | Haemodialysis machine, Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable UPS | Set | √ | √
7.6 | Dialyzer Polysulfone F5, F6, F7, compatible to Dialysis Machine | Piece | √ | √
7.7 | Diasafe filter plus, compatible with Dialysis machine | Piece | √ | √
7.8 | Double Lumen Catheter (DLC) Jugular 13.5 or 14.5 cm.11.5F, with 18G introducer needle 1no, injection cap 1no | Set | √ | √
7.9 | DLC Jugular 13.5 or 14.5 cm.11F (3.7 mm) x 15 cm, with 18G introducer needle 1no, injection cap 1no | Set | √ | √
7.1 | DLC Subclavian set with wings 16cm, 11.5F (flexible guide wire 1no, Catheter with wings 1 no, 12F dilator 1no, 18G Introducer needle 1no, Injection cap 1no.) | Set | √ | √
7.11 | DLC Femoral 19.5 cm, 11.5F (flexible guide wire 1no, Catheter with wings 1 no, 12f dilator 1no, 18G Introducer needle 1no, Injection cap 1no.) | Set | √ | √
7.12 | Frekaderm spray 250 ml | Piece | √ | √
7.13 | Frekaderm spray 500ml | Piece | √ | √
7.14 | Infusion Line set (i/v set) | Set | √ | √
7.15 | Leukopore 2.5cm X 9.14m | Roll | √ | √
7.16 | Leukopore 5cm X 9.14 m | Roll | √ | √
7.17 | Micro shield hand rub 250 ml | Piece | √ | √
7.18 | Micropore 1" (2.5 cm x 9.14 m) | Roll | √ | √
7.19 | Tranpore 0.5" | Roll | √ | √
7.2 | Port cap (Dialyzer cap -rubber) | Piece | √ | √
7.21 | Reverse Osmosis (RO) Water Hardness testing kit | Piece | √ | √
7.22 | Syringe 20 ml for heparin pump (disposible) | Piece | √ | √
7.23 | Syringe 50 ml (disposable) | Piece | √ | √
7.24 | Tegaderm (6 cm X 7cm), Sterile (Transparent film dressing, frame style) | Piece | √ | √
7.25 | Tegederm 15cmx20cm, Sterile (Transparent film dressing, frame style) | Piece | √ | √
7.26 | Tran spore 1.25cm X 9.1m | Roll | √ | √
7.27 | Transducer A/V pressure monitor compatible with Dialysis machine | Piece | √ | √
7.28 | Peritoneal Dialysis transfer set, compatible with Peritoneal Dialysis machine | Set | √ | √
7.29 | Peritoneal Dialysis Catheter set (Adult), compatible with Peritoneal Dialysis machine | Set | √ | √
7.3 | Stainless Steel dressing tray with cover, size: small | Piece | √ | √
7.31 | Stainless Steel dressing tray with cover, size: medium | Piece | √ | √
7.32 | Twist lock ISO fitting [MPC-45] dialyzer blood port cap | Piece | √ | √
7.33 | Electronic weighing scale,capacity 100 kg [DIGITAL] | Piece | √ | √
7.34 | AMSD- 650 Dialyzer [Gamma sterlized] | Piece | √ | √
7.35 | Micron filter compatible | Piece | √ | √

## GC
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
1 | General Consumables
1.1 | General Consumables
1.1.1 | Catheter, Foley 2 ways (Siliconized) - size 6 | Piece | √ | √ | √ | √
1.1.2 | Catheter, Foley 2 ways (Siliconized)- size 8 | Piece | √ | √ | √ | √
1.1.3 | Catheter, Foley 2 ways (Siliconized)- size 10 | Piece | √ | √ | √ | √
1.1.4 | Catheter, Foley 2 ways (Siliconized) - size 12 | Piece | √ | √ | √ | √
1.1.5 | Catheter, Foley 2 ways (Siliconized)- size 14 | Piece | √ | √ | √ | √
1.1.6 | Catheter, Foley 2 ways (Siliconized)- size 16 | Piece | √ | √ | √ | √
1.1.7 | Catheter, Foley 2 ways (Siliconized) - size 18 | Piece | √ | √ | √ | √
1.1.8 | Catheter, Foley 2 ways (Siliconized)- size 20 | Piece | √ | √ | √
1.1.9 | Catheter, Foley 2 ways (Siliconized) - size 22 | Piece | √ | √ | √
1.1.10 | Catheter, Foley 3 ways (siliconized) - size 14 | Piece | √ | √
1.1.11 | Catheter, Foley 3 ways (Siliconized) - size 16 | Piece | √ | √
1.1.12 | Catheter, Foley 3 ways (Siliconized) - size 18 | Piece | √ | √
1.1.13 | Catheter, Foley 3 ways (Siliconozed) - size 20 | Piece | √ | √
1.1.14 | Catheter, Foley 3 ways (Siliconized) - size 22 | Piece | √ | √
1.1.15 | Catheter, Plain (Siliconized) - size 5 | Piece | √ | √ | √ | √
1.1.16 | Catheter, Plain (Siliconized) - size 6 | Piece | √ | √ | √ | √
1.1.17 | Catheter, Plain (Siliconized) - size 8 | Piece | √ | √ | √ | √
1.1.18 | Catheter, Plain (Siliconized) - size 10 | Piece | √ | √ | √ | √
1.1.19 | Catheter, Plain (Siliconized) - size 12 | Piece | √ | √ | √ | √
1.1.20 | Catheter, Plain (Siliconized) - size 14 | Piece | √ | √ | √ | √
1.1.21 | Catheter, Plain (Siliconized) - size 16 | Piece | √ | √ | √ | √
1.1.22 | Catheter, suction (PVC), with X-Ray opaque line - size 6 | Piece | √ | √ | √ | √
1.1.23 | Catheter, suction (PVC), with X-Ray opaque line - size 8 | Piece | √ | √ | √ | √
1.1.24 | Catheter, suction (PVC), with X-Ray opaque line- size 10 | Piece | √ | √ | √ | √
1.1.25 | Catheter, suction (PVC), with X-Ray opaque line - size 12 | Piece | √ | √ | √ | √
1.1.26 | Catheter, suction (PVC), with X-Ray opaque line- size 14 | Piece | √ | √ | √ | √
1.1.27 | Catheter, suction (PVC), with X-Ray opaque line- size 16 | Piece | √ | √ | √ | √
1.1.28 | Catheter, suction (PVC), with X-Ray opaque line- size 18 | Piece | √ | √ | √ | √
1.1.29 | Chest drainage tube - size 10 | Piece | √ | √ | √
1.1.30 | Chest drainage tube - size 12 | Piece | √ | √ | √
1.1.31 | Chest drainage tube - size 14 | Piece | √ | √ | √
1.1.32 | Chest drainage tube - size 16 | Piece | √ | √ | √
1.1.33 | Chest drainage tube - size 18 | Piece | √ | √ | √
1.1.34 | Chest drainage tube - size 20 | Piece | √ | √ | √
1.1.35 | Chest drainage tube - size 22 | Piece | √ | √ | √
1.1.36 | Chest drainage tube - size 24 | Piece | √ | √ | √
1.1.37 | Chest drainage tube - size 26 | Piece | √ | √ | √
1.1.38 | Chest drainage tube - size 28 | Piece | √ | √ | √
1.1.39 | Chest drainage tube - size 30 | Piece | √ | √ | √
1.1.40 | Chest drainage tube - size 32 | Piece | √ | √ | √
1.1.41 | Chest drainage bag water seal, 1000 ml | Piece | √ | √ | √
1.1.42 | Colostomy bag (transparent) with WAFFER (Flange) tail closer (clip), perforated marking size of 25mm - 64mm | Set | √ | √ | √
1.1.43 | Condom catheter (male external device) - Small | Piece | √ | √ | √ | √
1.1.44 | Condom catheter (male external device)- Medium | Piece | √ | √ | √ | √
1.1.45 | Condom catheter (male external device) - Large | Piece | √ | √ | √ | √
1.1.46 | Disposable Wooden Tongue Depressor with paper cover | Piece | √ | √ | √ | √
1.1.47 | Disposable Sharp Container | Piece | √ | √ | √ | √
1.1.48 | Disposable Face mask, three ply with back-tie | Piece | √ | √ | √ | √
1.1.49 | ECG paper 50mmx20 metre for 6108T machine | Roll | √ | √ | √
1.1.50 | ECG Paper Schiler AT/AT-2/Plus | Packet | √ | √
1.1.51 | Gel, high conductivity for ECG | Piece | √ | √ | √
1.1.52 | AQUA Gel for USG, CTG, fetal Doppler | Piece | √ | √ | √ | √
1.1.53 | Adhesive Cardiac Electrode pad (disposable) | Piece | √ | √ | √
1.1.54 | Gauze Than 36"x20 yds | Packet | √ | √ | √ | √
1.1.55 | Disposable Gloves non-sterile - small | Pair | √ | √ | √ | √
1.1.56 | Disposable Gloves non-sterile - medium | Pair | √ | √ | √ | √
1.1.57 | Disposable Gloves non-sterile - large | Pair | √ | √ | √ | √
1.1.58 | Gloves sterile, latex with plastic wrap - size 6.5 | Pair | √ | √ | √ | √
1.1.59 | Gloves sterile, latex with plastic wrap - size 7 | Pair | √ | √ | √ | √
1.1.60 | Gloves sterile, latex with plastic wrap - size 7.5 | Pair | √ | √ | √ | √
1.1.61 | Gloves sterile, latex with plastic wrap - size 8 | Pair | √ | √ | √ | √
1.1.62 | Gloves sterile, latex, elbow-length, size: 6.5 | Pair | √ | √
1.1.63 | Gloves sterile, latex, elbow-length, size: 7 | Pair | √ | √
1.1.64 | Gloves sterile, latex, elbow-length, size: 7.5 | Pair | √ | √
1.1.65 | Gloves sterile, latex, elbow-length, size: 8 | Pair | √ | √
1.1.66 | Gluco strip compatible with Glucometer | 50strips/ pkt | √ | √ | √ | √
1.1.67 | IV infusion dripSet, Specification: 3mm non-toxic PVC tubing with big PVC moulded chamber with 15mm spike cover filter with airvent spike | Piece | √ | √ | √ | √
1.1.68 | Stop Cock, 3 ways connector for IV drips. | Piece | √ | √ | √
1.1.69 | IV Cannula 2 ways, size 14 with injecting port | Piece | √ | √ | √ | √
1.1.70 | IV Cannula 2 ways, size 16 with injecting port | Piece | √ | √ | √ | √
1.1.71 | IV Cannula 2 ways, size 18 with injecting port | Piece | √ | √ | √ | √
1.1.72 | IV Cannula 2 ways, size 20 with injecting port | Piece | √ | √ | √ | √
1.1.73 | IV Cannula 2 ways, size 22 with injecting port | Piece | √ | √ | √ | √
1.1.74 | IV Cannula 2 ways, size 24 with injecting port | Piece | √ | √ | √ | √
1.1.75 | IV Cannula 2 ways, size 26 with injecting port | Piece | √ | √ | √
1.1.76 | IV Cannula 2 ways, size 24 with injecting port without wings | Piece | √ | √
1.1.77 | IV Cannula 2 ways, size 26 with injecting port without wings | Piece | √ | √
1.1.78 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned without side opening at the tip, Length 52cm - size 4 FG | Piece | √ | √
1.1.79 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip, distal end coned without side opening at the tip, Length 52cm - size FG 5 | Piece | √ | √ | √ | √
1.1.80 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned without side opening at the tip, Length 52cm - size FG 6 | Piece | √ | √ | √ | √
1.1.81 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes, Length 52cm - size FG 8 | Piece | √ | √ | √ | √
1.1.82 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes, Length 52cm - size FG 10 | Piece | √ | √ | √ | √
1.1.83 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes - size FG 12 | Piece | √ | √ | √ | √
1.1.84 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes- size FG 14 | Piece | √ | √ | √ | √
1.1.85 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes - size 16 | Piece | √ | √ | √ | √
1.1.86 | Nasogastric tube (Siliconized) with funnel and luer connector with Radio-opaque line and tip,distal end coned with two lateral eyes- size FG 18 | Piece | √ | √ | √ | √
1.1.87 | Rectal Tube (Flatus), size 6 | Piece | √ | √ | √ | √
1.1.88 | Rectal Tube (Flatus), size 8 | Piece | √ | √ | √ | √
1.1.89 | Rectal Tube (Flatus), size 10 | Piece | √ | √ | √ | √
1.1.90 | Rectal Tube (Flatus), size 12 | Piece | √ | √ | √
1.1.91 | Rectal Tube (Flatus), size 14 | Piece | √ | √ | √
1.1.92 | Rectal Tube (Flatus), size 16 | Piece | √ | √ | √
1.1.93 | Rectal Tube (Flatus), size 22 | Piece | √ | √ | √
1.1.94 | Rubber (Mackintosh) roll of 25 meter | 25mtr/ roll | √ | √ | √ | √
1.1.95 | Condom | Gross | √ | √ | √ | √
1.1.96 | Urine Drainage Bag with T-type, bottom outlet and sampling port, capacity:1000/2000ml | Piece | √ | √ | √ | √
1.2 | GC- Dressing
1.2.1 | Adhesive plaster 7.5cm x 5mts | Roll | √ | √ | √ | √
1.2.2 | Bandage rolled 5cm x 3mtr Length | Pkt of 12 | √ | √ | √ | √
1.2.3 | Bandage rolled 10cm x 3mtr Length | Pkt of 12 | √ | √ | √ | √
1.2.4 | Bandage rolled 15cm x 3mtr Length | Pkt of 12 | √ | √ | √ | √
1.2.5 | Cotton (zig-zag) 400g | Packet | √ | √ | √ | √
1.2.6 | Elastic Adhesive Bandage 10cm x 6mts (stretch length), flesh colour with fast edges | Packet | √ | √ | √
1.2.7 | Adhesive surgical tape (paper) 25mm x 9.1 mts | Roll | √ | √ | √
1.2.8 | Adhesive surgical tape (paper) 50mm x 9.1 mts | Roll | √ | √
1.2.9 | Adhesive surgical tape (paper) 75mm x 9.1 mts | Roll | √ | √
1.2.10 | Transparent dressing (Tegaderm) :15 cm x 20 cm/6 inch x 8 inch | Piece | √ | √
1.2.11 | Transparent dressing (Tegaderm) :6 cm x 7cm/ 2 3/8inch x 2 3/4inch. | Piece | √ | √
1.2.12 | Transparent dressing (Tegaderm) :4.3 cm x 4.3cm/ 2 3/8inch x 2 3/4inch. | Piece | √ | √
1.3 | GC- Linen
3.1 | Hand towel 12 inch x 12 inch - white | Piece | √ | √ | √ | √
3.2 | Patient Trolley matttress, foam with water proof synthetic cover 72 inch x 21 inch x 2 inch with strips at both ends | Piece | √ | √ | √
3.3 | Mattress, coir with water proof synthetic cover : size length 78" x Width 37" x Breadth/thickness 4" | Piece | √ | √ | √ | √
3.4 | water proof synthetic mattress cover : size length 78" x Width 37" x Breadth/thickness 4" | Piece | √ | √ | √ | √
3.5 | Pillow cover (cotton): size 25 inch x 17 inch - Sky Blue | Piece | √ | √ | √ | √
3.6 | Synthetic foam pillow, washable: size 24 inch x 16 inch, 1.5 kg | Piece | √ | √ | √ | √
3.7 | Bed sheet with elastic band at four corners Size: length 78" x Width 37" x Breadth/thickness 4" | Piece | √ | √ | √ | √
3.8 | Wollen Blanket: Green colour. Weight-2kg Size: 60 inches × 90 inches.composition 80% wool and 20% other fibre | Piece | √ | √ | √ | √
3.9 | Draw sheet: size 72 inch x 36 inch - Green Colour | Piece | √ | √ | √
1.4 | GC-Utility
1.4.1 | Apron: plastic with cloth piping, back and neck tie, standard size | Piece | √ | √ | √ | √
1.4.2 | Apron plastic with cloth piping with back and neck tie. Full sleeves with elastic band at wrist. Size small | Piece | √ | √
1.4.3 | Apron plastic with cloth piping with back and neck tie. Full sleeves with elastic band at the wrist. Size Medium | Piece | √ | √
1.4.4 | Apron plastic with cloth piping with back and neck tie. Full sleeves with elastic band at wrist. Size Large | Piece | √ | √
1.4.5 | Polypropylene Bio-hazard bag, Autoclavable and Bio-degradable with biohazard symbol, size: Length 87cm x Breadth 75cmltr (Red colour). Thickness not less than 9micron. | Roll | √ | √ | √ | √
1.4.6 | Polypropylene, Bio-degradable: size Length 87cm x breadth 75cm (Green colour). Thickness not less than 9micron. | Roll | √ | √ | √ | √
1.4.7 | Baby liquid medicine dispencer. Maximum capacity 1 tsp/1ml. Simple | Piece | √ | √
1.4.8 | Digital forehead (Sensor) medical thermometer: Forehead Thermometer with sensor; High accuracy; 6-8 seconds Measurement; Memory function; LCD Display; Fever alarm; Batteries included | Piece | √ | √ | √ | √
1.4.9 | Digital Thermometer with battery |  | √ | √ | √ | √
1.4.10 | Safety goggles: With high impact resistance, panoramic and crystal clear view. Glare reducer. Anti scratch properties. | Piece | √ | √ | √
1.4.11 | Face mask N95: Comfortable, Lightweight. Particulate respirator with 95% filtration efficiency. Contoured, aluminium nose Piece with foam | Piece | √ | √ | √
1.4.12 | Bone marrow aspiration Needle, sterile - size 16G | Set | √ | √
1.4.13 | Bone marrow aspiration Needle, sterile - size 18G | Set | √ | √
1.4.14 | Bone marrow aspiration Needle, sterile - size 20 G | Set | √ | √
1.4.15 | Paraffin gauze dressing, 10 x 10 cm, sterile | Pkt of 10 dressings | √ | √ | √
1.4.16 | Paraffin gauze dressing, 15 x 15 cm, sterile | Pkt of 10 dressings | √ | √ | √
1.4.17 | Paraffin gauze dressing, 20 x 20 cm, sterile | Pkt of 10 dressings | √ | √ | √
1.4.18 | Paraffin gauze dressing, 30 x 30 cm, sterile | Pkt of 10 dressings | √ | √ | √
1.4.19 | Tourniquet | Piece | √ | √ | √
1.4.20 | Wrist band for patient- Paediatric and neonate: Precision ID Bracelets, Insert Style. | Piece | √ | √
1.4.21 | Wrist band for patient- Adult: Precision ID Bracelets, Insert Style. | Piece | √ | √
1.4.22 | Needle 16G (Disposable) | Piece | √ | √ | √ | √
1.4.23 | Needle 18G (Disposable) | Piece | √ | √ | √ | √
1.4.24 | Needle 20G (Disposable) | Piece | √ | √ | √ | √
1.4.25 | Needle 22G (Disposable) | Piece | √ | √ | √ | √
1.4.26 | Needle 24G (Disposable) | Piece | √ | √ | √ | √
1.4.27 | Syringe with needle 1ml (Disposable) | Piece | √ | √ | √ | √
1.4.28 | Syringe with needle 2ml (Disposable) | Piece | √ | √ | √ | √
1.4.29 | Syringe with needle 5ml (Disposable) | Piece | √ | √ | √ | √
1.4.30 | Syringe 10ml with 21G needle (Disposable) | Piece | √ | √ | √ | √
1.4.31 | Syringe without needle 20ml (Disposable) | Piece | √ | √ | √ | √
1.4.32 | Syringe without needle 50ml (Disposable) | Piece | √ | √ | √ | √
1.4.33 | Cover all: Large | Packet | √ | √
1.4.34 | Cover all: Medium | Packet | √ | √
1.4.35 | Cover all: Small | Packet | √ | √
1.4.36 | Face shield | Piece | √ | √
1.4.37 | Piles binding rubber bands | Piece | √ | √
2 | General Equipment
2.1 | Air mattress: Electric, 220V, 50Hz with pump, PVC 3ply, thickness 0.3-0.4mm. | Set | √ | √ | √
2.2 | Ampoule cutter: Anodised Aluminum.Silicone (O-ring). Stainless steel spring, Compatible ampoule sizes* 1-15ml . Slim light weight design, pocket size | Piece | √ | √ | √ | √
2.3 | Bed pan plastic: Type plastic be perfection with lid. Length 35.0cm, width 28.5cm, front height 6.0cm, rear height 8.5cm | Piece | √ | √ | √ | √
2.4 | Bed Fowler's with wheel -Semi/half: Size, 206cm (L) X90cm (W) x 60 cm (H) function: one function movements and controls. Back lift manual, Top, two sections. Tubular head and foot bows of unequal height. Frame: mild steel, finish with epoxy power coated, stainless steel panels with provisions for IV rod in two sections & wheels | Piece | √ | √ | √ | √
2.5 | Bed Fowler's with wheel -Full: Overall size: 78" L x36"W x 24" H ( 206L x 90W x60 H cms ) Back rest and knee rest adjustable maneuvered by 2 separate screw handle from foot end. Four section perforated CRCA MS sheet top. Provisions for stainless steel IV rod with four locations. Removable stainless steel tubular head & foot bows with laminated panels. Bed mounted on 100mm Dia castor two with brakes & wheels. Pre-teated epoxy powder coated. | Piece | √ | √ | √
2.6 | Bed Fowler's with wheel- Full : Orthpedic patient bed with balken beam and pulleys with adjustable height | Piece | √ | √
2.7 | Bedside Locker: Full Body SS (Stainless Steel) size: 16"x 16" x 32" HT 16" x 16" x 32" HT Body machine pressed from 20 SWG M.S. sheet with three sides closed. Have one box and one drawer with SS.Mounted on 5 cm dia. swinging castors or legs on rubber shoes. Size: 46 x 46 x 81 cm. | Piece | √ | √ | √ | √
2.8 | Bedside Screen 4 fold: with spring and cloth size 168 cm H x 244 cm W. Frame work made of ERW tubular pipe in four folds Piece fold 60cms dia and two inner folds fitted with one castor Piece of 5cms dia, Supplied with spring wire and curtains, pretreated and epoxy power coated. | Piece | √ | √ | √ | √
2.9 | Burns cradle | Piece | √ | √ | √
2.1 | Cardiac Table:Rectangular Frame Work Mounted On 5cms Castors30×16 Laminated Top Adjustable By Side Knob. Pre-treated and epoxy powder coated. | Piece | √ | √ | √
2.11 | Cheatle forceps : Size, 10 inch, Stainless Steel | Piece | √ | √ | √ | √
2.12 | Cheatle forcep stand: Stainless steel | Piece | √ | √ | √ | √
2.13 | Trolley for clean linen: Polymer design, Dimension 724mmDx219mmLx1736mmH, Capacity of 1.36 cubic metres | Piece | √
2.14 | Trolley for dirty linen : Box type, Dimension: 1060x610x610mm, Stainless steel frame work with push handle at one end, mounted on 5cms castors with wheels. Body: Superior quality plastic (ABC material) | Piece | √ | √
2.15 | Dressing Drum: Size, 12 inch, Stainless Steel: Domed top and bottom with sliding removable band and drop down-hasps (stainless steel). Made of stainless steel grade 316, thickness 0.5mm. | Piece | √ | √ | √ | √
2.16 | Dressing Drum: Size, 15 inch, Stainless Steel:Domed top and bottom with sliding removable band and drop down-hasps (stainless steel). Made of stainless steel grade 316, thickness 0.5mm. | Piece | √ | √ | √ | √
2.17 | Dressing Drum: Size, 9 inch, Stainless Steel: Domed top and bottom with sliding removable band and drop down-hasps (stainless steel). Made of stainless steel grade 316, thickness 0.5mm. | Piece | √ | √ | √ | √
2.18 | Digital weighing Scales (bathroom) 120 kg | Piece | √ | √ | √ | √
2.19 | Electronic Weighing scale: LCD Display, with tray, capacity 25 kg. Tolerance +-0.5 gm, Precision minimum 10gm, Power supply: 220V-230V, Power cord of round 3 pin type plug top | Piece | √ | √ | √
2.2 | ECG Machine (Portable) : Single Channel ECG Recording, 12 leads, Last memorized ECG Recording and additional copies available, Automatic and Manual Recording Modes. Dual Power supply: Mains and Battery operation. Power supply 220-230V. Power cord with round 3 pin top . Built in recharge | Set | √ | √ | √
2.21 | ECG machine( Table Model) : Dimensions: 400x330x100 mm, Weight: 5.0 kg (5.35kg with full paper tray). Mains Supply: 100 to 115/220 to 240 VAC, 50/60 Hz. Battery: Built-in 12 V lead-acid battery (rechargeable). Battery capacity: 4 hours normal use- 300 printouts. Power Consumption: Recording 40 VA max. Leads: Standard/Cabrera. Paper speed: 5/10/25/50 mm/s (direct). Sensitivity: 5/10/20 mm/mV, either automatically adjusted or manually selected. Chart Paper: Thermoreactive – Z-folded, 210 mm wide, perforation 280 mm. Printing Process: High-resolution thermal print head, 8 dots per mm / 200 dots per inch (amplitude axis). 40 dots per mm / 1000 dots per inch (time axis 25 mm/s). Recording Tracks: 6 channels, positioned at optimal width on 200 mm, automatic baseline adjustment. Automatic Lead Programs: Data Record: Printout of all 12 leads, Listing of ECG recording data. Version C: ECG measurement results (intervals, amplitudes, electrical axis), Sokolow Index, average complexes with optional measurement reference markings, and interpretation. | Set | √ | √
2.22 | ECG Storage: Circular input memory for 10 s, 12-lead ECG. Memory Option: Memory for c.45 ECG recordings with transmission facilities over an RS-232 interface. Frequency Range of Digital Recorder: 0 to 150 Hz (IEC), 0 to 150 Hz (AHA). ECG Amplifier: Simultaneous, synchronous registration of all 9 active electrode signals (=12 standard leads). Sampling frequency: 5µV. Dynamic range: ±9.5 VAC. Max. electrode potential: ±300 mVDC. Time constant: 3.2 s. Frequency response: 0.05 to 150 Hz (-3 dB). Input impedance: >2.5MOhms at 10 Hz. Myogram Filter (muscle tremor filter): 25 Hz or 35 Hz, programmable (not active on average waveform). The stored ECGs can be printed with or without filter. Distortion-free suppression of superimposed 50 or 60 Hz sinusoidal interferences by means of an adaptive digital filter. Line Frequency Filter: Inpatient Input: Fully floating and isolated, defibrillation protected. | set | √ | √
2.23 | Safety Standard: CF according to IEC and complying with the following. RL 93/42/EEC. EN 60601-1:1990. IEC 601-1. IEC 601-2-25:1993. pr EN 1441:1994. CISPR 11: 1985, EN 55011:1992. IEC 801-2:1991, IEC 801-3:1984, IEC 801-4:1988. IEC 801-5. Safety Class: I according to IEC 601-1 (with internal power supply). IIa according to RL 93/42/EEC, CE-0123. This device is not designed for outdoor use (IP 20). Environmental Conditions: Temperature, Operating: 10° to 40°C. Temperature, Storage: -10° to 50°C. Relative humidity: 25 to 95% (non condensing). Atmospheric pressure: 700 to 1060 hpa.Power supply 220-230V. Power cord with round 3 pin top. Control Panel: Rubber keys. | set | √ | √
2.24 | Enema can SS with marking: Measuring: 50ml - 1000 ml with 3 extra tubing | Set | √ | √ | √ | √
2.25 | Examination table with mattress and Foot step: Stainless Steel with rubber topping | Set | √ | √ | √ | √
2.26 | Pulse Oximeter, Finger Tip | Piece | √ | √ | √ | √
2.27 | Forceps Dressing: Thumb ( Non-toothed): Size 8 inch | Piece | √ | √ | √ | √
2.28 | Forceps Dressing: Gross Maier : Size 8 inch | Piece | √ | √ | √ | √
2.29 | Galli pot: 100ml, Stainless Steel | Piece | √ | √ | √ | √
2.3 | Gauze cutting machine: features XD-100 can cut from a single ply of fabric, right up to 1" high, making it a very versatile - lightweight machine. Octagonal cutting blade.Built-in sharpening mechanism .Extra-thin base plate.High-power electric motor .Carbide-tipped counter blade | Piece | √ | √
2.31 | IV stand with wheels: Stainless steel adjustable with four hooks/hanger | Piece | √ | √ | √ | √
2.32 | Jug: Plastic, graduated, 1 Ltr | Piece | √ | √ | √ | √
2.33 | Kick bucket: Foot step with wheel, Stainless Steel | Piece | √ | √ | √ | √
2.34 | Kidney dish: Size: 10 inch, Stainless Steel | Piece | √ | √ | √ | √
2.35 | Kidney dish: Size: 8 inch, Stainless Steel | Piece | √ | √ | √ | √
2.36 | Knee Hammer with plastic handle | Piece | √ | √ | √ | √
2.37 | Medicine cup - graduated 30 ml, plastic | Piece | √ | √ | √ | √
2.38 | Medicine Trolley with four drawers (SS): Overall approx. Sizes : 76Lx46Wx8H cms, Frame work of MS Tube mounted on 10cms castors,Two SS shelves with three side railing on top shelf, Two drawers under the each shelf, Whole body frame SS (Stainless steel) | Piece | √ | √ | √
2.39 | Emergency crash trolley
2.4 | Nebulizer machine: Nebulizer compressor system, Spare nebulizer chambers with mask, Power supply 220-230V. Power cord with round 3 pin top | Piece | √ | √ | √ | √
2.41 | Nebulizer chamber along with tubings and mask, Adult | Piece | √ | √ | √ | √
2.42 | Non electric weighing scale with tray – capacity upto 25 kgs, tolerance ± 1 gm, Precision minimum 10 gm. | Piece |  |  | √ | √
2.43 | Oxygen trolley: B-type | Piece | √ | √ | √ | √
2.44 | Oxygen trolley: D-type | Piece | √ | √ | √ | √
2.45 | Patient chart holder | Piece | √ | √ | √ | √
2.46 | Patient transport stretcher trolley: 210Lx65Wx65Hcms, with safety side railings, IV rods, Cylinder cage, Adjustable height by hydraulic foot operated pump, 15cm dia castors, pretreated & epoxy powder coated | Piece | √ | √ | √
2.47 | Proctoscope: For adult, Stainless Steel | Piece | √ | √ | √
2.48 | Proctoscope: For Pediatric, Stainless Steel | Piece | √ | √ | √
2.49 | Piles banding forcep | Piece | √ | √
2.5 | Scissor: For slough cutting, Stainless Steel | Piece | √ | √ | √
2.51 | Scissor: Tailor, 9 inch | Piece | √ | √ | √ | √
2.52 | Scissor: Bandage, 7 inch SS | Piece | √ | √ | √ | √
2.53 | Sphygmomanometer : Dial type Single-hand operation Non-stop pin reliable manometer. Big dial face 65mm & easy to read, Latex free cuff | Set | √ | √ | √ | √
2.54 | Spotlight on stand, fibre optic | Piece | √ | √ | √ | √
2.55 | Sputum cups with lid: Size: 100ml, Stainless Steel | Piece | √ | √ | √ | √
2.56 | Stethoscope bivalve | Piece | √ | √ | √ | √
2.57 | Stretcher: Hand carrying, aluminum pole | Piece | √ | √ | √ | √
2.58 | Suction machine: Portable, Stainless steel body with wheels, high pressure tubes, pressure regulating switch. twin suction bottle of 2-3 lits. Capacity, display of vacuum from 100 to 700 mm Hg. Power supply AC 220-230V. Power cord with round 3 pin top | Piece | √ | √ | √ | √
2.59 | Suction machine (foot operated) | Piece | √ | √ | √ | √
2.6 | Suture removing scissor: Size: 8 inch | Piece | √ | √ | √ | √
2.61 | Swab container with cover: Size, 4 inch, Stainless Steel | Piece | √ | √ | √ | √
2.62 | Tray with cover: Size, 10" x 12", Stainless Steel | Piece | √ | √ | √ | √
2.63 | Tray with cover: Size, 12" x 18", Stainless Steel | Piece | √ | √ | √ | √
2.64 | Tray with cover: Size, 9" x 5", Stainless Steel | Piece | √ | √ | √ | √
2.65 | Trolley, dressing: Stainless Steel: tubular pipe framework. Mounted on 100mm diameter castor wheels, two with brakes. Two SS shelves with railings. Fitted with SS basin and pail. Dimensions (LXWXH): 910X460X800mm | Set | √ | √ | √
2.66 | Urinal plastic: With marking from 50 - 1000ml | Piece | √ | √ | √ | √
2.67 | Vial Opener: Stainless Steel | Piece | √ | √ | √ | √
2.68 | Weighing scale stand model: For weight & height, 120 kg (Sliding Lever mechanism) | Piece | √ | √ | √ | √
2.69 | Wheel Chair with commode | Piece | √ | √
2.7 | Oxygen Concentrator : Small, quiet and light weight, with oxygen sensing device. Oxygen purity 85 - 99% at all flow rates, Flow rate : 0.5 - 5 LPM, Sound level: low, Electrical requirement: Power 240V,50Hz, Accessories: Ozygen mask with tubing for adult, nasal soft-tip cannula with tubing for pediatric and adult patients, oxygen bubble humidifier, warranty 3 - 5 years. Power supply: 220-240V. Power cord with round 3 pin top | set | √ | √
2.71 | Needle Cutter (Mini Destructor for needles and syringes): Balcan | piece | √ | √ | √ | √ | discuss with aue pem

## Sutures
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Sutures
1 | General Sutures
1.1 | Catgut Plain 2-0 NW 3/8 circle round bodied 30mm, 76 cm | 12Pcs/ Pkt | √ | √ | √ | √
1.2 | Chromic Catgut 1, 1/2 circle round bodied, 40mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.3 | Chromic Catgut 2-0, 1/2 circle round bodied 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.4 | Chromic Catgut 3-0, 1/2 circle reverse cutting 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.5 | Chromic Catgut 4-0, 3/8 circle round bodied 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
1.6 | Chromic Catgut 5-0, 3/8 circle round body 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
1.7 | Nylon 1-0, reverse cutting 45 mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.8 | Nylon 2-0, 3/8 circle reverse cutting 45mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.9 | Nylon 3-0, 3/8 circle reverse cutting 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.1 | Nylon 4-0, 3/8 circle cutting 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
1.11 | Nylon 6-0 reverse cutting 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
1.12 | Prolene 1,1/2 circle reverse cutting Heavy 100cm | 12Pcs/ Pkt | √ | √ | √
1.13 | Prolene 2-0, 1/2 circle round body 30mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.14 | Prolene 3-0, 1/2 circle round body 25mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.15 | Prolene 3-0, reverse cutting 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.16 | Prolene 4-0, 1/2 circle round body 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.17 | Prolene 6-0 round bodied 8mm, 60cm double needled | 12Pcs/ Pkt | √ | √
1.18 | Prolene 7-0, 8mm round bodied 60cm double needle | 12Pcs/ Pkt | √ | √
1.19 | Silk 2-0, 3/8 circle reverse cutting 45mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.2 | Silk 2-0, 1/2 circle round bodied 30mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.21 | Silk 3-0, 3/8 circle reverse cutting 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √ | √
1.22 | Silk 3-0, 1/2 circle round body 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.23 | Silk 4-0, 3/8 circle round bodied 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.24 | Silk sutupak 2-0 (120 cm) | 12Pcs/ Pkt | √ | √
1.25 | VICRYL 1, 1/2 circle round bodied 40mm needle heavy, 120 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.26 | Vicry 1-0, 1/2 circle round bodied 40mm needle, 90 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.27 | Vicryl 2-0, 1/2 circle round bodied 30mm, 90 cm suture length | 12Pcs/ Pkt | √ | √
1.28 | Vicryl 2-0 reverse cutting, 45mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.29 | Vicryl 3-0 reverse cutting, 26mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.3 | Vicryl 3-0, 1/2 circle cutting, 22mm, 76 cm suture length | 12Pcs/ Pkt | √ | √ | √
1.31 | Vicryl 4-0, 1/2 circle round bodied needle 20mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
1.32 | Vicryl 5-0, 1/2 circle round bodied 16mm, 76 cm suture length | 12Pcs/ Pkt | √ | √
1.33 | Vicryl 3-0, 1/2 circle round bodied, 20mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √
2 | Eye Sutures
2.1 | Nylon 10-0 | 12Pcs/ Pkt | √ | √
2.2 | Silk 5-0, 3/8 circle round bodied 16mm | 12Pcs/ Pkt | √ | √
2.3 | Silk 6-0, 1/4 circle black braided 8mm | 12Pcs/ Pkt | √ | √
2.4 | Silk 6-0, 3/8 micropoint spatulated 6mm double | 12Pcs/ Pkt | √ | √
2.5 | Vicryl 5-0 on ap2 needle | 12Pcs/ Pkt | √
2.6 | Vicryl 6-0, micropoint 1/4 circle spatulated | 12Pcs/ Pkt | √ | √
2.7 | Vicryl 8-0, micropoint 3/8 circle spatulated | 12Pcs/ Pkt | √ | √
2.8 | Silk 8-0, micropoint 3/8 circle spatulated | 12Pcs/ Pkt | √ | √
3 | Orthopedic Surgery |  | √ | √ | √ | √
3.1 | Ethibond 1-0 reverse cutting 45mm | 12Pcs/ Pkt | √ | √ | √
3.2 | Ethibond 2 taper cut 45mm | 12Pcs/ Pkt | √ | √ | √
3.3 | Ethibond 2-0 tapercut 45mm | 12Pcs/ Pkt | √ | √ | √
3.4 | Ethibond 5 tapercut 50mm | 12Pcs/ Pkt | √ | √ | √
3.5 | Braided Polyblend suture (fibre wire) size 5, 1/2 circle, 45mm, 96.5 cm | 12Pcs/ Pkt | √ | √
3.6 | Braided Polyblend suture (fibre wire) size 2, 1/2 circle, 45mm, 96.5 cm | 12Pcs/ Pkt | √ | √
3.7 | Braided Polyblend suture (fibre wire) size 2.0, 1/2 circle, 30mm, 96.5 cm | 12Pcs/ Pkt | √ | √
3.8 | Braided Polyblend suture (fibre wire) size 2.0, 1/2 circle, 40mm 75-90 cm | 12Pcs/ Pkt | √ | √
3.9 | Braided Polyblend suture (fibre wire) size 4.0, 1/2 circle, 20mm, 70cm | 12Pcs/ Pkt | √ | √
3.1 | Nylon 4-0 1/2 circle reverse cutting, 20 mm needle, 76 cm suture length | 12Pcs/ Pkt | √ | √
3.11 | PDS (PolyDioxanone), size: 1, double strand, 45 mm with length of 90 cm | 12Pcs/ Pkt | √ | √
4 | Surgical Department |  | √ | √ | √ | √
4.1 | PDS (PolyDioxanone), size: 1, double strand, 45 mm with length of 90 cm | 12Pcs/Pkt | √ | √
4.2 | PDS (PolyDioxanone), size: 1.0, 1/2 circle round body 25mm, 76 cm suture length | 12Pcs/Pkt | √ | √
4.3 | PDS (PolyDioxanone), size: 2/0, 1/2 circle round bodied 30mm, 90 cm suture length | 12Pcs/Pkt | √ | √
4.4 | PDS (PolyDioxanone), size: 3/0, 1/2 circle round bodied, 20mm needle, 76 cm suture length | 12Pcs/Pkt | √ | √
4.5 | PDS (PolyDioxanone), size: 4/0,1/2 circle round bodied needle 20mm, 76 cm suture length | 12Pcs/Pkt | √ | √
4.6 | PDS (PolyDioxanone), size: 5/0 , 1/2 circle round bodied 16mm, 76 cm suture length | 12Pcs/Pkt | √ | √
4.7 | Monocryl 2-0, 3/8 circle round bodied 30mm, 76 cm | 12Pcs/Pkt | √ | √
5 | Pediatric Department |  | √ | √ | √ | √
5.1 | PDS (PolyDioxanone), size: 4/0,3/8 circle cutting needle 19mm, 76 cm suture length | 12Pcs/Pkt | √
5.2 | Vicryl 4-0, 3/8 circle cutting needle 19mm, 76 cm suture length | 12Pcs/Pkt | √
5.3 | Vicryl 5-0, 3/8 circle cutting 13mm, 76 cm suture length | 12Pcs/Pkt | √
6 | Dental Department |  | √ | √ | √ | √
6.1 | Plain catgut 6.0 (Fast absorbing), PC-1 13mm Needle, 3/8 circle, 45 cm suture length | 12Pcs/Pkt | √
6.2 | Plain catgut 5.0 (Fast absorbing), PC-1 13mm Needle, 3/8 circle, 45 cm suture length | 12Pcs/Pkt | √
6.3 | PDS (PolyDioxanone), size: 4/0,3/8 circle reverse cutting needle 19mm, 76 cm suture length | 12Pcs/Pkt | √
6.4 | Monocryl 6.0, 3/8 circle, 11mm needle, Reverse cutting 45 cm suture length | 12Pcs/Pkt | √

## RCDC
Non Drugs List
Sl.no | Generic Name and Specification | Pack size | NRH | RRH | DH | BHU II | Remarks
Royal Centre for Disease Control
1 | National Tuberculosis Reference Laboratory
1.1 | Biological Safety Cabinet: Class II, Type B1, NU_427-400E NuAIRE Bio-Safety Cabinet- Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
1.2 | Biological Safety Cabinet: Class II, Type B2, NU-427-400E NuAIRE | Piece
1.3 | Autoclave: ES-315 TOMY ES-315 TOMY High pressure steam sterilizer ES-315 use for sterilization and decontamination of laboratory equipments and waste. The vertical autoclave with a top-open lid, which can be operated easily with a single hand or foot. Operating temperature range (during sterilizing): 105 - 132°C, Operating pressure range: 0 - 186kPa, Maximum operating pressure 216kPa, Temperature control: Digital, microprocessor controlled, Temperature display / display range: Digital / -15 - 180°C, Pressure display / display range: Analog / 0 - 400kPa, Heat source: 2.0kW electric heater, Timer control: Digital, microprocessor controlled, Timer display: Digital. Safety devices: • Inside the chamber overheat prevention, • Outer wall of the chamber overheat prevention, • Overpressure prevention, • Temperature sensor disconnection prevention, • Empty heating prevention, • Leakage breaker, • Safety valve. Malfunction prevention devices: • Low water level detection, • Exhaust valve knob open/close detection, • Insufficient sterilization detection, • Chamber lid open/close detection. Chamber material: SUS304. Capacity / material of exhasut bottle: 3ℓ / polyethylene. Type / material of lid gasket: Accommodating internal pressure / silicone rubber. Dimensions of main unit (mm): 490W x 560D x 1090H (Height from floor to control panel:875). Net weight: 80kg. Rated Voltage AC 120 / 220 / 230 / 240V. Power input: 20A (120V), 10A (220V), 9A (230 / 240V), Power supply requirements: Single-phase 220 / 230 / 240V AC (50 / 60Hz) 15A or above. Power Comsumption (calorific value): 2kW (1720kcal/h). Accessories: • Stainless steel basket with slatted bottom plate 1, • Water level sensor 1, • Inspection sheet 1, • Warranty card 1, • Customer card 1, • Operator's manual 1,• Clear plastic accessory case 1, • Attachment screw for clear plastic accessory case 1. Stainless basket (φ300 x 182mm) | Piece
1.4 | Autoclave: Ospery/ J8153 | Piece
1.5 | Refrigerated Centrifuge: Zentrifuge-Hittich, Rotana 460 R | Piece
1.6 | Refrigerated Centrifuge: Zentrifuge-Hittich, Rotana 460 R | Piece
1.7 | Micro-Centrifuge: MIKRO-200 Hettich, Germany | Piece
1.8 | TwinCubator: Hain Lifescience, Germany | Piece
1.9 | Thermal Cycler: GTQ Cycler 96,Hain Lifescience, Germany | Piece
1.1 | Water Bath: GFL, Type-1002, No. 11264413J | Piece
1.11 | Water & Oil Bath, MEMMERT | Piece
1.12 | Weighing Balance, CA-224 contech Instruments ltd CA-224 | Piece
1.13 | Weighing Balance, KERN PLS-310-3F, Germany | Piece
1.14 | Weighing Balance, UMT2 Mettler UMT2 | Piece
1.15 | Ph Meter, lab 850 INTERTEK lab 850 | Piece
1.16 | Light Microscope, CH20I OLYMPUS CH20I | Piece
1.17 | Liquid culture system, MGIT 960 Becton Dickinson MGIT 960 | Piece
1.18 | Isspissator (Open System), VFIN GRANT INSTRUMENT LTED VFIN | Piece
1.19 | Isspissator (Open System): TB-GD100IN GRANT INSTRUMENT LTED TB-GD100IN | Piece
1.2 | Isspissator (Closed System): J-IWL JESICO J-IWL | Piece
1.21 | Refrigerator: 2-8°C, Samsung | Piece
1.22 | Refrigerator, 2-8°C, Vision Scientific, VS 1302L | Piece
1.23 | Ultra Low Temperature Refrigerator: MDF-433V-PE PANASONIC MDF-433V-PE | Piece
1.24 | Freezer: Electrolux, TCW 1151 | Piece
1.25 | Universal Centrifuge: Z323K Hermele Labortechnik Z323K | Piece
1.26 | Shaker: KS501D10.010381 IKA-WERKE KS501D10.010381 | Piece
1.27 | Hot plate: HT1 BIBBY STERILIN LTD HT1 | Piece
1.28 | Incubator: MEMMERT Model 600 D6062 | Piece
1.29 | Incubator: MEMMERT INE600 | Piece
1.3 | Incubator: MEMMERT INE-800 | Piece
1.31 | Incubator: MEMMERT INE-800 | Piece
1.32 | Incubator: MEMMERT | Piece
1.33 | Hot Air Oven: UFB400 MEMMERT UFB400 | Piece
1.34 | Slide drying bench/Slide Warmer: with black anodised working plate, temp. Up to 100°C, regulation in 1°C steps, digital display, overheating protection, size WxDxH 160x290x30mm: with black anodised working plate, temp. Up to 100°C, regulation in 1°C steps, digital display, overheating protection, size WxDxH 160x290x30mm | Piece
1.35 | Safety Gas Burner & Safety gas tubing: 143000 INTEGRA BIOSCIENCES AG 143000: Safety gas burner, with foot switch, IR sensor, with rest heat display and splash guard. can be operated with touch free IR-Sensor, button function or foot pedal. With the DualKnob for adjusting gas and air flow, the flame can be precisely regulated. can be operated with stationary natural gas and propane / butane gas supplies, cartridge gas or gas cylinders. Safety Gas Tubing: safety gas tube 750mm long, with clamps. Inner diameter 9mm | Piece
1.36 | Glass brushing machine: Electrically operated, with a quick release chuck taking a wide variety of interchangeable nylon brushes for cleaning syringes, test tube, McCartney bottles, and round bottles up to 250ml. Mounted on a cast aluminium base with a perspex anti-splash shield, a white enamelled tray and incorporating a 1/16hp motor for 250V/50Hz single phase supplies. Dimensions: (approx.) 62.5x27cmsx25cms.; Approx.wt.:50-60 kgs., Accessories: Brushes, nylon, for round bottles capacity 30,120,250ml; brushes, nylon, for McCartney bottles- diameter 25mm. | Piece
1.37 | Vortex mixture: Vortexer is designed with a solid-state electronic control that provides adjustable speed from 40 to 3200rpm. Units can be operated in touch or contimnuous mode using the three-way switch. Supplied with 3-inch Platform and Pop-Off Cup. The 3-inch Platform accepts single or multiple test tubes, beakers and various flasks. The removable cover, made of thermoplastic rubber, is highly impervious to chemicals. The Pop-off Cup can accommodate tubes of virtually any size or shape from micro tubes to those up to 40mm in diameter and 200mm in length. Power required: 230 V, Hz: 50, AMPS: 0.5. Weight: 4 Kg (8.8 lb) Base Dimensions: (DxWxH) 165x122x165mm (6.5x4.8x6.5in). | Piece
1.38 | High Precision Weighing Balance for DST: Electronic analytical balance: ISO/GLP compliant print out, in conjunction with printer or PC, Data interface port RS232C, Display Backlit LCD, Calibration/ Adjustment: internal, isoCAL, Pan size Ø 80 mm, Repeatability ≤± 0.02 / 0.05 / 0.1 mg, Linearity ≤± 0.03 / 0.1 / 0.2 mg | Piece
1.39 | LED Microscope: Zeiss’ Primo Star iLED microscope, with following product specifications: high-grade optics; reflected rather than transmitted blue light for fluorescent applications; battery power on AC failure; easy switching between fluorescent and brightfield viewing; utility outside a darkroom; >10,000 hour bulb life; LED and filter sets that would allow use of Auramine O and acridine orange for TB and parasitologic examinations. Binocular microscope for use with electric light via power line or alternatively via battery pack. Battery pack for uninterruptible power supply, and is usually included as an accessory  Observation tube: binocular, 30 deg inclination (viewing angle) and 360 deg rotation  Stage: rectangular, built in mechanical stage with vernier scale (minimum: 14 mm x 135mm). No polymer belts, metal cables, timing belt systems or non-metallic components are acceptable in the drive mechanism. Coaxial controls must be low mounted for ease of use. Stage finger assembly is to be slide friendly so that it does not damage or break slides | Piece
1.4 | Slide Cabinet: Slide Cabinet: To store atleast 20,000 slides of 75x25mm slide - Horizontal with anodized aluminum slide tray of 350x210x35mm, uniformly slotted rows are numbered 1 to 50 to identify the slides.Power supply: AC 220V- 230V, 50HZ Power cord with round 3 pin top and suitable voltage stabilizer | Piece
1.41 | Trolley, stainless steel: deep pressed, rounded cornered, stainless steel shelves. Sound deading pads below each tray. Strong, welded rigid construction. Technical specification: No. of shelves/trays/tiers or bins: 2; Dimensions [w x d x h], mm: 1050x558x948. | Piece
1.42 | Deep freezer: Deep freezer [[-20ºC]: Upright Model: CFC free high efficiency double refrigeration system for cooling and freezing filled in the bottom. Temperature: -36oC to 0oC + - 0.5oC Temperature Control: (i) Digital temperature controller (ii) Microprocessor Control/Microcontroller-for temperature setting (iii) Alarms for-Voltage, Over heat, Over cool Size: 500 litres with minimum 3 Stainless Steel, rust free shelves. Fixed in casters for easy maneuverability. Electricity: 230 volts AC, 50Hz single phase. Refrigeration system: Heavy-duty maintenance free refrigeration system with hermetically sealed refrigeration compressors and reliable cascaded refrigeration to minimize noise and vibration. Air-cooled with security lock to prevent unintentional switch off. Short cooling time of 4 to 5 hours at 400C ambient temperature. The equipment should be of continuous duty and frostfree. Access port for CO2 back up. Audiovisual Electronic Alarm System independent of power supply Electrically heated doors for quick opening of frozen doors. | Piece
1.43 | PCR hood: PCR hood: Features: UVTect™ microprocessor controller; Class 100 (ISO 5) vertical laminar flow air; Polycarbonate and polypropylene design to reflect UV energy; Full access folding sash; Unibody plastic design - no joints or gaps in construction; Digital UV light timer 0-59 minutes; Lab event timer; 24″, 32″, and 48″ widths; nominal depth 24″; 220V AC versions. UVTect™ Features: Variable speed blower control; 254 nm ultraviolet light with user set digital timer; Constant HEPA filter monitoring; Constantly monitors and alerts effectiveness of the shortwave UV bulb life; Lab event timer. External Dimension (in inches): width x height x depth (32"x39"x24"); Internal Dimension (in inches): width x height x depth (31"x27"x23") | Piece
2 | Infectious Disease Serology Laboratory
2.1 | CENTRIFUGE: 5000 RPM (Gx2441-G000) | Piece
2.2 | INCUBATOR: SANYO (10202410) | Piece
2.3 | WATER BATH: Memmert (DIN 40050-IP-20) | Piece
2.4 | VORTEX: Genie-2/Abdos (2.41159/2.126764, VB3B008722) | Piece
2.5 | Rotatest Shaker: 4159B | Piece
2.6 | REFRIGERATOR (+2 -8): Vision Scientific, VS-1302L | Piece
2.7 | ELISA Plate Washer (12 manifold): Biotek, 207837/207836 | Piece
2.8 | ELISA Plate Washer (8 manifold): Biotek, 191866/183278 | Piece
2.9 | ELISA Plate reader: Sunrise 1002000585 | Piece
2.1 | ELISA Plate reader: ELx808, 207370 | Piece
2.11 | Freezer (-20): Whirlpool, CF41T300348002360 | Piece
2.12 | Balance: Mettler Toledo, B540527801 | Piece
3 | ENTERIC DISEASE LABORATORY
3.1 | Autoclave: HV-50 HICLAVE, HIRAYAMA | Piece
3.2 | Autoclave: ES-215, TOMY | Piece
3.3 | Automatic pipetter (E fill) | Piece
3.4 | Bunsen burner (Footswitch): Fireboy plus, INTEGRA FIREBOY, Switzerland | Piece
3.5 | BACTEC: BACTEC 9050, BD | Piece
3.6 | Compound Light microscope: CH20i BIMF, Olympus | Piece
3.7 | Distillation Plant: GFL, Germany | Piece
3.8 | Dark field microscope: Lab.A1, Zeiss | Piece
3.9 | Densimat Turbidometer: 99234, Densimat, Biomerieux | Piece
3.1 | Electronic weighing balance: PL202-S, Mettler Toledo, Switzerland | Piece
3.11 | Fluroscence Microscope: BX51TRF, Olympus | Piece
3.12 | Gel Electrphoresis chamber: PowerPAC basic, BIORAD | Piece
3.13 | Gel Doc imaging system: XR+ GelDOC, BIORAD | Piece
3.14 | Incubator: B-30, Memmert, Germany | Piece
3.15 | Incubator: Memmert, Germany | Piece
3.16 | Incubator: INE-800, Memmert, Germany | Piece
3.17 | Incubator: 100-180, Memmert, Germany | Piece
3.18 | Incubator: Memmert, Germany | Piece
3.19 | Loop sterilizer and holder: OTO6, Dinkelberg Analytics | Piece
3.2 | Microcentrifuge: MIKRO-200, Hettich | Piece
3.21 | Micro-oven: SANYO | Piece
3.22 | Microcentrifuge(spinner): PRISM MINI, C1801, PRISM | Piece
3.23 | Micropipette 2-10µl | Piece
3.24 | Micropipette 5-40µl | Piece
3.25 | Micropipette 20-200µl | Piece
3.26 | Micropipette 20-1000µl | Piece
3.27 | PH meter: Lab 850, Schott Instrument | Piece
3.28 | PCR Thermal cycler: PIKO 24, Thermoscientific | Piece
3.29 | PCR Thermal cycler: T100 Thermal cycler, BIORAD | Piece
3.3 | PCR Cabinet: UVP PCR cabinet | Piece
3.31 | Refrigerator: YPM-110P, SANDEN, INTERCOOL | Piece
3.32 | Refrigerator: 54-7-12105, SNOWLAND | Piece
3.33 | Upright freezer: 900 series -86C, Thermoscientific | Piece
3.34 | Vortex mixer: MX-S, swirlex, Abdos | Piece
3.35 | Vortex mixer: 21915, Whirlimixer, NICKEL ELECTRO LTD | Piece
3.36 | Vortex mixer(3000rpm): VB821200180150 | Piece
3.37 | Waterbath: W 14, Grant Instrument LTD, UK | Piece
3.38 | Bio-Safety Cabinet: Bio-Safety Cabinet- Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
3.39 | Autoclave: Autoclave: Chamber Capacity: 53 Litters, Temperature setting range: 105°C to 132 °C,Temperature control: Digital, Microprocessor controlled, Pressure display range: 0 to 400 kPa, Maximum operating pressure: 216kPa, Operating pressure range: 0-186kPa, Sterilization heat source: 2.0kW electric heater, Dimensions of main unit (mm): 490 W x 560D x 1090H (Height from floor to control panel:875) , Chamber diameter: 325 x 740mm, Net Weight: 80 kg, Power input: 20A (120V), 10A (220V), 9A (230/240V) | Piece
3.4 | Distillation Plant: WATER STILL 4 LIT/HOUR WITHOUT STORAGE TANK 230V /1 PH 50/60HZ AND HOSE SET 2940 GFL, Germany | Piece
3.41 | Refrigerator/Freezer: Refrigerator/Freezer: side to side refrigerator/freezer, Poduct details: GWB207FVQA, Capacity: - Refrigerator 349L / Freezer 178L, Energy Consumption: - 540kWh / Energy Saving Rating, Colour: Silver/Stainless steel, Dimensions: 894(w) x 1753(h) x 725(d) mm, Freezer Capacity 178L, Fridge Capacity 349L, Frost Level No Frost, Interior Light (Freezer)- Yes, Interior Light (Fridge) -Yes, Weight 100kg, External Features: - Silver/stainless steel / LED Digital Display Interior Features: - 5x Digital Sensors / Express Freezing / No Frost Cooling / Child Lock / Door Open Alarm / 12kg Freeze Capacity / N Climate Class Refrigerator Compartment: - Multi-Flow / 3x Tempered Glass Shelves / Door Baskets / Veg Box - Moist Balance Crisper / Egg Bank Freezer Compartment: - Multi Air Flow / Ice Maker / Door Baskets / 3x Tempered Glass Shelves / 1x Drawer | Piece
3.42 | Micro oven: 30 litre, 900 watt silver free standing microwave oven with an electronic timer, safety door locks, turnable table, hub and bearing washer, control panel, auto defrost, child lock | Piece
3.43 | Aspirator bottle: Polythene heavy duty aspirator bottle for storing distilled water with stop cock, carrying handle, 50 litres capacity | Piece
3.44 | Electronic weighing balance: Readability: 0.01g, Maximum load:210g, Adjustmment: Cal Ext, Repeatability: 0.008g, Linearity: 0.02g, Balance size: 195x225x67mm | Piece
3.45 | Electronic weighing balance: Readability: 0.001g, Maximum load: 81g, Adjustmment: Cal Ext, Repeatability: 0.0008g, Linearity: 0.002g, Balance size: 195x225x137mm | Piece
3.46 | Safety Gas Burner & Safety gas tubing: Safety gas burner, with foot switch, IR sensor, with rest heat display and splash guard. can be operated with touch free IR-Sensor, button function or foot pedal. With the DualKnob for adjusting gas and air flow, the flame can be precisely regulated. can be operated with stationary natural gas and propane / butane gas supplies, cartridge gas or gas cylinders. Safety Gas Tubing: safety gas tube 750mm long, with clamps. Inner diameter 9mm | Piece
4 | INFLUENZA/MOLECULAR LABORATORY
4.1 | Class II, Type B1-Biological Safety Cabinet: Model No-NU-427400E ( Brand-LABGARD) | Piece
4.2 | Class II, Microbiological Safety Cabinet: ENVAIR | Piece
4.3 | Water Bath | Piece
4.4 | Ependrof Centrifuge 5810 R: Sl No-5811BM567216 | Piece
4.5 | Inverted Microscope: Nikon | Piece
4.6 | Micro-Centrifuge: Hero Lab | Piece
4.7 | Centrifuge 5415 D | Piece
4.8 | Vortex Mixture | Piece
4.9 | 2-8°C Refrigerator | Piece
4.1 | ABI Machine ( 7500 Fast Dx Real Time PCR Instrument: Applied Biosystems | Piece
4.11 | Thermal Cycler-Rotor Gene: QIAGEN | Piece
4.12 | Thermal Cycler-Bio Rad: Bio-RAD | Piece
4.13 | Micro-Oven | Piece
4.14 | Electrophoresis Chamber | Piece
4.15 | Gel Doc ( Molecular Imager Gel Doc): Bio- RAD | Piece
4.16 | Freezer -30°C: Sanyo | Piece
4.17 | PCR Hood: ESCO | Piece
4.18 | Combi-Spin | Piece
4.19 | Fluroscent Microscope | Piece
4.2 | Egg Incubator | Piece
4.21 | Deep Freezer -80°C: SANYO | Piece
4.22 | Freezer -20-40°C: INSEGNIA | Piece
4.23 | Deep Freezer -80°C: INSEGNIA | Piece
4.24 | Dry Ice Machine | Piece
4.25 | Humidified Incubator | Piece
4.26 | Benchtop centrifuge: Benchtop centrifuge for microtitre plates: Spin speed up to 2000 rpm for microtitre plates ; Timed operation (1 to 90 minutes) with automatic switch-off; digital display; power supply 230V; weight 13-14 Kg; accompanied with rotors (changeable) for accomodating at least 2 standard 96 well-microtitre plates; | Piece
4.27 | Personal microcentrifuge: Personal microcentrifuge for 0.2 ml tubes: Holds two 8-strip or 16 individual 0.2 ml tubes, 6000 rpm, 2000 x g; Convenient on/off switch | Piece
4.28 | CO2 Cylinder: CO2 Cylinder for Dry Ice maker with siphon ( FILLED WITH CO2 GAS): Cylinder type 50 lb, Grade Industrial, Purity 99.5%,Volume (cu.Ft) 436, Volume (Kg) 22.7, Volume (lb) 50, Pressure (PSIG) 835 | Piece
5 | Water Quality Testing Laboratory
5.1 | Ammonia Gas sensing electrode: 95-12 Orion 95-12 | Piece
5.2 | Arsenic Kit: Lovibond | Piece
5.3 | Conductivity/Tds Meter: 44600 Hach 44600 | Piece
5.4 | Del Agua water testing kit: Del Agua Del-501-00 | Piece
5.5 | Digital Burette: Bibby | Piece
5.6 | Digital Burette: Titrettee | Piece
5.7 | Distillation Plant: 2001/4 GFL 2001/4 | Piece
5.8 | DO Meter: ELE International | Piece
5.9 | Electronic Balance: Metler Toledo BD202: Metler Toledo BD202 | Piece
5.1 | Filter Manifold: Millipore | Piece
5.11 | Filter Manifold: Gelman Sceiences: Gelman Sciences | Piece
5.12 | Hot Air Oven: SANYO Gall OMT150xx2 | Piece
5.13 | Hotop: Tarsons MC01-5020 | Piece
5.14 | Incubator: WTW | Piece
5.15 | Incubator: SANYO MIRI 162 | Piece
5.16 | Oven(Thermolyne): Sybron F6030 CM | Piece
5.17 | Oximeter: WTW oxi330 | Piece
5.18 | Oximeter(BOD Meter): WTW OxiTop | Piece
5.19 | Ph Meter: Jenway 3070 | Piece
5.2 | Ph Meter: Schott Instruments Lab 850 2001759 | Piece
5.21 | pH/ORP meter: Hanna HI 98172 | Piece
5.22 | Photometer: Palintest | Piece
5.23 | Pipette Washer: ABC LABO | Piece
5.24 | Refrigerator: Samsung RT26M | Piece
5.25 | Sample Digester: Hach 44336-62 | Piece
5.26 | Spectrophotometer(UV-160A): Shimadzu PH26402A2 | Piece
5.27 | Spectrophotometer: Hach DR2000(44800-60) | Piece
5.28 | Turbidity Meter: Hanna HI93703 | Piece
5.29 | Turbidity Meter: Jenway 6035 | Piece
5.3 | Turbidity Meter: Lutran TU-2016 | Piece
5.31 | Vaccum Pump: KnF VDE0530 | Piece
5.32 | GARMIN(GPS): OREGON-650 | Piece
5.33 | Arsenic meter: Lovibond | Piece
5.34 | Bio-Safety Cabinet: Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
5.35 | Laboratory filtration pump: The lab filtration pump that transfers filtered liquids directly to waste; Ideal for use both on the work bench and in the laminar flow hood where its small footprint; No need to use additional containers to capture the waste liquid; Quiet design:Optimized to reduce the noise down; Maintenance-free: pump technology and the materials of construction provide reliable vacuum performance and make the pump maintenance-free. The diaphragm is made of highly durable chemically resistant polytetrafluoroethylene | Piece
5.36 | Manifold filter: Manifold filter 6-place for rubber stopper tulip; Overview: Vacuum manifolds support simultaneous filtration of six test samples. Use Millipore glass, plastic or stainless steel filter holders fitted with standard No. 8 perforated stoppers over holder outlets. Manifold is available with the new Microfil head for use with presterilized, ready-to-use funnels and membranes; Application: General microbiological analysis, Water Monitoring; Materials:Stainless Steel. | Piece
5.37 | Acid Digestion Fume Hood with base cabinet: Acid Digestion Fume Hood with base cabinet: should be designed to perform digestion applications; should built with unplasticised polyvinylchloride (u-PVC) internal surfaces (including the work surface) and a polycarbonate sash window. UL classified; Full uPVC or PP interior; Excellent protection against acids; Excellent aesthetics; 5 degree sloped front; Superior containment at 0.3m/s; High energy savings; Superstructure: Tri-wall construction for maximum robustness; U-PVC internal chamber constructed of unplasticised polyvinylchloride; Chain and Sprocket Sash support system; Hood lighting should be pre-wired. Electronically ballasted lighting, energy efficient, instant start. Typical light intensity on work surface should be >1076 lux in zero ambient conditions; U-PVC exhaust collar ensures superior chemical resistance. Safety Features: Shaping vanes increase airflow "sweep" at the critical area at the sidewall to improve containment, especially when laboratory personnel walk fast in front of the hood; Sash made from polycarbonate to prevent hydrofluoric acid fume etching; Sash stop limits sash movement beyond 457mm (18.0"), 'encouraging' user to work at safe positions; When sash is raised above 457mm (18.0"), it should automatically and gently fall back to the safe level unless held in place. Allowing the sash to be raised temporarily to the full open position for set up of equipment and apparatus in hood, while enforcing regular operation of the hood with lowered sash; Aerodynamic foil entry provides maximum airflow to "sweep" the critical boundary layer near the work surface level to reduce turbulence and eliminate backflow; Serviceability: U-PVC internal work zone allowing for easy cleaning of all surfaces; Removable baffle system allowing for thorough cleaning inside the hood; Chain and sprocket system should require minimal service; Internal access panels on side walls should facilitate access to plumbing connections for service especially when hoods are installed next to each other; Removable front panel facilitates easy access to lighting and other electrical components mounted above the hood chamber. Local ducting and exust fan should be included. | Piece
5.38 | Glassware Dryers: External – (h) x (w) x (d) -1626 x 1168 x 610(mm); Internal – (h) x (w) x (d)-1430 x 1165 x 590; Capacity (Liters)-1000; Air circulation-by hot air fan; Voltage- 230v | Piece
5.39 | Autoclave: High pressure steam sterilizer ES-315 use for sterilization and decontamination of laboratory equipments and waste. The vertical autoclave with a top-open lid, which can be operated easily with a single hand or foot. Operating temperature range (during sterilizing): 105 - 132°C, Operating pressure range: 0 - 186kPa, Maximum operating pressure 216kPa, Temperature control: Digital, microprocessor controlled, Temperature display / display range: Digital / -15 - 180°C, Pressure display / display range: Analog / 0 - 400kPa, Heat source: 2.0kW electric heater, Timer control: Digital, microprocessor controlled, Timer display: Digital. Safety devices: • Inside the chamber overheat prevention, • Outer wall of the chamber overheat prevention, • Overpressure prevention, • Temperature sensor disconnection prevention, • Empty heating prevention, • Leakage breaker, • Safety valve. Malfunction prevention devices: • Low water level detection, • Exhaust valve knob open/close detection, • Insufficient sterilization detection, • Chamber lid open/close detection. Chamber material: SUS304. Capacity / material of exhasut bottle: 3ℓ / polyethylene. Type / material of lid gasket: Accommodating internal pressure / silicone rubber. Dimensions of main unit (mm): 490W x 560D x 1090H (Height from floor to control panel:875). Net weight: 80kg. Rated Voltage AC 120 / 220 / 230 / 240V. Power input: 20A (120V), 10A (220V), 9A (230 / 240V), Power supply requirements: Single-phase 220 / 230 / 240V AC (50 / 60Hz) 15A or above. Power Comsumption (calorific value): 2kW (1720kcal/h). Accessories: • Stainless steel basket with slatted bottom plate 1, • Water level sensor 1, • Inspection sheet 1, • Warranty card 1, • Customer card 1, • Operator's manual 1,• Clear plastic accessory case 1, • Attachment screw for clear plastic accessory case 1. Stainless basket (φ300 x 182mm) | Piece
5.4 | Vortex Mixer: 3-inch Platform, Pop-off Cup, Speed (RPM), 600-3200, (2700 for 50Hz models), Dimensions (base), (D x W x H), 165 x 122 x 165mm, (6.5 x 4.8 x 6.5in) | Piece
5.41 | Incubator: Temperature range: Ambient + 5°C to 80°C. Temperature control: Digital P.I.D. controller. Dual LED display. Capacity: 50 - 100 liters.Power supply: AC 220V- 230V, 50HZ ; shelf provided:-2-3 | Piece
5.42 | HPLC: For detection of organic contamination and Pesticides in water. Ultra high pressure pump optimized for fast liquid chromatography, capable of pressures up to 1000 bar (15000 psi) and fl ow rates up to 5 ml/min. The electronically controlled pump drive provides for low pulsation, a highly precise fl ow rate and a stable baseline. Flow rate- 0.01 – 5.0 ml/min or 0.001 ml increments; Flow rate accuray- ± 1 %;n Flow rate precision- < 0.1 %; Features: touchscreen, active pressure and pulsation compensation, RFID automatic pump head detection (GLP), interchangeable pump heads; | Piece
5.43 | ICP-OES: Detector Type-High performance solid-state CID86 chip; Wavelength Range- 166 to 847nm; Auxiliary Gas- Mass flow control, 0-2 L/min; Data Acquisition Modes-Precision mode, Speed mode; Depth (Metric)- 750mm; Features- Intelligent introduction and rinse, Speed data acquisition mode, Get ready: automatic performance checks; Full-frame Capability- Yes; Height (Metric)- 590mm; Nebulizer Gas- Mass flow control, 0-2 L/min; Peristaltic Pump- 4-channel, mini peristaltic pump, Speed: 0 to 125rpm; Plasma Gas- Mass flow control, 0-2 L/min; Plasma Viewing-DUO or Radial; RF Source-27.12MHz solid state, 750 to 1500w output power (DUO restricted to 1350w); Spectral Bandpass- 7pm at 200nm; Spectrometer-Simultaneous echelle type, 52.91 grooves/mm ruled grating, 383mm effective focal length, 9.5° UV fused silica cross dispersion prism; Standard Sampling Kit- Concentric glass nebulizer, Glass cyclonic spray chamber, Semi-demountable EMT torch, 1.5mm bore quartz center tube (Radial), 2mm bore quartz center tube (DUO); Width (Metric)-840mm; Function- Detection of trace elements in water | Piece
6 | National Drug Testing Laboratory
6.1 | UHPLC: knauer | Piece
6.2 | Dissolution Apparatus: 708 DS Agilent Technologies | Piece
6.3 | Disintegration Apparatus: Labindia | Piece
6.4 | Autotitrator: Metrohm | Piece
6.5 | Autotitrator: Fischer Scientific | Piece
6.6 | UV-Vis Spectrophotometer: Perkin Elmer | Piece
6.7 | Friability Apparatus: Pharma test | Piece
6.8 | Analytical Balance: Shimadzu • Perform basic weighing with below balance, counting, percent, animal weighing, and density determination functions • Anti-theft feature secures balance to work station • Functional in 8 languages • Document critical weights in GLP mode via RS-232 communication with PC or printer • Sliding doors and removable shield chamber parts make cleanup easy • Backlit high-contrast display delivers excellent viewability in any lighting condition • Calibration:external• Capacity (g): 220g • Dimensions (" W):11.9• Dimensions (" H):13• Dimensions (" D):9.1 • Display: back light LCD• Readability (g):0.0001• Linearity (g):0.0002 • Repeatability (g):0.0001 • Weighing units: g, mg, kg, ct, oz, troy, lb, pennyweight, grain, Newton, momme, mesghal, tael • Stabilization time:205 secs• Power (VAC):120/220 • RS-232 interface: yes with miniprinter• Output:RS232 miniprinter • Accesorries: miniprinter | Piece
6.9 | Microscope-olympus | Piece
6.1 | pH meter - schaut instruments | Piece
6.11 | Water Deionizer- millipore | Piece
6.12 | Water Deionizer/Distillation Unit: millipore | Piece
6.13 | Centrifuge: Fischer Scientific | Piece
6.14 | Automatic shaker- werke | Piece
6.15 | Automatic shaker- bibby scientific | Piece
6.16 | Sonicator - rivotek | Piece
6.17 | Hot Plate- Fischer Scientific | Piece
6.18 | Polarimeter: :sugar scale ( with automatic temperature compensation).Measurement readings= angle of rotation, international sugar scale ( without temp compensation) , international sugar scale ( without automatic temp compensation) , specific optical rotation, concentration and purity. • Measurement range = angle of rotation: -89.999 to +89.999° , international sugar scale : -130 to +130 °z. • Resolution= angle of rotation: 0.001°(Default 0.01°z) , international sugar scale : 0.001°z ( default 0.01°z). • Measurement accuracy = angle of rotation displayed value +_ 0.01° ( -35 to +35°) , relative precision +_ 0.2% (-35.01 to -89.99° , +35.01 to +89.99°).. • International sugar scale displayed value +_ 0.03°z ( -101 to +101°z.. relative precision +_ 0.2% (-130 to - 101.01°z , +101.01 to +130 °z) • Length of observation tube= upto 200mm • Temperature = 10 to 30°c• Temp accuracy = +_ 0.5°c• Temp resolution = 0.5 °c • Display panel = backlit color LCD • Graphic display = indicating the conformity of measurement values of measurement with limit settings. • Output= 1) digital printer ( dot printer matrix printer) 2) RS-232 C input / output ports to a computer system. • Light source = halogen lamp• Accessories = observation tube 100mm n 200mm • Power supply = AC 100v to 240 V• Dimension = 48.5 × 28.5× 14.5 mm , 14.4 kg | Piece
6.19 | Particulate Matter Counter: Benifits: Defect Reduction,100% sample volume for greater accuracy,A precise sample volume within 0.1 ml is achieved with the syringe sampler, producing repeatable results,SamplerSight-Pharma software comes with a Validation Information Notebook, providing comprehensive software development documentation and full IQ, OQ, and PQ protocolsSamplerSight-Pharma complies with FDA21 CFR 11User-Friendly, Reduce operator error and ensure process ,accuracy with recipes for repeat sample processing to USP, EP, or JP standards • Automated particle counter calibration wizard for full calibration or routine verification • Alarm levels for pass/fail criteria ensure quality control Comprehensive System: Easy-to-use, pull-down menus,Security password maintains system integrity, Particle measurements are reported per ml or per container values,Small sample volume minimizes waste of expensive product,Various custom report formats are available.with a computer system provided. Features: Sizing capability from 1.5 to 125 µm, Fifteen user-selectable sizing channels,Meets and exceeds all USP, EP and IP requirements • Sample volumes from 0.2 ml to 1 liter • Adaptors for SVI and IV bag testing • Automated sensor calibration Applications: Pharmaceutical parenteral monitoring to USP, EP, IP and FDA standards,Parts/medical device cleanliness testing, Laboratory water sampling,WFI and Purified Water testing,Filter efficiency testing. | Piece
6.2 | FTIR: 1. Interferometer:25 mm, permanently aligned, Michelson, 45°, mechanical flexure 2. Spectrometer interface:USB-2 3. Size :16 x 31 x 13 cm (6 x 12 x 5 in) (W x D x H with ATR) 4. Weight :3.8 kg (8 lb) (with ATR) 5. Sample interface:DialPath, TumblIR, Diamond ATR, Diffuse Reflectance, Transmission 6. Software: PC software, Automated IQ/OQ, 21 CFR Part 11 compliant, Resolutions Pro for advanced data analysis. 7. Power:220 – 250 V AC, 60/50 Hz 8. Spectral range: KBr 6300–350 cm-1 and ZnSe 5100–600 cm -1 9. Spectral resolution: < 2 cm-1 10. Wavenumber accuracy:0.05 cm-1 Measured with ASTM 1921 11. Wavenumber reproducibility:0.005 cm-1 Measured with ASTM 1921 12. computer system and Printer attached. 13.Warranty: Five (5) year interferometer, Five (5) year laser, Three (3) year source Location dependent,Hardware support period. Seven (7) years from date of last unit manufacture. After this time,parts and supplies will be provided if available 14.Software support: Software upgrades to add additional functionality .. IQ, OQ and PQ to be conducted at the site of installation. | Piece
6.21 | Water Bath: Chamber Capacity: 34.5 liters (9.1 gallons) Temperature Display: Ambient +5 °C to 99.9 °C Dimensions (L x W x H) Work Area: 12 x 27 x 7.5 in (30.5 x 68.6 x 19.0 cm) Overall without cover: 14 x 36.5 x 9.5 in. (35.6 x 92.7 x 24.1 cm) Shelf: 11 11 x 26.5 in. (27.9 x 67.3 cm) Electrical Specs (50/60Hz) 120V : 1550 Watts 230V: 1550 Watts Temperature controller provides +/-0.05 °C uniformity at 37 °C and sensitivity of +/-0.05 °C with stainless steel gable cover • Water flow directed around bath perimeter • Elevated sample platform for more thorough circulation • Easy-to-maintain, corrosion resistant construction • Automatic sensor to protect bath from overheating • Single switch control for heater and circulator pump • Meets ASTM E715 Class IIA standards for uniformity | Piece
6.22 | Fumehood: Fume hood- should be designed to perform digestion applications; should built with unplasticised polyvinylchloride (u-PVC) internal surfaces (including the work surface) and a polycarbonate sash window. UL classified; Full uPVC or PP interior; Excellent protection against acids; Excellent aesthetics; 5 degree sloped front; Superior containment at 0.3m/s; High energy savings; Superstructure: Tri-wall construction for maximum robustness; U-PVC internal chamber constructed of unplasticised polyvinylchloride; Chain and Sprocket Sash support system; Hood lighting should be pre-wired. Electronically ballasted lighting, energy efficient, instant start. Typical light intensity on work surface should be >1076 lux in zero ambient conditions; U-PVC exhaust collar ensures superior chemical resistance. Safety Features: Shaping vanes increase airflow "sweep" at the critical area at the sidewall to improve containment, especially when laboratory personnel walk fast in front of the hood; Sash made from polycarbonate to prevent hydrofluoric acid fume etching; Sash stop limits sash movement beyond 457mm (18.0"), 'encouraging' user to work at safe positions; When sash is raised above 457mm (18.0"), it should automatically and gently fall back to the safe level unless held in place. Allowing the sash to be raised temporarily to the full open position for set up of equipment and apparatus in hood, while enforcing regular operation of the hood with lowered sash; Aerodynamic foil entry provides maximum airflow to "sweep" the critical boundary layer near the work surface level to reduce turbulence and eliminate backflow; Serviceability: U-PVC internal work zone allowing for easy cleaning of all surfaces; Removable baffle system allowing for thorough cleaning inside the hood; Chain and sprocket system should require minimal service; Internal access panels on side walls should facilitate access to plumbing connections for service especially when hoods are installed next to each other; Removable front panel facilitates easy access to lighting and other electrical components mounted above the hood chamber. Local ducting and exust fan should included. | Piece
6.23 | Hardness Tester:: Tablet hardness tester 1. Tablet size= 2mm to 30mm with accuracy of +0.5% 2. Hardness measurement by load cell up to= 500N 3. Units of measurement= N,kgf, and lbf 4. PC connectable 5. Programmable upper T1 and lower T2 tolerance limit 6. Automated Semi calibration 7. Can display of test value and calculate average 8. LCD display with back light 9. Acrylic glass cover with disposable container 10. Operates with rechargeable batteries and 6 v dc adaptor 11. Accessories =IQ,PQ ,OQ documents, printer,6V dc adaptor | Piece
6.24 | Trolley, stainless steel: deep pressed, rounded cornered, stainless steel shelves. Sound deading pads below each tray. Strong, welded rigid construction. Technical specification: No. of shelves/trays/tiers or bins: 2; Dimensions [w x d x h], mm: 1050x558x948 | Piece
6.25 | UV-Spectrophotomter: Spectrophotometer with ES Software and PC System and printer • Principle :Double-beam, double monochromator, ratio recording spectrophotometers. • Optical System:All-reflecting optical system (SiO2 coated) with holographic grating monochromators. Littrow mounting, sample thickness compensated detector optics. • Detector: Photomultiplier (UV/Vis) • Wavelength range: 190-1100 nm • Accuracy: ±0.1 nm • Absorbance Range: -3 to 4 A • Bandwidth: 0.5-4 nm, variable • Stray Light: <0.01% at 220 nm • Photometric Accuracy: ±0.001 A • Dimensions: 65W x 56D x 23.3H cm (25 x 22 x 9") Weight: 25 kg (57 lbs.) IQ, OQ and PQ to be conducted at the site of installation. | Piece
6.26 | Table Top Digital Timer: • 10-hour timing range tabletop clock with timer • Includes alarm, count-up, time-out function, and four-mode stopwatch • One-inch digital display for easy visibility • Runs on one C battery | Piece
6.27 | Digital LCD Alarm Thermometer Hygrometer Temperature Humidity Meter Features: Convenient wall mount and mini design Can check the max / min temperature and humidity Memorize or erase the last minimum and maximum temperature / humidity Ideally placed on the desk Indoor temperature, humidity, displayed at the same time, it has °F / °C to choice Hourly chime function, alarm clock function Ideal household decoration and ensure you the most desirable ambient temperature and humidity Multifunction instrument can display temperature, time and humidity all at the same time Powered by one AAA battery. (not included) Specification: Range (indoor): 0°C~50°C; 32°F~122°F (Temperature); 10%~99% (Humidity) Accuracy (temperature): ±0.1°C Accuracy (humidity): ±5% Resolution: 0.1°C (temperature); 1% (humidity) Indoor temperature, humidity, time is displayed at the same time °C / °F LCD screen dimension: 65 x 35 mm Dimension: Approx. 8 x 7.5 x 2cm (W x H x T) Max key: the highest measurement environment Min key: the lowest temperature / humidity | Piece
6.28 | Hot air oven: Interior: Stainless steel, material 1.4301 (ASTM 304),with all-round deep-drawn ribs to integrate the large-area heating with ceramic-metal sheath. Insertions: Stainless steel grid shelves Housing: Textured stainless steel, rear zinc-plated steel,with 4 feet, sizes 450, 750 and 1060 mounted on lockable castors, with mains cable and plug. Fresh air: Admixture of pre-heated fresh air by electronically adjustable air flap. Stainless steel interior Volume approx. l/cu. ft. :74/2.6 Width (A) mm/in: 400/15.8 Height (B) mm/in:560/22.0 Depth (less 39 mm/1.5 in for fan) (C) mm/in :330/13.0 Stainless steel grid shelves: 2 Max. number of grid shelves/shelves number: 6 Max. load per grid shelf/shelf : kg/lbs 20/44.1 Max. load of chamber kg/lbs :120 /2640 Textured stainless steel exterior Width (D) mm/in : 585/23.0 Height (size 450, 750, 1060 with castors) (E) mm/in: 944/37.2 Depth (without door handle),door handle + 56 mm (2.2 in) (F) mm/in : 514/20.2 Working-temperature range °C/°F: at least 5 (9.0) (UN/UNplus) at least 10 (18.0) (UF/UFplus) above ambient temperature to +300 (572) Setting temperature range °C/°F :+20 (68) to +300 (572) Setting accuracy °C/°F: up to 99.9 (211.8): 0.1 (0.2) / from 100 (212): 0.5 (0.9 | Piece
6.29 | HPLC: Physical specifications Features Description Pump system Hydraulic system :Dual pistons in series pump with proprietary servo-controlled variable stroke design and smooth motion control for active damping. Pump resolution step size :300 pl step size Settable flow range: 0.001 – 5 mL/min, in 0.001 mL/min increments (- executed in 300 pl/step increments) Flow precision :≤ 0.07 % RSD, or 0.01 min SD, whatever is greater (0.2 - 5.0 mL/min); based on retention time at constant room temperature Flow accuracy :}1% or }10 μL/min, whatever is greater; pumping degassed H2O at 10 MPa (100 bar) Maximum operating pressure:Operating range up to 120 MPa (1200 bar) up to 2 mL/min, ramping down to 80 MPa (800 bar) up to 5 mL/min Pressure pulsation :< 1 % amplitude or < 0.5 MPa (5 bar) whatever is greater at 1 mL/min water Compressibility compensation:Automatic, pre-defined, based on mobile phase selection Gradient formation: Low pressure quaternary mixing Delay volume: Standard configuration: < 350 μL With optional V380 Jet Weaver Mixer: < 500 μL Composition range: Settable range: 0.0 – 100.0 % Recommended range: 1.0 – 99.0 % or 5 μL/min Composition precision :< 0.15 % RSD, or 0.02 min SD whatever is greater (1 mL/min); based on retention time at constant room temperature Composition accuracy :}0.4 % absolute (1 - 99 % B, 0.5 - 2.0 mL/min with water/caffeine tracer, 400 bar) Integrated degassing unit:Number of channels: 4 Internal volume per channel: 1.5 mL Materials in contact with solvent:TFE/PDD copolymer, FEP, PEEK, PPS, stainless steel, polyimide Communications: Controller-area network (CAN), RS232C, APG remote: ready, start, stop and shutdown signals, LAN Safety and maintenance :Extensive diagnostics, error detection and display through company LabAdvisor, leak detection, safe leak handling, leak output signal for shutdown of the pumping system. Low voltage in major maintenance areas. GLP features :Early maintenance feedback (EMF) for continuous tracking of instrument usage in terms of seal wear and volume of pumped mobile phase with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors. | Piece
6.3 | HPLC continuation: Autosampler Physical specifications Weight: 15.5 kg (34.2 lbs) Dimensions :200 x 345 x 440 mm (height x width x depth) (8 x 13.5 x 17 inches) Line voltage: 100 – 240 VAC, ± 10 % Wide-ranging capability Line frequency :50 or 60 Hz, ± 5 % Power consumption :300 VA Maximum (apparent power) Power consumption (active power): 200 W Maximum Ambient operating temperature: 4–55 °C (39–131 °F) Ambient non-operating temperature: -40–70 °C (-40–158 °F) Humidity: < 95 %, at 25–40 °C (77–104 °F) Non-condensing Operating altitude: Up to 2000 m (6562 ft) Non-operating altitude :Up to 4600 m (15091 ft) Performance specifications Injection range: 0.1 – 20 μL With standard loop 0.1 – 40 μL With extended loop in 0.1 μL increments Precision Typically :< 0.25 % RSD from Measured with injections 5 – 20 μL, Typically < 0.5 % RSD of benzyl alcohol. from 2 – 5 μL volume, Typically < 0.7 % RSD from 1 – 2 μL volume Accuracy :± 1 % (10 μL, n-10) Pressure range: Up to 1200 bar Cooling/Thermostatting Compatible with G1330B Autosampler Thermostat, 4-40°C Sample viscosity range 0.2 – 5 cp Sample capacity: Capacity 2 x well plates (MTP) + 10 x 2 mL vials, 108 x 2 mL vials in 2 x 54 vial plate plus 10 additional 2 mL vials, 30 x 6 mL vials in 2 x 15 vial plate, plus 10 additional 2 mL vials, 54 Eppendorf tubes (0.5/1.5/2.0 mL) in 2 x 27 Eppendorf tube plate Injection cycle time :Typically < 19 s without needle wash and typically <25 s with 3 s needle wash; conditions: default draw (100 μL/min) and eject speed (400 μL/min). Injection volume :< 5 μL (water) | Piece
6.31 | HPLC contiuation: Thermostatted Column Compartment G1316C Physical specifications Weight :11.2 kg (24.7 lbs) Dimensions: 140 x 345 x 435 mm (height x width x depth) (5.5 x 13.5 x 17 inches) Line voltage :100 – 240 VAC, ± 10 % Wide-ranging capability Line frequency :50 or 60 Hz, ± 5 % Power consumption :320 VA / 150W / 512 BTU Maximum Ambient operating temperature: 0–55 °C (32–131 °F) Ambient non-operating temperature: -40–70 °C (-40–158 °F) Humidity :< 95 %, at 25–40 °C (77–104 °F) Non-condensing Operating Altitude: Up to 2000 m (6562 ft) Non-operating altitude :Up to 4600 m (15091 ft | Piece
6.32 | HPLC continuation: Performance specifications Temperature range :10 °C below ambient to 80 °C (5 mL/min) 10 °C below ambient to 100 °C (2.5 mL/min) Temperature stability: ± 0.05 °C Temperature accuracy: ± 0.8 °C ± 0.5 °C With calibration Column capacity • 2 columns of 300 mm length Up to three 1290 Infinity with individual solvent heating TCCs can be clustered • 3 columns of 300 mm length in a system. These will in series appear with a single joint • 4 columns of 100 mm length user interface in the with opt. Low dispersion control software.* heat exchanger • 6 columns (max. 300 mm) with triple-TCC clustering and column selection valves • 8 columns (max. 100 mm) with dual-TCC clustering and column selection valves Warm-up/cool-down time 5 min from ambient to 40 °C 10 min from 40 – 20 °C Temperature zones • 2 (single TCC) • 4 (dual-TCC cluster) • 6 (triple-TCC cluster) Individual standby temperatures definable for cluster configurations. Dead volume • 1.6 μL low dispersion heat The use of low dispersion exchanger heat exchangers is rec- • 3 μL left heat exchanger ommended. 3 μL and 6 μL • 6 μL right heat exchanger heat exchanges are available for backwards compatibility. Available valves • 8 position / 9 port, 8x column The optional valve drive selection valve, min. (G1316C #058) needs to be 2 1290 Infinity TCCs required, present. Valve kits have to (G4230A 400 bar, be ordered as separate (G4230B 1200 bar) products. • 2 position / 6 port (G4231B),1200 bar • 2 position /10 port (G4232B), 1200 bar | Piece
6.33 | HPLC continuation: Diode Array Detector G4212A Physical specifications Weight: 11.5 kg (26 lbs) Dimensions: 140 x 345 x 435 mm (height x width x depth) (5.5 x 13.5 x 17 inches) Line voltage :100 – 240 VAC, ± 10 % Wide-ranging capability Line frequency: 50 or 60 Hz, ± 5 % Power consumption :160 VA / 130 W / 444 BTU Maximum Ambient operating temperature: 5-40 °C (41–104 °F) Ambient non-operating temperature :-40–70 °C (-40–158 °F) Humidity Operating: < 80 % Non-condensing Operating altitude Up :to 2000 m (6562 ft) Non-operating altitude :Up to 4600 m (15091 ft) For storing the module | Piece
6.34 | HPLC continuation: Performance specifications Detection type :1024-element photodiode array Light source UV-lamp Equipped with RFID tagthat holds lamp typical information. Wavelength range :190 – 640 nm Short term noise (ASTM) Single < ± 3 x 10-6 AU at 230/4 nm, and Multi-Wavelength slit width 4 nm, TC 2 s ASTM Drift < 0.5 x 10-3 AU/hr at 230 nm Linear absorbance range > 2.0 AU (5 %) at 265 nm Typically 2.5 AU (5 %) Wavelength accuracy ± 1 nm After recalibration with deuterium lines Wavelength bunching 2 – 400 nm Programmable in steps of 1 nm Number of signals 8 Slit width 1, 2, 4 , 8 nm Programmable slit Diode width ~ 0.5 nm Signal data rate up to 160 Hz Spectra Data rate up to 160 Hz Flow cells Max-Light Cartridge Cell Cartridge type, equipped (10 mm, V(s) 1.0 μL), with RFID tags that holds 60 bar (870 psi) cell information. pressure maximum Max-Light Cartridge Cell (60 mm, V (s) 4.0 μL), 60 bar (870 psi) Max-Light Cartridge Test Cell Accesorries: Computer system with a printer. IQ, OQ and PQ to be conducted at the site of installation. | Piece
6.35 | Karl Fischer titrator: Measuring interface A measuring interface with one high-ohm measuring input (Ind.) for pH, redox or ISE electrodes, one measuring input (Ref.) for a separate reference electrode, one measuring input (Temp.) for temperature sensors (Pt1000 or NTC), one measuring input (Pol.) for a polarizable electrode and one connector (iConnect) for an iConnect Titrando: • Type plate:Contains specifications concerning mains voltage, instrument type and serial number. • USB connector (USB 1 and USB 2) USB ports (type A) for connecting printer,keyboard, barcode reader, additional Titrandos, USB Sample Processor, etc • Controller connector Connection of a Touch Control or PC with installed PC software. Mini DIN, 9-pin • MSB connector (MSB 1 bis MSB 4 Serial Bus. For connecting external dosing devices, stirrers or remote boxes. Mini DIN, 9-pin. • Mains connection socket • Measuring interface 1 (Input 1) • Electrode connector (Ind.) For connecting pH, redox or ISE electrodes with integrated or separated reference electrode. Socket F. • Electrode connector (Ref.) For connecting reference electrodes, e.g. Ag/AgCl reference electrode. Socket B, 4 mm. • Temperature sensor connector (Temp.) For connecting temperature sensors, e.g. Pt1000 or NTC sensors. Two B sockets, 2 mm. • Electrode connector (Pol.) For connecting polarizable electrodes, e.g. double Pt electrodes. Socket F. • iConnect connector (iConnect) For connecting electrodes with integrated data chip Titration vessel / 50-150 mL: Material: Clear glass ,Height (mm): 80, Outer diameter (mm): 78 ,Volume (mL): 20 ...90 Titration vessel / 20-90 mL: Material: Clear glass,Height (mm): 80,Outer diameter (mm): 78,Volume (mL): 50 ...150 KF titration vessel lid:For volumetric KF titrations. Material: PTFE,Note, Material: Insert, KF adsorber tube:For KF equipments. Including cover and O-ring.,Height (mm): 113,Outer diameter (mm): 8 Septum / 12 mm / 5 pieces:Set of 5 pieces. Material: Silicone rubber.Height (mm): 2,Outer diameter (mm): 12 Stirring bar / 16 mm: Stirring bar with magnetic core, PTFE covering, 16 mm long.,Material: PTFE,Outer diameter (mm): 4,Length (mm): 16 Stirring bar / 25 mm:Stirring bar with magnetic core, PTFE covering, length 25 mm,Material: PTFE,Outer diameter (mm): 5,Length (mm): 25 Holding clamp for bottles:Holding spring for reagent bottles and exchange units. Glass weighing spoon:Material: Glass,Length (mm): 120 | Piece
6.36 | Karl Fischer titrator continuation: Molecular sieve:Molecular sieve. Bottle of 250 g. Pore size: 0.3 nm. Without moisture indicator. For Rancimats and Karl Fischer instruments. Stirrer:Magnetic stirrer with base plate, support rod and electrode holder. With permanently attached cable for MSB (Metrohm Serial Bus). Ti Stand Compact titration stand with magnetic stirrer and built-in pump for adding solvent and aspirating the titration vessel content. Suitable for Karl Fischer titrations with Titrandos. MSB connector. Including extensive accessories, solvent and waste bottle. Dimensions in mm (B/H/T): 106/ 101/ 220.Remark Dimensions: Height with support rod 412 mm Touch Control Operating element for the Titrandos, USB Sample Processors and the 846 Dosing Interface. Touch-sensitive color display. Simple and intuitive user guidance. Two PCMCIA card slots. Default dialog language English, one additional language can be added later. Including CompactFlash card with adapter, demo version of the PC Control software. KF Thermoprep The KF Thermoprep is used for thermal sample preparation in Karl Fischer titration. Many substances are not suitable for a direct Karl Fischer titration because they are not soluble, react with the KF reagent or because they release their water content only very slowly or only at high temperatures. With the aid of tightly sealed sample vials the sample are brought to the oven and can afterwards be determined by every volumetric or coulometric KF titrator. Dosing unit 5 mL Dosing unit with integrated data chip with 5 mL glass cylinder and light protection, mountable to a reagent bottle with ISO/DIN GL45 glass thread. FEP tubing connection, antidiffusion buret tip.Volume (mL): 5 Dosing unit 10 mL Dosing unit with integrated data chip with 10 mL glass cylinder and light protection,mountable to a reagent bottle with ISO/DIN GL45 glass thread. FEP tubing connection, antidiffusion buret tip. Volume (mL): 10 Dosing unit 20 mL Dosing unit with integrated data chip with 20 mL glass cylinder and light protection, mountable to a reagent bottle with ISO/DIN GL45 glass thread. FEP tubing connection, antidiffusion buret tip. Volume (mL): 20 | Piece
6.37 | Karl Fischer titrator continuation: KFT:Karl Fischer titration for water determination. As measuring mode can be selected: – Ipol (voltametric measurement with adjustable polarization current) – Upol (amperometric measurement with adjustable polarization Voltage pH electrode Redox electrode ISE electrode Reference electrode Polarizable electrode Temperature sensor iConnect Printer IQ, OQ and PQ to be conducted at the site of installation. | Piece
6.38 | Bio-Safety Level 3
6.39 | Bio-SAfety Cabinet: Bio-Safety Cabinet-Class 2 Type A2, with base console stand/wheels, width 4 foot (1.2 m). The basic equipment shall consist of exhaust HEPA filter, ‘Supply HEPA filter, negative pressure plenum, positive pressure plenum front opening. Work access opening Standard Opening Height 10" (254 mm),/ 12" (305 mm) / 8" (203 mm) Standard Inflow Velocity 105 fpm (0.53 m/s). Viewing Window: Standard Safety Plate Sliding Glass Fully Closed to 21" (533 mm) . Power 230 V, 50/60 Hz; Face dimensions: 4ft. (L) X 2ft. (W) X 6ft . The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or equivalent ISO within the work area. Face Velocity: At least 100 fpm Light Intensity at work surface: 800 lux or more over the entire work surface. Noise level : <70 dBA. UV germicidal lamp intensity: >40 microwatt/sq.cm over the entire work surface.Construction : Main body, side and rear panel: Electrogalvanized Steel or Mild Steel, oven baked epoxy powder coated finish.Switches and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp. External height 2200 mm including support stand, allowing an available space of at least 400 mm from the top of the BSC to the ceiling. Higher versions may be accepted, provided the 400 mm over the BSC is available to measure air velocity above the exhaust filter, and to have enough space for changing the filter and for ducting and/or a thimble connection to outlet. The Biosafety Cabinet should be tested and comply with the following requirements (at site):- 1. Down flow velocity and Volume Test; 2. Inflow Velocity Test; 3. Airflow Smoke Pattern Test; 4. HEPA Filter Leakage Test; 5. Cabinet Leakage Test; 6. Electrical Leakage : Ground Circuit Resistance and Polarity Test; 7. Lighting Intensity Test; 8. Vibration Test; 9. Noise Level Test; 10. UV Lamp Intensity Test; 11. Alarms and indicators test (if provided); 12. The differential pressure gauge should be calibrated; Spare accessories required : HEPA Filter (one)-dimensions same as above, Prefilters (two) | Piece
6.4 | Incubator- TOMY High pressure steam sterilizer ES-315 use for sterilization and decontamination of laboratory equipments and waste. The vertical autoclave with a top-open lid, which can be operated easily with a single hand or foot. Operating temperature range (during sterilizing): 105 - 132°C, Operating pressure range: 0 - 186kPa, Maximum operating pressure 216kPa, Temperature control: Digital, microprocessor controlled, Temperature display / display range: Digital / -15 - 180°C, Pressure display / display range: Analog / 0 - 400kPa, Heat source: 2.0kW electric heater, Timer control: Digital, microprocessor controlled, Timer display: Digital. Safety devices: • Inside the chamber overheat prevention, • Outer wall of the chamber overheat prevention, • Overpressure prevention, • Temperature sensor disconnection prevention, • Empty heating prevention, • Leakage breaker, • Safety valve. Malfunction prevention devices: • Low water level detection, • Exhaust valve knob open/close detection, • Insufficient sterilization detection, • Chamber lid open/close detection. Chamber material: SUS304. Capacity / material of exhasut bottle: 3ℓ / polyethylene. Type / material of lid gasket: Accommodating internal pressure / silicone rubber. Dimensions of main unit (mm): 490W x 560D x 1090H (Height from floor to control panel:875). Net weight: 80kg. Rated Voltage AC 120 / 220 / 230 / 240V. Power input: 20A (120V), 10A (220V), 9A (230 / 240V), Power supply requirements: Single-phase 220 / 230 / 240V AC (50 / 60Hz) 15A or above. Power Comsumption (calorific value): 2kW (1720kcal/h). Accessories: • Stainless steel basket with slatted bottom plate 1, • Water level sensor 1, • Inspection sheet 1, • Warranty card 1, • Customer card 1, • Operator's manual 1,• Clear plastic accessory case 1, • Attachment screw for clear plastic accessory case 1. Stainless basket (φ300 x 182mm) | Piece
6.41 | High Spec Centrifuge-Max. Speed 15200rpm Capacity 1.6L (4 x 400) Refrigerated No Max. RCF 25830 x G Hertz 50/60Hz Voltage 230V Control Microprocessor Display Digital Height (English) Exterior 14.2 in. Height (English) Open Cover 34.3 in. Height (Metric) Exterior 36cm Height (Metric) Open Cover 87cm Length (English) Exterior 23.8 lb. Length (Metric) Exterior 60.5cm Memory 99 programs (5 with direct access key) Motor Type Direct, brushless induction, low-profile Net Weight (English) 126.8 lb. Net Weight (Metric) 57.5kg Timer 99 hrs., 59 min. plus HOLD Width (English) Exterior 17.3 in. Width (Metric) Exterior 44cm Certifications/Compliance IVD compliant; IEC 61010-1, IEC 61010-2-020, IEC 61010-2-101; 230V models meet EN 292, EN 61326, EN 55011B. UL, CE, CSA, Certified Biosafety - See more at: http://www.thermoscientific.com/en/product/heraeus-multifuge-x1-centrifuge-series.html#sthash.JsGJZZUE.dpuf | Piece
6.42 | Autoclave-Chamber Capacity: 53 Litters, Temperature setting range: 105°C to 132 °C,Temperature control: Digital, Microprocessor controlled, Pressure display range: 0 to 400 kPa, Maximum operating pressure: 216kPa, Operating pressure range: 0-186kPa, Sterilization heat source: 2.0kW electric heater, Dimensions of main unit (mm): 490 W x 560D x 1090H (Height from floor to control panel:875) , Chamber diameter: 325 x 740mm, Net Weight: 80 kg, Power input: 20A (120V), 10A (220V), 9A (230/240V) | Piece
6.43 | Freezer (-80 deg); uprightfreezer, thermo Scientific. Model: 900 series- 86° C | Piece
6.44 | 2-8 Refrigerator-Laboratory refrigerator of Capacity 1500L, temperature range 2℃-8℃, accuracy 0.1℃, Audio and visual alarm for: Height or low temperature; System & Sensor failure; Door ajar(except BXC-V50M); Power failure alarm , Auto defrost,Rigid polyurethane insulation material, Height efficiency air-cooled condenser, made of copper , Three-ply transparent glass door with lock ,Stealing joint strip with magnetic force ,Power Supply 220V±10%, 50Hz/60Hz with adequate lighting, Stainless steel Material (Internal Cold-rolled steel with anti-bacteria powder coating, External Cold-rolled steel with anti-bacteria powder coating ) | Piece
7 | Food Lab
7.1 | Analytical balance: Capacity up to 320g and Readability 0.01 mg / 0.1mg. MonoBloc HighSpeed cell, full auto internal adjustment. Touchscreen, weighing-in guide, test management. Dismentalable weighing chamber, dishwasher proof. | Piece
7.2 | Spectophotometer: Measuring wavelength range:220 to 800 nm. Spectrum band width:3 nm. Pathlength: 0.2 mm, 0.7 mm (manual selection). Light source: Xenon flash lamp. Monochrometer: Holographic grating. Detector: Photo diode array. Auto wiping function. Auto sample mount function. Quantitation rate:Pathlength 0.2 mm, 1 to 75 OD (50 to 3,700 ng / µL). | Piece
7.3 | Vortex: Speed: 3000rpm Tube Capacity Power:Voltage 230V Platform Material Foam Hertz 50/60Hz Amperage 0.3A | Piece
7.4 | Stomacher | Piece
7.5 | Double-door refrigerator: Transparent Double door high performance refrigerator with temperature Control Microprocessor. Min. capacity (volume) 1200L. Must be frost free and CFC free. Temperature Range (Metric) 1° to 8°C Power:Voltage 250V/50Hz. | Piece
7.6 | Deep Freezer (-86): Temperature Range (Metric) -50° to -86°C. Capacity :28 cu. ft. Min. capacity (volume) 700-800L. Power Voltage: 230V Amperage 12A. | Piece
7.7 | Deep Freezer (-40): Temperature Range (Metric) -20° to -40°C. Capacity :17.3 cu. ft. Min. capacity (volume)250-300L with CO2 & Liquid Nitrogen option. Power Voltage: 230V Amperage 12A. | Piece
7.8 | Centrifuge with refrigerator: Max. Speed 15200rpm Capacity 1.6L (4 x 400) Refrigerated No Max. RCF 25830 x G Hertz 50/60Hz Voltage 230V Control Microprocessor Display Digital Height (English) Exterior 14.2 in. Height (English) Open Cover 34.3 in. Height (Metric) Exterior 36cm Height (Metric) Open Cover 87cm Length (English) Exterior 23.8 lb. Length (Metric) Exterior 60.5cm Memory 99 programs (5 with direct access key) Motor Type Direct, brushless induction, low-profile Net Weight (English) 126.8 lb. Net Weight (Metric) 57.5kg Timer 99 hrs., 59 min. plus HOLD Width (English) Exterior 17.3 in. Width (Metric) Exterior 44cm Certifications/Compliance IVD compliant; IEC 61010-1, IEC 61010-2-020, IEC 61010-2-101; 230V models meet EN 292, EN 61326, EN 55011B. UL, CE, CSA, Certified Biosafety - See more at: http://www.thermoscientific.com/en/product/heraeus-multifuge-x1-centrifuge-series.html#sthash.JsGJZZUE.dpuf | Piece
7.9 | Waterbath: Microprocessor controlled with integrated autodiagnostic system with fault indication, digital timer from 1 min. up to 999 hours for:ON, DELAYED ON, HOLD or HOLD set temperature with HOLDING time facility.Resolution of display and setting accuracy: 0,1 °C below 99,9 °C, 1 °C above 100 °C. Material 1.4301 (ASTM 304), corrosion resistant.Temperature range: Min. 5 °C above ambient up to +200 °C heating capability. Dimension(mm): 600 x 350 x 230 mm approx. or 40-45L. Power: 230-250V/50Hz. | Piece